## Sheet1
id | OSF dataset | OSF link | NIFTI filename | Data type | contact | site name | MRI vendor | MRI version | MRI field | sample type | phantom version | phantom serial number | phantom temperature | age | sex | sequence name | sequence type | matrix size | resolution | dimension | TR | echo time | TI | bandwidth | T1 - genu (WM) | T1 - splenium (WM) | T1 - deep GM | T1 - cortical GM | T1 - NIST sphere 1 | T1 - NIST sphere 2 | T1 - NIST sphere 3 | T1 - NIST sphere 4 | T1 - NIST sphere 5 | T1 - NIST sphere 6 | T1 - NIST sphere 7 | T1 - NIST sphere 8 | T1 - NIST sphere 9 | T1 - NIST sphere 10 | T1 - NIST sphere 11 | T1 - NIST sphere 12 | T1 - NIST sphere 13 | T1 - NIST sphere 14
1.001 | matthewgrechsollars_ICL_NIST | https://osf.io/7k96d/download/ | 20200121_matthewgrechsollars_ICL_NIST/20200121_matthewgrechsollars_ICL_NIST_Magnitude_T1map.nii.gz | Magnitude | Matthew Grech-Sollars | Imperial College London | Siemens | Prisma VE11C | 3 | NIST | 130 | 102 | 21.89 |  |  | se_TI50/400/1100/2400 | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2400] | 130 |  |  |  |  | [2176.4 2201. 2169. 2192.4 2211.2 2250.6 2208.1 2106.6 2222.3 2187.3 2173.7 2255.1 2182.7 2132.7 2323.5 2146.3 2175.9 2157.6 2120.4 2104.6 2106.9 2165.4 2141.4 2139.1 2160.4 2333.6 2074.7 2101.8 2162.8 2147.7 2139.8 2151.1 2181.6 2193.3 2064. 2337.3 2160.5 2205.1 2216. 2157.7 2127.3 2097.6 2195. 2142. 2175.7 2109.9 2211.5 2268.2 2110.2 2186.2 2158.8 2175.9 2177.5 2129.2 2166.1 2119.3 2132.5 2091. 2343.5 2021.5 2083.6 2123.4 2144.5 2136.2 2143.4 2152. 2141.4 2105.5 2307.3 2020.9 2222.2 2193.5 2168.7 2178.2 2163.4 2159.5 2077. 2192.5 2269.9 2100. 2086.7 2163.5 2102. 2214.8 2108.4 2190.8 2079.7 2344. 2091.5 2162.9 2120.4 2180.3 2098.5 2142.4 2187.3 2208. 2107.5 2235.7] | [1437. 1481.8 1451. 1428.2 1395.4 1411.7 1392.4 1461.3 1440.3 1416. 1436.4 1421.4 1413.6 1495.4 1395.2 1405.4 1473.6 1455.5 1401.4 1412.8 1399.6 1423.8 1357.8 1421.5 1452.5 1393.6 1427.6 1415.3 1430.6 1395.9 1423.7 1441.8 1445. 1439.6 1399.9 1477.9 1375.9 1479.6 1470.9 1429.7 1445.8 1439.6 1433.3 1430. 1409.9 1406.8 1397.1 1450.1 1389.1 1505.7 1407.2 1429.6 1401.7 1443.6 1426.7 1420.3 1441.9 1430. 1440.9 1350.7 1533. 1448.5 1413.9 1415.8 1410.7 1415.4 1422.2 1412.8 1421. 1446.7 1386.2 1420.1 1443.1 1405.7 1428.4 1438.4 1422.4 1406. 1420. 1403.9 1418.1 1411.8 1399.8 1423.2 1391.9 1438.9 1362.1 1447.7 1441.2 1436. 1423.2 1395.5] | [ 993.4 978.5 977.9 990.4 993.2 991.9 990.9 981.7 1006.5 966.3 986.1 977.2 998.6 978.9 975.7 996.2 950.3 974. 987.6 1000.8 984.1 988.9 992.4 986.8 1002.9 980.9 980.7 977.1 993. 971.2 974.4 989.6 981.7 970.3 975.2 978.5 983.6 1011.9 972.4 981.3 998.2 982.8 989.5 971.1 999.9 971.4 989.5 973.2 980. 978.6 989.4 977.8 993.9 976.4 985.4 971.1 1018.2 967.2 1011.7 982.1 992.8 987.3 992. 981.4 986.2 998.8 976.8 1001.8 1009.5 991. 984.3 973.9 981.4 970.6 996. 999.6 978.2 988.8 990.5 989.1 1001.1 1005.4 976.6 970.2 1007.2 979.6 977.4 978. 1000.2 1015.8] | [700. 692.5 697.2 694.8 698.1 692. 702.9 689.4 706.7 684.4 696.9 695. 692.7 695.5 688.9 700.7 693.9 702.5 697.5 695.5 688. 697.5 704. 687.3 698. 686.2 698.2 699. 696.9 694.5 692.3 693.3 694.5 696.3 702.1 677.9 712. 694.4 708.4 696.8 698.4 693.5 698.4 695.3 692.7 697.5 670.9 706. 697.6 699.4 693.2 703.1 688.7 698.6 686.4 696.5 697.2 694.5 675.5 704.4 702.7 701.6 702. 691.6 695.8 689.8 703.4 685.8 693.3 679.6 681. 694. 687.5 687.4 680.7 692.8 691.7 714.1 700.1 690. 693.2 697.4 704.3 702.6 671.9 677.6 689.1 686.1 684.8 676. 675.9 689.3] | [486.3 485.5 487.2 489.1 492.1 494.1 488.8 494.9 485.2 490.2 497.4 481.7 486.1 481. 492.7 486.1 492. 487.3 498.1 498.6 488.8 494.9 489. 499.9 487. 494. 490. 492.5 499.9 491.9 485. 489.6 485.9 489.7 485.7 496.1 484.4 489.8 487. 483.6 490.9 495.2 487.4 502.2 484.7 496.9 482.4 493.8 491.9 499.2 484.7 487.7 498.1 483.5 490.3 485.3 490.7 489. 494.8 487.4 486.5 499.1 494.3 496. 494.5 494.5 492.6 490.4 496.8 492.7 495.7 502.5 482.5 487.6 487. 496.6 490.8 494.6 488.9 495.7 492.4 491.6 492.9 493. 492.8 491.4 495.1 493.4 487.1 491.8 492.6] | [345.5 346.6 344.9 345.4 343.7 352.4 342.7 345.4 343.6 345.4 345.7 350. 344.4 346.6 351.6 350.2 352.5 349.4 345.9 348.7 345.3 346.7 347.3 347.2 347.9 343.2 343.9 340.6 347.4 345.7 344.7 344.4 346.8 351.2 341.1 348.5 348.9 350.6 348.5 348.2 346.7 346.3 343.6 345.9 343.4 348.6 345. 350.3 347.2 343.8 344.7 343.7 347.9 344.2 347. 345.9 354.4 336.2 353.6 342.5 350.8 348.3 348.7 347.7 346.8 346.8 349.3 345.5 339.3 342.3 351.9 340.5 341.4 346.5 341.4 346.2 346.5 345. 352.5 338.4 341.4 352.7 349. 351.4 347.8 347.7 346.6 350. 353.3 342.3 340.2 345.8 343.7 342.9 347.7 348.1 349.5 346.2 346.8 349.3 349.1 344.8] | [245.8 248.4 245.6 243.3 238.6 236.1 240.1 238.5 244.6 244. 242.7 247.8 245. 246.3 243.3 245.7 240.6 247.4 236.4 241.9 238.6 237. 241.3 241.8 240.8 244.7 239.4 247. 244.1 241.4 246.4 246.6 242.9 242.2 245.7 242.1 237.2 239.9 247.5 240.5 243.3 237.1 239.3 238.8 241.1 240.5 243.8 239.2 251.5 241.4 246.6 244.8 243.8 242.5 243.7 244.5 241.6 240.7 240.6 247. 238.3 238.7 237.8 237.9 243.7 238.8 241.4 245.1 236.3 243. 248.6 247.6 247.4 246.4 242.3 241.9 244.3 236.5 237.3 241.1 237.5 237.5 242. 243.7 243.8 242.7 247.2 245.1 240.3 237. ] | [175.1 173.4 178.8 180.8 181.3 177.8 176.1 177.3 178. 173.1 172.8 174.3 171.2 172.5 173.7 175.3 177.6 182.5 169.3 183.2 174.5 181.7 175.3 182.6 173.8 172.9 178.4 182.5 170.3 178.5 171.2 176.9 172. 177.9 172.3 175.2 171.9 183.3 168.9 181.9 174.3 180.4 171.8 184.2 172. 182.4 171.3 182.1 168.9 188.1 178.7 170.5 180.7 171.2 176.9 170.6 175.9 171. 180.2 169.8 186.3 179.4 172.2 179.4 176.1 180.8 174. 179. 175. 173.1 176.1 178.5 173.7 174.2 175.8 176.1 171.1 176.3 172.2 173.3 176.7 178.7 179.3 176.2 174.2 175.5 180.7 174.3 173.3 173.2 174.4 178. ] | [183.8 181.6 185.3 184.6 118.4 118.1 116.5 116.1 121.5 121.2 127.3 175.2 171. 184.4 158.9 171.3 119.8 124.4 178. 163.6 117.8 117.6 120.4 111.5 119.5 168.1 121.3 180.9 176.7 118.9 187.3 122.2 183.1 166.3 178.1 161.7 156.3 129.5 171.2 119. 171.3 116.6 118.7 117.1 161.8 113.4 121.8 115.9 115.2 121.9 182.8 175.9 121.9 182.8 123.1 180.8 120.2 183.6 118.9 169. 119.8 187.1 114.3 163.8 114.3 120.8 116.8 117.9 116.7 162.8 126.4 177.2 120.6 183.3 172.8 181.9 123.1 172. 118.9 117.6 130. 119.2 118.1 115.6 117.6 163.5 155.3 128.7 179. 170.7 119. ] | [108.6 96.3 81.1 72.4 90.5 100.2 94.6 81. 101.5 102.6 95.4 90.1 96.1 118.2 77.4 87.4 81.4 83.1 69.4 80.7 100.6 60.6 95.3 102.7 80.5 90.1 77.4 105.5 88.3 107.9 84.1 103.4 71.3 128.3 2.4 4. 103.2 80.6 89.1 76.2 98.8 66.3 102.7 54.3 132.5 1.8 61.7 94.4 79.4 109.7 80.5 106.8 82.4 114.1 9.5 127.8 1. 70.5 102.8 71. 104.5 75.4 100.5 63.2 107.9 3.2 128.8 2.8 94.7 89.6 99.3 92.1 105.1 83. 116.4 1.2 134.7 75.8 104.5 71.7 97.9 3. 105.6 79.6 125.6 92.9 88.8 101.8 103.1 71.5 109. 70.7] | [101. 86.5 2.5 1.1 2.6 56.5 68.2 71.5 71.7 75.8 90.6 78.8 72.4 9.5 80.3 79.8 2.4 2. 1. 1. 60.2 2. 74.5 2. 109.3 72.9 87.4 86. 90.2 69.4 70.3 64.3 91.2 2. 106. 9.5 2.8 9.5 3.3 1.2 2.1 57.2 7. 3.2 2.8 82. 45.2 64.2 99. 86.8 89.4 59.3 63.5 92.8 68.8 2.5 69.4 67. 58.7 2.5 10.2 5.6 62.8 2.8 69.4 2.8 91.1 4.7 59.5 88.4 90. 92. 73.7 84.1 70.9 87.1 1.8 110.6 2.4 3.1 2.1 9.1 72.5 3.3 70.8 1.2 105.8 82.8 91.9 89.9 1. 2.5 1.9 83. 9.5 2.7 76.1 75.3] | [ 4. 2.8 4.2 77.1 69.6 1.2 80.8 63. 2. 69.9 67.4 2.1 2.7 4. 81.8 4. 2. 44. 4.5 89.2 86.3 3.6 96.8 3.6 83.9 68.1 74.2 9.5 2. 3.1 89.9 1.1 4. 2.1 72.5 9.5 76.7 2.7 90.6 4.8 134.3 101.8 4. 98.6 1.8 96.4 4.5 110.3 4.9 105.1 9.1 136.5 3.4 85.5 3.6 62. 2. 77.6 48. 74.6 3.1 79.7 3.3 111.5 82.8 2.8 72.2 3.4 71.3 78.3 68.2 1.3 1.1 102.2 6.1 87.1 3. 9.5 44. 10. 6.6 96.2 3.4 104.7 2.4 61.3 75. 14.6 52.3 67.9 2.5 63.4 10. 63.3 2.4 1.2 72.2 78.5 95. ] | [ 75.6 80.7 79.2 72.4 106. 90.5 3.7 3. 67.3 1.3 102.1 3.6 106.7 56.7 96.7 103. 2. 111.1 4. 111.8 47.6 116.3 4. 106. 2.7 69.3 3.5 83.3 6. 70.4 3.3 93.1 3.1 112.6 116.5 5. 102.6 2.5 123.2 3. 93.7 80.6 94.6 7.3 83.7 103. 4. 100.2 2.1 103. 2.5 86.3 4. 106.2 9.5 135.6 63. 82.9 89.6 1.1 119.5 3. 108.9 3. 125.9 2.8 149. 3. 54.8 82.5 8. 9.1 2. 52. 4.9 68.7 10. 111.4 123.2 3.6 64.2 86. 79.2 73.6 62.5 78.7 2.4 101.3 48.2 52. 2.5 3.6 1.2 61.7 1. 71.5 85.2 99.9] | [ 57.9 71.3 1.2 2.8 1.5 97.1 4.4 84.3 2.5 166.6 108.6 1.1 71.9 4.5 52.7 1.8 131.5 2.4 142.7 1.2 9.1 82.8 2. 58.3 5. 68.8 4.7 86.6 55.1 90.1 40.9 91.5 91.1 1.3 1.1 86.5 8. 66.3 3.7 90. 2.5 89.4 62.9 85.6 48.7 108.4 5.4 101.4 9.1 88.5 100.9 5. 75.5 53.5 3.3 53.3 6.3 4. 4.9 5. 64.2 4.9 92.4 3.3 2. 81.2 3. 1. 110.2 2.5 146.8 7. 1.3 70.4 10.6 2. 60.2 1.2 132. 2.7 163.8 1. 56.3 1.3 8. 3.8 2.7 87.6 79.2 4. 2.8 2. 1.2 9.1 1.4 80.5]
1.002 | matthewgrechsollars_ICL_NIST | https://osf.io/7k96d/download/ | 20200121_matthewgrechsollars_ICL_NIST/20200121_matthewgrechsollars_ICL_NIST_Complex_T1map.nii.gz | Complex | Matthew Grech-Sollars | Imperial College London | Siemens | Prisma VE11C | 3 | NIST | 130 | 102 | 21.89 |  |  | se_TI50/400/1100/2400 | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2400] | 130 |  |  |  |  | [2176.4 2201.5 2169. 2192.5 2211.2 2250.6 2207.9 2107.5 2222.8 2187.2 2174.3 2255.1 2182.9 2132.6 2324.2 2146.4 2176. 2157.7 2120.6 2105. 2106.9 2165.4 2141.4 2139.1 2160.5 2333.6 2074.9 2101.8 2162.8 2148.4 2140.3 2151.1 2181.9 2193.4 2064. 2337.6 2160.9 2205.3 2216. 2158. 2127.5 2098.1 2195. 2142.3 2175.8 2110.1 2212. 2268.7 2110.6 2186.8 2159.4 2175.9 2177.6 2129.4 2166.1 2120.2 2132.5 2091.2 2343.8 2021.5 2084. 2123.6 2144.8 2136.2 2143.9 2152.2 2141.4 2105.5 2307.2 2020.9 2222.7 2193.5 2168.8 2178.2 2163.5 2159.6 2077.3 2192.8 2270.1 2100.1 2086.9 2163.6 2102.1 2215.4 2108.5 2191. 2079.7 2344. 2091.5 2163.4 2120.5 2180.2 2098.7 2142.4 2187.3 2208. 2107.8 2235.8] | [1437. 1481.8 1451. 1428.2 1395.4 1411.7 1392.5 1461.3 1440.3 1416. 1436.4 1421.5 1413.6 1495.4 1395.2 1405.4 1473.6 1455.6 1401.3 1412.9 1399.7 1423.6 1357.8 1421.4 1452.5 1393.6 1427.6 1415.4 1430.6 1395.9 1423.6 1441.8 1444.9 1439.6 1399.9 1478. 1375.9 1479.7 1470.9 1429.6 1445.8 1439.6 1433.3 1430. 1410. 1406.8 1397.2 1450.2 1389.1 1505.7 1407.3 1429.6 1401.7 1443.8 1426.8 1420.4 1441.9 1430.1 1440.9 1350.7 1533. 1448.5 1413.9 1415.8 1410.7 1415.4 1422.4 1412.8 1421. 1446.7 1386.2 1420.1 1443.3 1406.1 1428.5 1438.4 1422.3 1406. 1420. 1403.8 1418.1 1411.8 1399.9 1423.2 1392. 1438.9 1362.1 1447.7 1441.3 1436. 1423.2 1395.5] | [ 993.4 978.5 977.9 990.4 993.2 991.9 990.9 981.6 1006.5 966.3 986.1 977.3 998.6 978.9 975.8 996.2 950.4 974. 987.6 1000.8 984.1 988.9 992.4 986.8 1003. 980.9 980.7 977.1 993. 971.2 974.4 989.6 981.7 970.3 975.2 978.4 983.6 1011.8 972.4 981.2 998.2 982.8 989.5 971.1 999.9 971.4 989.5 973.2 980. 978.6 989.4 977.8 993.9 976.4 985.4 971.2 1018.2 967.2 1011.7 982. 992.8 987.3 992. 981.4 986.2 998.8 976.8 1001.9 1009.5 991. 984.3 973.9 981.4 970.6 996. 999.6 978.2 988.8 990.6 989.1 1001.2 1005.4 976.6 970.2 1007.1 979.6 977.4 978. 1000.2 1015.8] | [699.9 692.4 697.2 694.8 698.1 692. 702.9 689.4 706.7 684.3 696.8 694.9 692.7 695.5 688.8 700.6 693.9 702.5 697.5 695.5 687.9 697.4 704. 687.3 698. 686.2 698.2 699. 696.9 694.5 692.3 693.3 694.5 696.3 702.1 677.9 712. 694.4 708.2 696.7 698.4 693.4 698.4 695.3 692.7 697.5 670.9 705.9 697.6 699.4 693.2 703. 688.7 698.6 686.4 696.5 697.1 694.3 675.4 704.4 702.7 701.5 701.9 691.6 695.8 689.8 703.3 685.7 693.3 679.5 681. 693.9 687.5 687.4 680.7 692.8 691.7 714.1 700. 690. 693.2 697.4 704.2 702.6 671.8 677.6 689.1 686.1 684.8 676. 675.9 689.3] | [486.3 485.5 487.2 489.1 492.3 494.1 488.8 494.9 485.2 490.2 497.4 481.8 486.1 481. 492.8 486.1 492. 487.3 498.1 498.6 488.9 495.1 489. 499.9 487. 494.3 490. 492.5 499.9 491.9 485. 489.6 485.9 489.7 485.7 496.2 484.5 489.8 487.1 483.6 490.9 495.2 487.5 502.3 484.7 496.9 482.5 493.8 491.9 499.2 484.7 487.7 498.1 483.5 490.3 485.4 490.7 489. 494.8 487.4 486.5 499.2 494.4 496. 494.5 494.5 492.6 490.4 496.9 492.8 495.7 502.6 482.5 487.6 487. 496.6 490.8 494.6 488.9 495.8 492.4 491.6 492.9 493.1 492.8 491.4 495.1 493.4 487.1 491.8 492.6] | [345.5 346.6 344.9 345.4 343.7 352.4 342.7 345.4 343.6 345.4 345.7 349.9 344.4 346.6 351.6 350.2 352.5 349.4 345.9 348.7 345.3 346.7 347.3 347.2 347.9 343.2 343.9 340.6 347.4 345.7 344.7 344.4 346.8 351.2 341.1 348.6 348.9 350.6 348.5 348.2 346.7 346.3 343.6 345.9 343.4 348.6 345. 350.3 347.2 343.8 344.8 343.7 347.9 344.2 347. 345.9 354.5 336.2 353.6 342.5 350.8 348.3 348.8 347.7 346.8 346.8 349.3 345.5 339.4 342.3 351.9 340.5 341.4 346.5 341.4 346.2 346.5 345. 352.4 338.4 341.4 352.7 349. 351.4 347.8 347.7 346.6 350. 353.3 342.3 340.2 345.8 343.7 342.9 347.7 348.1 349.5 346.2 346.8 349.3 349.2 344.8] | [245.8 248.4 245.6 243.3 238.6 236.1 240.1 238.5 244.6 244. 242.7 247.8 245. 246.3 243.3 245.7 240.6 247.4 236.4 241.9 238.6 237. 241.3 241.8 240.9 244.7 239.4 247. 244.1 241.4 246.4 246.6 242.9 242.2 245.7 242.1 237.2 239.9 247.5 240.5 243.3 237.1 239.3 238.8 241.1 240.5 243.8 239.2 251.5 241.4 246.6 244.8 243.8 242.5 243.7 244.5 241.6 240.7 240.6 247. 238.4 238.7 237.8 237.9 243.7 238.8 241.4 245.1 236.3 243. 248.6 247.6 247.4 246.4 242.3 241.9 244.3 236.5 237.3 241.1 237.5 237.5 242. 243.7 243.8 242.7 247.2 245.1 240.3 237. ] | [175.1 173.4 178.8 180.8 181.3 177.8 176.1 177.3 178. 173.1 172.9 174.3 171.2 172.5 173.7 175.3 177.6 182.5 169.3 183.2 174.5 181.7 175.3 182.6 173.8 172.9 178.4 182.5 170.3 178.5 171.2 176.9 172. 177.9 172.3 175.2 171.9 183.3 168.9 181.9 174.3 180.4 171.8 184.2 172. 182.4 171.3 182.1 168.9 188.1 178.7 170.5 180.7 171.2 176.9 170.6 175.9 171. 180.2 169.8 186.3 179.4 172.2 179.4 176.1 180.8 174. 179. 175.1 173.1 176.1 178.5 173.7 174.2 175.8 176.1 171.1 176.3 172.2 173.3 176.7 178.7 179.3 176.2 174.2 175.5 180.7 174.3 173.3 173.2 174.4 178. ] | [127.7 128.7 129.9 128.7 118.4 118.1 116.5 116.1 121.5 121.2 127.3 125.5 124. 130.7 117.3 123.7 119.8 124.4 126.5 120.3 117.8 117.6 120.4 111.5 119.5 122.2 121.3 129. 126.5 118.9 131.9 122.2 129.9 121.2 126.9 118.9 117.1 129.5 123.1 119. 124.1 116.6 118.7 117.1 120.3 113.3 121.8 115.9 115.2 121.9 129.1 126.6 121.9 129.6 123.1 128.4 120.2 129.8 118.9 122.9 119.8 131.1 114.3 120.3 114.3 120.8 116.8 117.9 116.7 119.3 126.4 126. 120.6 129.4 125. 128.7 123.1 124.6 118.9 117.6 130. 119.2 118.1 115.6 117.6 120.7 114.9 128.7 127.5 123.4 119.1] | [ 98.7 96.4 81.1 72.4 90.6 91.6 86.8 81. 92.8 93.3 87.8 90. 96.2 106.2 77.4 80.6 81.4 83.1 69.5 80.8 92.2 57.6 95.3 94.5 80.5 90.1 77.4 95.3 81.8 97.9 84.1 93.7 71.3 114. 1.6 1.2 94. 80.6 89.1 76.3 90.6 66.3 93.8 54.3 117.4 4. 58.7 94.3 79.4 98.9 80.5 96.7 76.6 102.8 4. 113.8 2.5 70.5 93.4 71. 95.2 75.4 91.7 60.1 98.1 1.8 114.8 9.5 94.7 82.9 90.6 84.9 95.2 82.9 104.5 1.3 119.3 75.8 94.8 71.8 89.7 10. 96.1 79.6 112.2 85.5 88.8 92.6 93.6 71.6 99. 70.7] | [101. 86.5 6.2 10. 3.3 58.2 70.5 73.3 71.9 75.8 94. 78.9 72.3 10. 83.3 81.8 1.3 10. 10.6 9.5 61.8 3.2 74.5 1. 109.4 75.3 89.7 86. 90.2 69.4 70.3 65.6 91.2 6.9 106. 4.3 4.6 2.7 3.6 1.2 10. 57.3 2.8 2.7 2.7 84.3 44.9 66.2 99. 86.9 89.5 59.2 64.8 92.7 70.4 1.7 71.4 67. 60.2 9.5 2.8 4. 64.5 1. 71.3 1. 94. 2.4 61.3 88.4 90. 92. 73.7 84.2 70.9 87.1 4.2 110.5 10. 2.5 1.3 1.6 74.8 1.1 70.8 1.8 105.8 82.9 92. 89.9 10. 8.7 2.4 86.3 3.3 2.1 76.2 78.2] | [ 3.3 1.1 6.7 76.9 69.8 10. 80.8 71.9 3.2 69.7 77.1 10. 10.6 3.4 81.8 2.7 3.9 44. 2.8 106.5 86.4 2.6 96.9 10.3 83.8 77.8 74.1 5. 9.5 5.9 89.8 9.5 10.6 2.6 72.4 3.2 76.7 2. 90.6 1.3 134.3 102. 9.5 98.8 2.7 96.4 5.7 110.3 9.4 105.1 3.4 136.5 1.3 85.6 3.4 61.8 4.8 77.5 48.4 74.7 9.5 79.9 8. 111.5 82.8 1.6 72.2 3. 71.3 78.3 68.1 1.8 9.5 102.2 8.3 87.1 2.8 1.3 44.4 3.3 3. 96.2 2. 104.7 2. 68.8 75. 35.3 57.8 67.9 1. 63.3 10.2 63.2 1.2 35.8 83.5 78.5 94.9] | [ 75.5 81. 79.2 72.3 106.1 90.3 3.3 1.9 89.3 10.7 101.9 4. 106.6 71.3 97.8 103. 1. 111.3 10.7 111.8 51.9 116.4 9.5 106.1 3.5 69.2 3.3 83.5 8. 70.4 10. 93.4 10. 112.6 116.5 10.2 102.5 10. 123.2 8. 93.4 81.3 94.4 10. 116.5 103. 9.5 100.3 4.4 103.4 9.5 86.4 9.5 106. 1. 135.6 81.2 82.8 89.5 4.7 119.7 9.5 108.9 9.5 125.8 4.1 148.9 10. 68.7 82.5 4.2 4. 3.7 64.1 7.2 68.6 5. 111.5 123.1 11.1 82.4 86.3 109.4 99.2 62.5 78.9 10.6 101.2 57.8 63.7 2.6 5. 1.2 78.8 10.6 94.6 85.8 100. ] | [ 89.1 113.3 5.9 4. 1.5 96.7 3.4 84.5 6. 167. 108.7 4.7 120.9 2.8 79.6 2.3 131.2 7. 142.6 10. 10.6 82.6 2.9 87.3 10.3 112.3 2. 88.3 81.6 92.5 54. 90.9 91.9 2.8 10.9 86.5 8. 108.1 10. 175.4 10. 89.2 98.4 79.8 66.4 108.6 10.3 101.2 3. 167.5 106.3 5.5 75.4 59.1 1.9 51.6 10.6 9.5 10.7 3.8 100.5 8. 92.2 10. 2.6 81.7 9.1 10. 110.5 1. 147.5 2.3 10.6 70.1 6. 8.2 94.7 10. 132. 8. 163.7 3.8 84.3 1.5 10.6 8. 3.4 91.6 80.7 10. 10. 10. 8.7 10.7 5. 140.8]
2.001 | siyuanhu_casewestern_NIST | https://osf.io/46nd9/download/ | 20200124_siyuanhu_casewestern_NIST/20200124_siyuanhu_casewestern_NIST_Magnitude_T1map.nii.gz | Magnitude | Siyuan Hu | Biomedical Engineering, Case Western Reserve University | Siemens | Skyra | 3 | NIST | 130 | 6 | 19.1 |  |  | SE_IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | [5 empty] | [1818.4 2060.6 1959.5 1987.9 1934.6 2087.8 2027.9 1881.4 1906.4 1914.5 1981. 1907.9 1920.2 1906.3 1973. 1874.3 1993.3 1922.1 1996.4 1953. 1965.3 1941.2 1999.2 1902. 2021.8 1852.4 1953.4 1881.6 1913.2 1918. 1960.2 1868.1 1990.8 1950.4 1925.9 1993.6 1994.5 1920.7 1975.2 1926.3 1956.6 1935.9 1975.1 1919.3 1917.4 1947.2 2000.4 1973.6 1939.8 1954.2 1908. 1940.3 1936.7 1931.9 1912.7 1989.3 1921.1 1970.7 1959.7 1931.1 1990.5 1940.7 1965. 1929.9 1966.6 1905.6 1934. 1907.3 1964.7 1954.1 1938.9 1898.1 1902.8 1945. 1927.3 1950.3 1919.4 1903.5 1955.4 1982.2 1944.2 1983.5 1904.7 1962.6 1911.9 1897.2 1933.3 1948.3 1950. 1925.4 1922.9 1946.5 1986.8 1944.6 1942.7] | [1281.6 1308.8 1287.4 1301.3 1286.8 1318. 1292.9 1271.2 1276.1 1289.3 1280.9 1281.8 1277.1 1279.7 1282. 1311.1 1283.2 1313.8 1271.2 1301.6 1284.1 1297.9 1279.3 1302.5 1278.1 1289.1 1279. 1288.6 1286. 1294. 1295.1 1288.6 1281.1 1285.2 1288.4 1286.9 1285.7 1285.7 1280.8 1278.2 1285.3 1294.3 1284.1 1287.3 1286.3 1274.4 1294.2 1273.3 1295.4 1281.2 1292.1 1278. 1294. 1276. 1300.4 1270.7 1291.9 1281.1 1292.5 1267.7 1298.6 1262.5 1305.9 1263.3 1289.9 1279.6 1300.5 1281.6 1287.4 1270.1 1295.8 1279.5 1292.8 1280.8 1281.4 1280.1 1289.6 1272.1 1281.4 1278.5 1294.8 1292.8 1287.5 1303.6 1284. 1299.9 1280. 1285.7] | [842.4 832.9 830.5 831.1 835.4 818.3 823.3 831. 834. 831. 834. 824.1 823. 843.8 830.3 831.2 821.6 825.8 820.9 840. 832.5 817. 831.8 817.4 839.4 824.1 835.6 824.6 829.5 819.6 822.7 834.2 825.3 827.6 823.8 824.8 826.8 822.3 833. 819.1 835.1 818.9 822.1 824.6 825.5 835.6 823.4 834.9 822. 827.1 812.7 820.6 830. 823.9 818.8 823.1 818.6 821.8 818.3 835.6 269.6 814.2 820.8 819.6 829.7 830. 827.4 826.3 817.7 825.7 833.1 810.7 817.6 812.3 811.9 814.5 824.4 820.9 818.5 824.9 830.6 821.7 826.1 820.5 804.7 811. 273.2] | [413.1 406.7 406.2 408.7 404.1 407. 408.3 416.3 402.9 406.4 404. 407. 408.1 403.2 405.8 406.1 402.6 410.6 403.2 406.3 403.3 404.7 408.4 408.6 404. 407.6 407. 402.8 410. 404.2 407.5 403.4 404.1 401.8 409.7 412.7 404.2 403.7 408.5 401.9 406.7 402.8 404.9 405.8 405. 401.5 399.7 410.7 407.8 404.6 406.7 404.4 405.1 404.5 404.3 405.9 406.6 403.1 401.6 406.9 403.9 406.1 402.3 408.1 404.4 404. 399.9 408.4 403.9 407.9 402.2 402.5 403.5 404.4 403.5 407.2 401.8 405.1 404.5 406.6 402.2 403. 407.6 396.8 409.8 397.9 407.2 397. ] | [411.8 405.8 404. 403.3 399.1 397. 398. 398.1 398.8 403.2 405.2 401.6 406.4 403.4 403.9 398.6 402.4 401.6 400.5 405.2 395.7 403.5 398. 402.2 393. 405.3 409.4 402.4 406.7 399.5 404.6 401.7 399.1 402.1 400. 402.2 395.5 412.1 390.2 410. 397.5 405.6 398.4 405.9 399.8 401. 398.4 399.6 413.8 396.7 408.2 392.3 405.3 395.7 402.1 396. 402.2 397.4 399.5 389.5 408.6 400.8 408.7 399.3 404.7 399.6 401.9 398. 408.7 391.9 403. 396.6 401.9 395.4 397.9 398.8 405.7 402. 408.5 401.7 405.7 399.6 404.2 400. 392.9 399.3] | [243.9 257.4 265.5 260.9 266.9 264.5 268.3 256.4 264. 253.5 257.5 252.1 250.4 249.1 254.2 249.9 275.7 245.4 268. 252.4 265.5 252.1 265.4 250.8 259.7 261.4 255. 259.9 252.5 261.2 250.6 259.8 244.7 263.1 246.8 266.8 258.7 259.8 257.8 257.8 257.1 256.2 257.2 255.6 255.9 247.3 269. 250. 263.5 256.9 260.4 257. 255.2 255.5 253.8 255.5 255.5 269.1 254.2 262.2 254.8 255.2 251.9 257. 252.4 255.4 252.7 250.3 263.4 255. 265.3 255.9 257.5 256.5 255.8 256.3 251.5 263.1 261.7 256.2 252. 255.1 252.4 253. 248.3 257. 249. 269.3 258.3 265.2 255.2 265. 253. 251.3] | [194.7 99.2 102.7 103.9 97.1 118. 132.9 146.8 139.8 145.8 142.9 142.4 118.6 152.8 116. 126.2 112.6 123. 106.4 128.4 118.4 152.6 105.1 146.6 131.2 146.8 129.3 148.3 126.6 137.4 132.8 140.3 131.4 123.3 126.5 123. 126.2 123.7 139.6 130.4 152. 129.3 135.2 132.7 147.7 134.3 144.3 137.5 138.9 128.2 127.7 120.5 156.7 119.8 141. 120.2 142.4 118.3 142.7 126.1 150. 129.2 141.2 136.9 152.2 138.3 149.7 135.2 145.1 124.6 153.2 132.4 138.6 132. 137.1 127.8 139.8 139.4 139.9 145.2 151.1 147.3 147.6 147.2 131.7 150.3 128.3] | [14.1 9.5 10. 1.4 3. 4.9 14. 14. 1.3 3. 2.4 10.6 3.5 3.6 14. 6.7 10. 1.3 2.5 1.2 3.8 2.7 2. 9.5 10. 14. 2.2 4. 1.2 1.3 7. 9.5 14. 3. 10. 13.7 14.7 5.1 2.8 2.7 1. 1.2 13.4 1.3 3.6 7.3 3.7 4.4 3. 2.7 13.4 13.2 14. 14. 2. 10. 3. 10. 2.9 7.2 4. 10. 5.1 10. 7.2 1. 3.2 15. 1.3 15. 10.7 3.4 10. 3.3 9.5 2.5 10. 3. 15. 1.3 3.3 10.9 3.6 3.1 3. 1.2 2.7 13.7] | [ 3.3 10. 3. 10.6 2.7 13.2 13. 2. 6.3 5. 13.2 9.5 10.2 12.2 3. 13.2 10.2 1. 13.2 12.9 2. 15.2 2. 15.2 1. 2.5 13.2 5. 3. 12.9 13. 2. 13. 12.9 13.2 16. 1.3 13. 4. 10. 2.7 3.2 10. 5. 12.9 13.2 2.3 1. 4. 1.2 3.6 3.3 10. 2. 4.4 10.6 4. 5. 6.7 12.9 13. 10. 4.4 2. 3. 2.8 13. 13.2 10. 1. 1.3 4.8 13. 2.5 16. 6.7 13. 12.2 13.2 3. 13.2 13. 13.2 13.2 13. 12.7] | [ 1.2 10.7 1.2 3.3 3.3 3.6 9.5 97.4 2.6 2.4 6.3 15. 1.7 15.6 2. 12.2 1.2 2.4 10.9 4.6 2. 51.9 1.2 16. 15.6 1.3 6.7 2. 5.3 4. 1.2 4.6 5. 2. 2.8 15.8 9.5 10. 4.6 4. 6.7 1.3 1.8 10.7 2. 1.1 3.1 4.6 13.2 2.7 2. 1.1 2. 2.8 10. 2.4 1.1 13. 6.4 15.6 15. 2.7 1. 7.6 10.7 1.8 16. 2.6 15. 4. 13. 14.2 13. 4. 6.7 3.3 1.2 13. 13.2 15.6 2. 1.3 9.5 3.3 14.2 3.7 1.8 5. ] | [ 72.6 98. 9.5 97.6 66.4 59.3 70.3 83.5 13.1 3.2 11. 47.1 60. 15. 41.9 1.2 3. 65.6 73. 10. 82.3 1.2 77.9 61.2 76.2 10. 93. 1.3 13.2 9.5 82.2 12.2 93. 5. 76.8 13. 103.1 11.1 85.3 5.6 88.8 10. 54.7 2.5 53.9 68.7 12.2 1.3 97.7 9.5 91.4 3. 85.5 13. 78.2 10. 81.5 107.9 2.8 89.8 1. 76.5 9.5 2.6 4. 55.2 2. 1.3 71.1 5. 76. 10. 84.2 9.5 65.6 6.3 13. 16. 2.9 13.2 9.5 10. 62. 9.5 51.9 62.1 64.3 63.8 55.4 63.9 62.1 16. ] | [ 83.1 1.3 71.6 77.7 1. 96.8 94.5 84.6 53. 1.2 55.8 103.8 68.7 70.7 4.9 57.7 3.2 91.5 64.2 72.5 1. 40.9 83.9 77.8 94.1 1.4 54.2 3.3 68.5 68.4 83.2 7. 2.7 3. 4. 70.8 78.5 68.3 76.8 2.1 64.9 2.6 102.7 64.5 103.1 78.7 72.7 58.3 4.1 66.7 51. 85.8 69.3 1.6 3.4 6.3 81. 76.1 88.3 69.5 2. 10.6 4.3 1. 107.3 4. 66.2 2.7 1.2 1.4 3.2 48.7 73.9 9.5 4.2 2.7 1.1 80.2 52.3 84.5 2.4 73.8 1.1 5.3 1.3 1.1 1.3 2.5 88.2 68.9 97.1 79.4 87.9] | [ 9.5 2.7 115.4 93.9 101.7 114.1 108.8 88.5 46.2 2. 74.7 105.1 6.7 2.1 4. 78.6 76.3 68.8 2.5 1.2 2.8 2.7 116.2 105.4 95.8 5. 2.5 79. 62.2 107. 1.8 128.8 2. 1.1 4.2 3.3 82.9 113.9 62.3 58.9 73.4 1. 2.7 120.3 90.6 83.6 2.8 1.1 2.1 1.3 111.2 1.3 131.9 1.1 2.1 4.3 54.3 2.7 108.1 80.5 54.1 10. 9.5 4.5 2.7 94.9 52.6 1. 1.7 2.7 1.3 3.1 117. 7. 5.6 3.4 1.7 78.5 2. 118.2 91.9 91.6 106.4 89.1 3.2 2. 1.3 3.3 2.6 90.3 41.1 9.5 95.6 98.8] | [ 85.8 104.3 1.2 1.8 127.8 1.3 124.5 5.3 166.5 168. 9.5 128. 1. 188.2 101.5 179.7 217.3 94.9 3. 5. 57.8 85.7 128.3 2.7 2.8 9.5 3.5 177.1 117.1 127.6 3. 111.9 1.3 172. 104.4 172.6 6.6 65.5 1.8 120.7 2.5 121.8 1. 3.2 3. 7.4 213.2 211.6 185. 89.6 4.1 7.2 91.1 113.9 130. 158.7 2. 181.3 4. 121.4 1.1 60. 126.4 87.2 38.5 1. 4.1 2.6 160.7 2.7 6.6 4.1 1.8 91.2 91.8 157.5 2. 3.1 1.2 151.3 50.6 144.5 3.2 1.3 4. 1.1 90.7 8. 83.9 2.4]
3.001 | iveslevesque_muhc_mgh_NIST | https://osf.io/nk6ae/download/ | 20200203_iveslevesque_muhc_mgh_NIST/20200203_iveslevesque_muhc_mgh_NIST_Magnitude_T1map.nii.gz | Magnitude | Ives Levesque | McGill University Health Centre - Montreal General Hospital | Siemens | syngo MR E11 | 3 | NIST | 130 | 42 | 19 |  |  | T1_se_cor_TI | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 390 |  |  |  |  | [1903. 1943.1 1966. 1908.9 1903.9 1907.2 1878.4 1856.5 1841.4 1861.7 1902.3 1935.1 1945.2 1872.5 1842.2 1817. 1862.4 1894.3 1891.8 1853.3 1921. 1897.3 1869.5 1846.8 1845.5 1830.2 1890.6 1864.7 1889.6 1865. 1881.5 1884.2 1857.9 1919.1 1917.7 1938.3 1918.6 1892.1 1871.6 1922.3 1868.9 1842.3 1896.3 1822.8 1892.2 1899.8 1862.7 1857.7 1868.8 1861.1 1884.5 1876.7 1861.9 1889.1 1837.9 1852.3 1883.7 1882.3 1889. 1916.1 1850.6 1862.8 1910.6 1883.4 1866.4 1842.1 1876.2 1860. 1842.7 1833.9 1813.4 1860.2 1794.3 1896.8 1866.7 1902.8 1887.8 1872.6 1852.6 1849. 1872.9 1894.5 1889.2 1828.7 1862. 1871.4 1830.4 1902.6 1916.9 1899.1 1876.5 1893.5 1866.1] | [1362.3 1370.5 1336.2 1315.2 1315.3 1306.9 1332.2 1334.2 1317.1 1358. 1353.4 1353.8 1338.5 1289.3 1329.9 1336.6 1333.7 1321.9 1333.5 1332.4 1340.4 1349.2 1338.3 1315.1 1321.9 1342.6 1335.3 1355.4 1332.5 1320.5 1337.9 1362.3 1346.8 1346. 1327.8 1339.2 1313.4 1353.7 1358.4 1327.9 1299.8 1359.6 1345.5 1326.7 1339.5 1326.2 1330.6 1360. 1341.9 1340.3 1361.8 1341.9 1342.1 1323.1 1323.8 1326.8 1326.3 1323.5 1312.7 1334.7 1349.5 1363.8 1312.1 1295.8 1369.4 1341.3 1327.6 1339.8 1353.5 1323.6 1339.6 1331.1 1342.5 1365.7 1325.4 1338.9 1346.2 1352.1 1313.9 1310.4 1377.6 1340.3 1330.3 1327.4 1337.9 1331.5 1342. ] | [923.8 940.4 931.6 933.2 896.9 922.9 939.5 949. 946.2 945. 937.8 956.2 914.3 943.9 924.6 962.6 941.3 947.9 957.9 932.7 938. 948. 930.1 943.7 952.9 947.7 955.5 955.2 967. 935.4 937.2 924.2 932.4 943.3 946.4 944.8 948.3 933.2 960.2 942.5 946.3 931.2 920.9 937.4 944.4 954.6 932.4 960.2 964.7 951.8 929.7 952.5 934.2 943.1 913.1 966.1 937.1 935.7 936.8 944.1 946.1 940.4 896.8 912.9 914.4 966.2 953.5 939.1 972.1 950.9 888.5 943.2 944.5 927.2 914.2 920.7 913.3 911.1 896.4 943.1 945.3 950.6 933. 934.1 940.3 924.7] | [654. 641.9 645.8 669.9 656.6 670.9 641.2 643.3 662.3 650.4 624.1 655.1 624.6 679.8 642.2 672.5 649.8 658. 655.4 639.4 659.1 624.2 673.5 649.4 668.3 630.8 654.2 661.1 643.7 645.6 618.1 629.3 659.3 634.1 657.5 675.4 669.5 632.2 651.7 628.7 639.7 691.1 644.5 641.1 668.3 632.9 641.3 637.4 649.3 674.6 598.1 691.1 697.2 614.7 635.3 652.5 668.7 664.1 644.3 603.6 651.4 676.6 663.6 689.3 641.9 617.1 612.8 614.5 618.1 676.8 680.5 475.3 635.4 684. 684.8 689. 689.8 688.1 649.4 600.1 624.5 608.3 609. 603.6 602.5 627.7 633.9] | [441.5 444.1 447.8 439.5 461.8 466.1 468.8 470.2 465. 486.4 438.3 468.3 454.6 449.3 458.6 454.3 451.1 458.2 449.3 458.5 455.3 470.1 463.4 467. 459.3 454.6 483. 447.7 460.4 452.4 459.2 455.6 456.2 456. 463.1 453.1 461.6 478.4 448. 462.8 457.5 459.5 455.5 458.1 461.8 462. 455. 463.3 466.5 467. 443.7 457.7 447.6 459.8 461.3 460.9 447.3 457.7 449.7 487.8 446.6 456.5 457.1 449.9 452.2 453.4 463.2 467.6 465.4 484.9 446.5 448.7 459.2 463.1 464.5 452. 443.3 466.8 481.7 462.6 444.5 437.2 441.6 455.4 473.7 479.9 480.3] | [343.7 331.3 326.6 350. 334.6 328.2 346.3 349. 344.1 339.6 334.3 346.1 341.9 331.8 344.8 336.3 334.4 330.6 341.4 347.7 333. 348. 326.4 344.3 332. 338.9 346.6 345.6 333.1 348.4 335.6 346.8 330. 347.5 335.2 349.1 331.2 347.7 343.7 343.7 344. 338.5 346.4 332.3 346.7 335.8 347.8 332. 346.9 341. 346.9 345.2 325.9 343.5 330.4 340.1 329.3 339.1 342.1 342.4 335.3 343.8 331.4 342.2 335.2 340.4 341.5 337.1 334.5 337.4 349.7 339.2 335.3 336.8 346.8 331.9 332.3 343.7 345.3 345. 332.6 331.4 348. 334.5 340.4 340.7 329. 326.9 329.7 342.5 347.5 338.5 337.9 337.2] | [233.2 239.1 241.6 237.5 237. 243.1 240.2 245.4 246.7 246.7 246.8 240.7 238.4 242.4 237. 239.7 240.9 240. 238.3 239.9 247.6 241. 238. 247.5 241.2 242.3 248.9 246.9 243.6 240.7 247.6 239. 234.3 248.1 238.6 244.8 240.3 238.3 246.2 239.5 244.5 238.5 245.3 238.9 242.4 241.5 242.9 247. 247. 243.8 242.2 245.4 239.6 246.6 242. 245.9 240.8 238.1 241.2 242.9 242.7 240.6 238.3 239.5 243.3 246.5 243.4 244.6 248.3 242.4 243.1 242.9 243.4 241.2 238.9 241.8 242.9 241.7 238.6 241. 240.9 241.7 241.4 240.8 241.6 231.9] | [172.9 177.2 183. 180.9 177. 175.1 177.1 173.1 182.1 171. 170.1 176.5 180.1 177.5 171.6 169.1 183.1 170.3 174.9 178.7 178.4 175.3 176.9 179.5 170.9 175.7 178.8 172.4 182.4 172.3 177.8 179.4 172.8 174.2 179.5 177.3 175.8 181.3 177.8 178.5 179.2 172.3 180.6 173.4 181.2 174.5 179.1 177.8 181.2 174.4 180.7 181. 171.6 182.2 174.5 181.9 176.9 179.6 175. 179. 172.5 178.4 179.4 176. 174. 182.6 181.1 170.5 178. 176.1 176.7 173.7 174. 179.9 172.4 179.9 172.5 173.3 181.4 179.8 178.9 173.9 180. 182. 176. 173. 173.7] | [182.6 181.7 180.5 127.3 119.7 112.3 117.3 118.5 119.5 164.4 131.2 114.9 116.3 183.3 182.2 184.2 175.4 186.5 119. 115.5 183. 119.5 118. 111.7 122.2 179.7 121.3 182.1 115.2 180.6 119.2 124.5 178.1 178.6 120.2 124.3 119.1 123.8 177.7 120.5 121.7 118. 178.2 109.5 122.6 174. 119.2 119.6 123. 183.9 119.9 166.8 117.3 181.4 123.1 183.6 119.5 123.9 119.6 188.5 118.3 176.8 122. 116. 121.9 117.5 115.1 124.8 183.8 124.8 113.4 183.7 124.9 118. 160.9 172.9 114.9 193.7 120.7 113.2 116.8 119.9 121.1 113.1 186.9 124.8 121.8] | [111.7 99.3 92.3 86.1 113.5 116.2 99.2 72.6 72.4 85.4 84.5 3.1 108.3 101.8 67.9 96.2 105.7 78.3 89.4 112.9 66.6 115.1 88. 89.3 93.5 78.3 90.9 75.9 66.8 94.4 49.4 97.2 81.9 94.7 68.5 103.7 102.8 103.1 78.7 70.2 85.7 116. 96.7 80. 90.9 79.7 99.9 1. 77.5 85.2 80.5 88.8 116.1 68.4 102.9 77.3 103.9 107. 108.4 70.5 104.5 91.8 124.2 77.5 84.6 109.1 83.7 75.1 57. 70. 110.6 49.3 96.1 75.7 79.6 103.7 95.7 101.6 107.2 44.1 94.5 76.4 49.3 75.4 76.8 108.7 92.7] | [ 47.2 9.5 84.5 95.7 70.7 1. 75.3 86. 93.7 3. 100.3 67.7 1.3 61.6 58.9 75.7 9.5 61.3 110.5 70.5 1. 86.3 73.1 2.5 1. 97.4 4.9 68.6 92.9 4.6 83.1 1.2 2. 73.7 3. 9.5 84.9 3.1 98.3 77.8 1.2 79.2 1. 91.7 9. 89.7 1.2 72. 2. 76.1 1.2 69.9 9.5 89.3 2.1 69.9 64.3 59.1 6.1 66.7 90.3 2. 72.5 1.2 79.6 87.2 77.7 2. 89.2 1.2 87.8 10. 1.2 87.3 1.2 1.8 1.1 49.1 62.7 2.7 92.7 1.2 3. 86.9 80.1 81.1 58.1 1.4 42.7 3. 50.3 3.2] | [ 3. 1.8 1.2 5.6 74.3 9.5 46.4 96.2 2.6 4.8 104.6 1.8 103. 46.2 3.1 2. 72.8 1.6 8. 85.6 3. 101.6 2. 70.3 3.3 69.6 80.2 108.7 57.7 2.6 1.2 1.8 73.5 9.5 64.7 2. 97. 2.5 59.1 1.8 77.4 107. 91.3 2. 91.1 6.9 81.9 9.5 5. 62.3 52.8 4.3 81.8 2.5 129.4 2.8 3.4 3.3 5. 55.6 1.2 74. 2.7 1.2 58.3 3.4 97.6 3.4 2.4 48.7 66.5 60.5 176.1 104.1 1.9 1.2 88.3 73.7 5. 2.7 1.2 69.1 88.5 86.2 1.3 143.7 104.7 89.2 84.5 123.4 161.2 76.2 71.5 5. 91.2] | [ 9.6 2. 2. 90.5 97.2 107.1 3.3 2.7 10.2 2.8 1.3 5. 97.9 1.2 9.5 1.6 7. 3. 105.6 110.9 3.1 4.1 4.2 4. 1. 64.5 2. 73.4 2.1 2.4 79.3 3. 77.4 1.3 138.5 5. 5. 2. 3. 1.2 1. 99.6 1.9 73.6 1.2 76.2 76.5 7. 89.6 1.3 2.1 84.2 2.5 2.7 100.9 4. 3.6 90.9 1.1 6.3 99.7 4.5 7.7 1.2 4. 1.8 69.7 2.5 82.1 3.4 1. 1.2 3.4 1.2 3.3 59.3 3.3 122.8 2.1 1.6 1.2 2. 103.5 98.9 4.6 3. 3.2 57.9 62.6 104.3 90.2 3. 67.8 1.8] | [ 97.6 1.3 10. 107.2 132.8 9.6 72.2 153.8 77.7 59.4 6.1 3. 116.5 71.3 1.3 1.2 10. 85. 1.3 1.6 2.4 78.8 83. 2.6 107.6 57.7 3.3 8. 75.1 3.4 3.4 1.2 167.8 1. 3.1 10. 1. 72.3 137.3 5.9 120.6 10. 144.8 4.1 176.6 2.7 156.7 2.8 154.4 9.5 129.5 5. 3.7 166.1 1.2 154.2 4.2 170.6 4. 123.7 3.2 136.8 1.9 1.5 192. 1. 67.5 6. 2. 5.1 157.6 4.1 78.6 1.2 162. 2. 1.9 118.6 1. 2.5 5. 76.1 81.3 9.5 4. 9.5 142.3 104.9 2. 2.7 1. 1.1]
4.001 | mrel_usc_NIST | https://osf.io/t5y28/download/ | 20200204_mrel_usc_NIST/20200204_mrel_usc_GE3T_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Nam Lee | Keck Medical Center of University of Southern California GE 1 | GE | Signa HDxt | 3 | NIST | 130 | 46 | 20.9 |  |  | 2D Spin Echo, IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 244.141 |  |  |  |  | [1887.3 1882.5 1834.1 1897.5 1888. 1857.8 1871.3 1896. 1867.8 1891.7 1852.9 1869.5 1916.7 1813.8 1844. 1910. 1896.9 1870.3 1846.7 1881.3 1864.5 1834.3 1942.6 1900.3 1888.2 1859.3 1872.8 1871.6 1834.2 1851.5 1915.1 1867.3 1824.9 1902.7 1898.4 1899.7 1833.7 1860.6 1878.9 1876.2 1885.4 1915.1 1891.5 1837. 1867. 1832.7 1875.4 1829.3 1878.2 1893.3 1920.6 1888.2 1857. 1888.2 1855.2 1858. 1831.3 1879.7 1876.1 1897.6 1848.6 1872.3 1871.3 1852.2 1862.9 1857.2 1815.9 1800.2 1803.5 1844.1 1861.6 1814.2 1840.6 1874.1 1859.2 1899.5 1813.8 1838.5 1834.8 1857.8 1878.3 1861.3 1863.5 1828.8 1862.1 1832.4 1806.9 1840.8 1829.3 1852.4 1927.5 1906.5 1865.2 1851.4 1832.2] | [1331.5 1334.5 1338.7 1344. 1330.8 1340.5 1338.5 1342.1 1330.7 1337.6 1353.8 1343.6 1302.7 1322.6 1342.2 1304.5 1326. 1342.2 1338.1 1339.6 1330.9 1301.1 1337.4 1309. 1330.7 1315.3 1327.3 1323.7 1335. 1349.2 1354.6 1315. 1329.3 1288.2 1315.5 1316.8 1337.1 1323.6 1331.5 1339.4 1369.5 1357.5 1351.7 1295.8 1326.5 1318.8 1331. 1337.6 1347.2 1322.6 1361.6 1362.5 1308.2 1326.7 1325.4 1328.8 1332.3 1351.7 1328.4 1300. 1345.9 1315.8 1261.8 1384. 1321.5 1314.3 1333.8 1341.1 1313.5 1305.9 1354. 1335. 1345.9 1329.2 1316. 1299.2 1310.9 1320.2 1349.6 1337.2 1361.9 1331.1 1327.6 1335.2 1361.4] | [943.4 954.3 931.4 925. 934.2 932.8 927.8 938.7 939.3 945.7 942.4 941.1 803.7 933.7 939.7 929.8 933.4 944.7 933.3 939.1 948.8 954.7 949.5 950.9 939.1 936.9 946.8 942.6 935.5 949. 942.3 947.9 938.4 931.1 937.6 943.1 945.4 939. 946.9 957.7 945.8 931.6 926. 938.1 949.4 943.6 927.7 939.3 985.3 943.1 936.9 926.5 923.4 935.5 945.5 944.1 939. 856.6 936.4 945. 934. 925.2 924.4 923.8 925.1 921.9 999.2 949.8 948.8 938.5 927.3 930.1 944.7 817.7 941.9 940.3 935.2 958.7] | [668.8 706.9 654.3 650.8 685. 696. 677.8 682.5 679.3 671.2 665.4 664. 668.6 668.4 676.4 673.5 674.8 689. 668.6 655.1 678.3 668.4 672.2 672.4 672.9 671.9 676.9 661. 668.5 682.4 673.8 679.5 674.3 670.2 670.2 670.1 671. 677. 691.8 674. 676.2 669.5 666.8 671. 675.4 668.3 671.4 687.1 675. 676.8 677.3 676.8 678.2 679.2 664.8 671.4 672. 660.2 668.4 672.2 669.7 671.8 677.6 684.4 677.4 678.9 660.2 665.6 672.5 673.8 661.6 671.7 680.5 673.9 677.5 678. 678.9] | [475.4 483.3 483.7 481.8 485.6 482.1 479.3 482. 482.2 475.4 478.5 479.7 479.6 480.4 470.6 481.5 479.9 482. 476.6 477.9 483.9 482.4 481.3 475. 481.8 476.9 476.1 476.4 476.2 482.1 485. 479.6 477.5 478.1 477.8 476.5 480.4 482. 480.3 485.9 483.6 479.3 479. 482.4 479.3 476.1 478. 483.2 480.4 483.6 484.3 477.2 475.5 488.9 480.5 480.5 481.3 482.5 481.1 483. 483.9 479.3 477.9 484.8 479.1 479.8 482.9 481.1 478.9 485.3 486.4 471.4 480.4 483. 481.4 481.3 477.4 477.4 475. 471.9 473.2] | [351.5 344. 341.9 338. 338.5 343.8 343.4 342.8 342.7 343.2 347.7 338.2 340.7 340.7 345.8 347.2 343. 344. 346.2 338.8 338.1 342.7 347.5 338.8 343.3 343.1 343.7 343.5 344.8 340.8 340.1 344.5 336.7 348.8 340.2 347.4 341.6 344.5 341. 342. 342.4 345.7 343.4 340. 344.6 342.3 347.3 343.7 343.8 338.3 339.6 339.8 350.2 348.7 341.8 341.3 339.3 341.8 343.8 342.4 338.3 344.3 341.7 344.4 344.2 339.7 341.5 344.9 342.6 340.6 341.4 338.8 343.8 341.7 344.5 343.6 348.1 344.8 347. 344.9 346. 341.3 349.4 338.9 343.9 340.5 343.6 338.5 342.6 338.8 344.5 337. 340.6 342.3 345.4 346.9] | [241.4 237.1 239.9 243. 241.4 241.9 239.3 239.5 244.7 245.2 245. 247. 241.5 237. 243.8 243.3 240.5 240.7 240.1 238.5 243.7 244. 238.6 238.5 235.8 240.7 239.9 242.1 240.6 239.6 241.7 240.1 240.8 238.9 237.9 239. 238.2 239.4 238.6 239.1 240.5 240.6 239. 240.4 241.1 239.1 240. 241. 242.4 242.4 241.1 241. 239.5 242.7 237.2 236.8 239.5 241. 240.9 241. 240.3 243.1 243.9 246.1 239.8 242.4 241.5 242.3 241.2 242.1 239.3 241.1 240.8 244.3 239. 240.7 241.6 241.2 241.3 241.8 238.4 237.1 241. 239.9 240.7 241.7] | [167. 169.7 172.5 171. 166.8 168. 173.8 172.5 173.6 172.8 173.1 171.4 154.5 161.1 178.5 173.7 175.3 276.8 176.4 177. 174.5 172.7 154.3 179.1 171.1 267.4 173.2 172.2 172.3 275.5 174.1 175.4 167.2 182.8 172. 175.1 172.3 273.2 172.4 175.6 173.3 173.3 170.5 179.3 170.4 173.4 172.7 174.5 172. 171.7 170.7 171.5 167.6 177.2 171. 178.8 176.4 176.3 173.2 171.9 174.4 173.3 165.8 329.2 170.3 176.8 173.2 174. 171.8 173.3 175.2 174.9 174.6 171.9 174.8 174.3 173.6 174.6 173.3 175.3 174.1 171.4 170.3 175.1 173.6 171.3] | [ 61.5 112.4 117. 114.3 117.3 110.7 126. 123.7 120.7 120.2 121.8 120.1 116.5 93.6 156.3 112.1 124. 120.7 120.9 119.6 119.3 121.9 110.3 114.9 118. 149.8 123.8 121.2 119.7 117.8 121.7 115.5 115.5 126.6 122.9 120.3 118.1 121.1 117.5 119.4 114.9 90.9 110.4 145.8 118.9 125.3 122.8 122.5 114.7 113.6 109.2 92.7 102.2 112.6 107.9 113.7 115.2 117.3 114.6 116.9 107.1 96.7 96.9 113.8 118.5 126.6 122.6 119.7 118.6 116.3 83.8 118.2 101.7 117.9 116.7 117.6 112.8 99.4 148.7 159. 150.2] | [ 69.5 89.5 90.8 88. 91.1 66. 90.3 87.6 89.5 84.9 84.8 74.4 86.8 82.1 89.1 81.1 82.4 85.2 87.5 83.7 81.7 103.5 81.6 78.1 79.2 81.8 76.7 84.2 88.9 80.3 63. 99.3 80.9 89. 88.3 87.2 57.6 64.1 81.2 84.2 88. 76.8 92.4 87.8 75.2 75.5 90.7 88.5 86.6 83.2 80.7 84.6 80.8 70.8 85. 89.6 83.6 85.5 83.2 98. 79.5 84.3 77.5 87.7 93.2 91.7 78.4 75.8 83.9 88.2 84.1 8. 77. 86.7 75.9 80.3 82. 96.6 89.1 90.7 91.8 78.8] | [55. 63.8 62.1 66.7 66.5 1.3 41.6 71. 1.8 1.3 66.9 66.5 2.5 1.2 75. 55.4 3.3 66.8 68.2 50.7 2. 5. 51.5 67.5 71.9 1. 53.5 60.4 74.9 2.8 3.3 48.3 49.5 72.6 65.2 2.7 62.4 63.2 77.2 57.6 69.1 76.3 53. 68.9 66.2 53.2 41.6 62.3 64.3 51.6 72.7 63. 51.6 64.8 55.6 4.6 47.9 2. 68.2 1.8 1.2 41.6 72.8 67.6 64.9 7.2 39.6 1.8 67. 69.3 9.5 1.8 53. 60.1 3.2 46.8 3. 2. 65.1 2.1 2. 62.6 58.9 1.6 2. 2.5 6.4 1.2 3.2 1.8] | [ 66.3 10. 63. 74.6 86.5 2.5 64.4 74.9 3.3 1.8 1.8 110.4 1.4 52.7 5.3 6.1 1.1 2. 1.2 9.5 82.3 4. 64.3 3.3 3.1 3.2 1.8 3.1 2.7 72.3 1.8 4. 4. 2.4 1.9 5. 2.8 52.8 2.1 71.3 1.3 61.3 65.1 76.7 87.1 76.5 43.5 54.8 3.6 79.9 72.8 68. 3.1 74.2 86.4 57.1 5. 55.9 67.1 5. 71.6 44.3 1.4 84.9 75. 54.4 58.6 56.1 1.2 9. 51.5 55.7 5. 60.1 65.8 78.8 104.4 10. 2. 64.3 1.1 1.4 65.7 57.7 60.7 93.7 64.6 75.9 64.6] | [ 1. 69.1 50.9 2.5 1.9 94.2 6.1 3.8 2. 1.8 5. 94.2 113.4 2.7 73.3 1.8 2.5 1.3 7.1 7.6 76. 5. 84.9 5.3 70.1 3.1 69.9 75.4 79.1 2.5 2.5 2. 10. 90.3 1.2 127.7 1.8 68.3 3.3 71.5 91.2 9. 1.8 101.6 3.3 103. 2.8 116.1 89.9 97. 2.8 7. 2.4 120.5 2.8 68.7 1.6 84.8 1.2 110.9 1.4 1.2 113.3 2.1 56.6 51.6 8.8 49.7 83.9 3.6 3.4 9.6 62.5 4. 55.3 1.2 71.9 2.1 5. 3. 91.2 105. 91.5 95.9 2.1 3. 8.6] | [ 60.7 58.7 47.6 100.5 40. 2. 1.6 3.2 56.5 65.3 75.2 2. 4.2 61.4 3. 5.5 1.1 3.6 72.1 121. 65.7 59.7 2.5 3.1 58.6 2. 3. 1.2 10.7 54.6 65.3 56. 1.3 54.8 47.9 1.2 8.8 1.8 1.6 75. 3.1 53. 2.4 2. 1.6 112.9 72.3 66.9 55.8 1.1 1.2 68.1 3.6 1.2 50.6 2. 3.3 66.7 61.1 1.6 1. 9.5 8.6 106.9 9.5 2. 68.3 57.5 1.2 68.2 148.7 2. 3.3 8.4 129.1 143. 61.6 2.8 4. 43.5 114.4 81.9 6. 1.2 1.3 1.2 1.3 56.2 1.4 1.8 1.3]
4.002 | mrel_usc_NIST | https://osf.io/t5y28/download/ | 20200204_mrel_usc_NIST/20200204_mrel_usc_GE3T_MR2_NIST_Magnitude_T1map.nii.gz | Magnitude | Nam Lee | Keck Medical Center of University of Southern California GE 2 | GE | Signa HDxt | 3 | NIST | 130 | 46 | 20.9 |  |  | 2D Spin Echo, IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 244.141 |  |  |  |  | [1851.2 1849.3 1879.3 1822.4 1801.9 1818.3 1814.7 1872.7 1891.6 1849.7 1852.5 1834.9 1842.1 1793.9 1853.9 1906.7 1909.5 1854. 1853.7 1839.4 1858.2 1833.6 1857.1 1841.9 1835.3 1833. 1863.3 1888.8 1882.7 1858.5 1877.4 1795.2 1779.4 1855.6 1845.8 1816.5 1835.2 1886.1 1905. 1844.3 1849.9 1873.6 1831.3 1774.4 1845. 1903.9 1883.6 1879.7 1885.3 1854.2 1842.3 1850. 1811.5 1856. 1826.2 1822.5 1905.9 1864.2 1835.1 1851.7 1821.2 1881.7 1851.9 1872.2 1860.3 1821.2 1832.3 1815.2 1831.3 1844.9 1841.2 1878.1 1893.9 1861.2 1855.5 1839.3 1811.3 1818.3 1856.9 1848.3 1858.3 1876.6 1812. 1817.4 1859.2 1860.7 1832.3 1880.1 1879.1 1838.1 1826.1 1834.3 1909. 1880.6 1899.9] | [1314.8 1311.9 1310.3 1316.1 1323.2 1310.9 1307.7 1294.6 1307.9 1305.7 1327.9 1315.8 1299.4 1299.9 1307.6 1306.4 1316.6 1300.4 1303.2 1319.9 1311. 1309. 1315.1 1306. 1295.7 1295.3 1308.7 1320.9 1330.4 1328.3 1316.1 1310.2 1309.5 1327.9 1327.8 1327.9 1320.9 1297.2 1306.9 1302. 1319.4 1304.4 1308.1 1314. 1304.3 1305.1 1318.8 1305. 1330.5 1313. 1317.8 1304. 1306.7 1311.6 1323.7 1320. 1320.8 1309.5 1317.9 1318.3 1313.9 1317.1 1315.1 1310.1 1312. 1319.8 1313.3 1314.7 1328.6 1304.9 1312.5 1306.2 1295.7 1314.2 1319.5 1318.6 1296. 1307.4 1297.2 1307.3 1309.7 1313.8 1299.9 1321.3 1321.3] | [935.4 927.4 926.9 936.6 915.9 923.7 929.8 923.6 923.2 932.3 931.9 925.2 927.1 925.5 933.7 934.9 931.3 924.2 931.3 932.9 924.1 934.6 932.3 924.3 919. 927.8 931.8 930. 914.3 920.3 921. 926.4 933. 933. 932.6 929.1 925.2 933.5 929.9 923.7 920. 922.4 925.1 922.7 926. 936.9 934.7 920.9 930.7 938.6 929. 929.3 931.2 924.7 927.6 927.6 927.4 929.4 930.6 927.1 924.3 923.6 919.7 932.2 933.7 931.7 931.7 933.7 933.9 926.9 920.1 918.8 923.9 928.2 925.7 933.4 930.4 929.6 924.7 914.5 919.2 928.4 932.6 922.1 934.4] | [680.8 678. 667.9 669.2 666. 665.6 666.2 665.4 667.6 671.3 672.4 674.9 675.4 669.4 672. 676.5 675. 675.9 678.5 675.4 669.8 674.6 674. 670.2 675.7 680.5 672.2 664.8 662.8 665.5 682. 678.4 674.1 668.1 672.6 679.6 671.9 669.4 669.6 673.1 680.9 676.3 670.3 663.9 671.7 675.3 672.3 671.9 667.6 670.3 672.8 662.2 657.3 661.9 670.5 669. 672. 673.9 672.3 671.9 673.5 671.3 669.8 674.8 667.1 668.7 670.8 671.2 672.2 674.3 673.2 675.4 676.5 668.3 668.8 669.9 666.6 668.6 672.6 671.8 675.1 673.6 668.9 671.5 671. ] | [484.2 480.1 485.4 485.4 481.7 479.9 480.1 479.3 481.8 480.1 478.5 482.7 484. 480.1 479.9 476. 479.3 482.4 477.1 479.3 477.1 478.9 483.3 485.5 481.9 484.1 481.2 484.7 477. 479.2 480.8 479.4 480.7 478.3 478. 485.5 485.1 484.3 480.6 479.2 483.1 481.3 479.4 475.6 475.5 479.4 482.1 479.1 477.1 472.1 473.8 479.9 480.2 483.1 485.3 484.3 479.5 481.1 482.2 481.8 480.7 482.9 481.7 477.8 481.3 484.1 480.7 485.7 473.8 480.7 483. 482.9 479.2 481.6 481.5 479.1 487.6 483.4 487.5 482.5 480. 483.8 484.6 477. 476.6 478.4] | [338.9 335. 341.5 338.6 340.6 341.7 342.4 340.7 334.1 338.8 338.7 340.1 339.7 336.9 339.5 338.4 335.4 338.7 339. 340.1 341.2 343.1 340.4 343.7 336.9 337.7 337.2 338.1 340.3 339.7 336.9 340.3 339.2 340.3 337.2 342.6 333.3 339.4 340. 343.2 343.4 340.1 344.1 340. 342.5 338.4 344.7 334.6 343.4 345.6 341.7 340.1 338.6 343. 335.9 342.1 339.3 346.7 337.9 343.8 340.4 340.4 339.1 342.3 339.8 342.4 338.6 344.4 341.7 338.6 337.4 339.1 340.4 341.1 337.1 341.7 344.3 338.4 339.8 336.9 340.9 341.3 342.4 341.2 343.7 335.2 341.2 341.4 342.8 348.3 340.4] | [238.8 237.7 237.4 236.8 236.9 235.7 241.2 239.2 239.8 238.6 239.1 234.9 238.4 238.4 239. 237. 237.3 238.9 237.4 235.6 238.2 239.4 239.3 236.8 236.6 237.5 239.6 237. 238.9 236. 239.3 235.4 237.7 235.4 238.8 237.8 237.9 235.3 239.2 237.9 238.1 241. 238.3 244.3 239.4 243. 238. 241. 239.4 238.9 237.2 242.2 235.9 238.5 236.6 240.7 237. 241.1 239.3 244. 239. 240.5 241.4 238.9 239.7 238.8 240.8 236.7 241.4 239. 238.1 237.1 237.5 235.6 236.1 239.3 242.1 240. 239.4 240.2 240.4 243.1 238.4 236.5 234.2] | [173.5 174.4 173.6 178.5 171.9 171.1 170.3 172.7 173.3 170.8 172.3 173.7 172.1 174.5 173.6 172.7 174.4 173. 172.1 177.3 170.4 172.7 170.2 171.2 171.4 171.6 171.1 173.1 173.4 169.1 176.5 172. 176.7 173.6 171.5 171.3 172. 171.5 174. 171.7 177. 172.1 289.6 175.2 173.6 174.1 173.9 172.5 173.8 172.4 175.4 170.5 176. 173.2 172.4 172.9 175.6 175.6 174.6 174.1 172. 177.9 177.8 279.5 173.5 174. 172.5 173.4 173.2 287.3 174.5 175.5 174. 175.9 174.1 172.8 175.3 173.6 170.4 173.1 172.2 173.9 175.3 173.9 175.1] | [106.7 109.9 114.2 117.8 123.1 123.7 124.4 124. 123.9 119. 116.3 116.8 112.2 110.2 111.3 113. 118.6 125.7 108.6 118.2 117.6 121. 121.1 119.7 119.5 120.8 118.1 121.8 115. 120.3 115.7 111.8 107.6 110.6 112.7 116.4 118.8 117.1 116.4 115.8 113.8 135. 118.2 119.5 120.2 115.6 114.4 118.2 112.8 119.2 120.7 117.3 115.9 114.5 112.1 114. 114. 116.5 122.2 118.5 117.1 115.4 118.8 121.5 122.7 119.3 118.6 120.1 117.4 121.1 115.6 116. 111.9 112.2 112.3 116.4 119.7 115.1 116.5 119.8 133.8 137.3 123. 117.5 116.4 111.7] | [ 3.2 63.4 3. 4.1 48.1 78.1 1.3 85.5 90.9 84.6 84.3 93. 93.9 85.5 85.3 70.8 72.3 73.9 77. 77.5 68.1 80.9 75.9 73.3 78.7 86.8 83. 81.5 83.5 85.7 86.2 84.3 81.2 64.6 78.4 67.2 62.9 64.2 62.8 62.9 75.4 88. 78.7 68.9 75.5 71.8 74.8 66.3 76.1 81.4 80. 78.6 76.7 69.2 67.8 66.3 62. 61.8 80.1 79.8 81. 83.2 63.3 72. 69.3 74.4 69. 65.9 76.4 86.5 80.8 64.6 77.9 73.7 65.3 65.3 57.2 2. 69. 85.9 71.9 64.9 71.8 78.8 63.4] | [ 6. 5. 71.8 73.7 85.8 89.1 83. 69.1 77.8 85.3 1.3 1.8 3. 3.3 9.1 77. 50.2 76.1 75.3 80.9 88.6 87.4 1.8 2. 1.8 72. 1.3 49.4 2. 2. 5. 5. 3.3 7.2 1.1 68.3 62.6 70.6 4. 55.3 63.6 46.7 56.5 5. 4.6 59.2 60.7 66.2 63.8 3.1 9.5 3. 4.9 2.1 2. 67.9 1.6 4.1 55.9 69. 78.3 58. 1.3 10.6 71.1 52.8 2.5 55.7 9.5 8.3 69.6 68.6 60.8 56.4 64.3 1.3 72.2 67.8 64.7 61.6 57.8 64.6 3. 1.1 86.8 80. 3.5 2.4] | [ 1.2 3. 3. 3.3 49.5 2.7 4.7 7. 6.3 8.8 1.3 2.4 62. 2. 68.6 72.1 62.1 45.9 1.3 3.3 2. 55.9 3.7 59. 65.6 98. 6.4 2. 2. 4.2 2.7 54.2 59.4 52.8 64.7 47.2 3.3 38.8 2. 1. 2.8 9.5 96.4 3.6 78.7 5. 55.7 59.9 6.3 61.2 66.9 65.1 3. 89.8 8.9 58.2 1. 3. 1.2 70. 9.5 7. 3.3 1.1 1.3 47.2 4. 2.4 10. 58.5 1. 2.4 75.4 81.4 2.5 2.1 6.6 1. 4.9 3.3 2.7 1. 9.5 1.5 73.9 43.3 1. 2.3 3.1] | [ 4.2 1.3 42.7 42.6 1.8 7.9 6.7 104.3 10. 8.3 79. 119.3 49.2 63.2 56.9 76.3 120.6 91.8 1. 57.2 65.6 2.4 100.1 67.3 2.3 10. 105. 3.6 2.8 3.7 3.4 7. 80.9 95.3 9.1 6.3 46.7 121.7 6.2 1.4 58.7 50.3 1.3 104.1 98. 82.2 2.1 137.1 57.2 2.8 4.7 10. 4.7 15.2 70.5 1.8 99.7 2.3 3.4 4. 8.3 5.6 2.6 92.4 92.8 3.9 109.6 1. 110.8 3. 3.3 61.6 78.2 10. 3.2 9.5 7.3 4.6 6.7 72.8 2.1 2.8 9.5 71.8 2.1 105.9 3.1 10. 9.5 63.4] | [ 10.7 3.6 58.9 121. 96.1 129.9 89.3 83.5 1.1 72. 3.9 1.2 9.2 2.4 7.5 44.9 4.2 3.5 1.2 77.6 1.2 143.6 70.3 122.8 4. 46.7 59.2 170.8 7.7 90.1 4.2 5.9 3.4 58.8 42.7 16.9 8. 55.9 173.5 2.7 100.1 3.1 174.4 100.6 139.2 57.4 139.9 1.1 174.1 3.2 70.9 1. 51.4 7. 2.7 2.7 76.3 1.2 67.1 203.9 1.1 189.2 55.7 143. 66.7 162.2 4. 50.5 60. 5. 5.6 3.3 2. 1.2 1.8 5.2 1. 3.3 96.1 4. 158.7 164.9 188.8 64.3 181.8 61.5 75.3 1.8 4. 6.2 5.3 53.2]
4.003 | mrel_usc_NIST | https://osf.io/t5y28/download/ | 20200204_mrel_usc_NIST/20200209_mrel_usc_GE3T_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Nam Lee | Keck Medical Center of University of Southern California GE 1 | GE | Signa HDxt | 3 | NIST | 130 | 46 | 20.9 |  |  | 2D Spin Echo, IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 244.141 |  |  |  |  | [1696.5 1761. 1776. 1747.1 1734.6 1698.1 1736. 1731.6 1694. 1734.7 1781.7 1787.7 1694.1 1744.8 1800.5 1804.7 1781.7 1720.1 1689.5 1743.5 1754. 1705.6 1722.6 1761.5 1776.8 1718.6 1745.4 1735.1 1684.7 1710.5 1745.2 1746.8 1722.4 1737.5 1707.4 1724.3 1713.7 1740.9 1784.2 1738.2 1706.5 1731.2 1737.9 1781.9 1755.4 1702.1 1726.8 1784.5 1768.7 1764.5 1742. 1717.8 1735.1 1749. 1775.6 1778. 1755.7 1708.6 1703.9 1739.1 1752. 1738.8 1734.4 1733.1 1766.4 1734.1 1748.1 1724.2 1738.9 1779.8 1786.7 1748.9 1723.6 1748.9 1741.3 1749.7 1781.8 1770.1 1755.4 1741.7 1705.1 1729. 1726.5 1745.8 1763. 1778.4 1741.3 1708.9 1727.7 1772.8 1734.4 1713.3 1812.9 1734. 1711.3 1771.3 1770. ] | [1242.7 1302.9 1302.6 1279.5 1290.1 1279.8 1282.9 1262.8 1261.8 1279.4 1265.3 1267.9 1269.6 1281.4 1264.2 1308.6 1285.4 1255.3 1256.6 1277.2 1279.2 1259. 1265.4 1262.3 1263.8 1274.1 1306.2 1305.6 1268.1 1277. 1279.6 1285.7 1287.5 1253.9 1287.1 1271.9 1277.4 1289.2 1293.8 1296.2 1273.9 1274.8 1288.9 1264.1 1292.2 1286.3 1268.8 1272.5 1287.5 1293.6 1290.7 1243.4 1234.2 1274.7 1275.9 1267.7 1283.4 1286.7 1272.1 1262.1 1273. 1282.8 1285.8 1289.6 1286. 1282.4 1277. 1298.3 1300.2 1276.8 1264.3 1273.4 1299.3 1289.1 1309.1 1264.4 1278.1 1301.4 1271.4 1281.8 1280.8 1261.8 1278.2 1275.2 1280.7 1262.3 1265.3 1287.2] | [927.9 930.1 929.8 937.1 930.7 929.4 934.6 930.4 923. 954.9 931.5 899.2 947.2 964.7 934.9 928.9 939.1 919.2 936.5 920.7 918.5 951.7 943.9 941.1 931.1 918.1 926.4 927.6 930. 932.8 936.3 920.7 905.7 913.9 918. 932.1 926.3 930.8 932.3 932. 931.8 924.6 911.3 911.7 914.5 931.6 919.3 930.8 920.4 939.3 953. 942.1 911.7 898.9 919.7 927.8 919.3 954. 933. 944.1 918.8 904.9 924.2 929.4 933.6 939.9 923.1 941.5 927.2 931.4 916.6 928.8 946.1 958.5 929.3 926.3 923.4 930.7 928.6 913.7 921.9 916.7 938.8 925.2 923.5 925. 954.7 947.2] | [673.3 669.8 668.1 669.2 675.5 673.3 665.3 674.4 672.7 669.2 671.6 675.3 677.4 678.7 676.9 677.4 669.1 667.8 674.8 672.8 671.3 672.4 675.4 663. 666.9 664.9 666.3 668.9 677.2 675.5 674. 671.7 665.1 658.7 665.1 668.5 673.9 672.2 670.7 669.2 674.4 679.3 669.1 671.3 675.3 677.7 674.9 679.2 669. 669. 666.3 675.6 674.4 668.8 672.6 672.8 669.2 680. 677.8 681.6 674.1 679.9 673.6 666.3 674.4 675.5 673. 678.2 679.4 676.7 671.6 682.2 676.5 677.8 675.6 674.7 678.4 678.7 676.2 665.3 670.6 670.5 680.3 676.3 674.2 672.3 671.6 675. ] | [467.4 495.4 486.6 481.5 482.8 480. 466.9 483.5 481.2 480.1 483.5 485.9 484.2 483.2 484.9 483. 479.6 488.2 487.7 481.6 480.4 477.4 480.2 488.1 468.8 477.1 476.9 483.4 482.5 484.9 485.9 484.4 483.4 481.2 479.9 482.3 477.9 480.5 480.5 484. 483.5 482.2 484.1 476. 487.4 479.7 473.9 475.9 477.4 481.9 484.5 483.7 488.5 480.7 487.5 479.4 480.8 482.2 483.5 478.5 480.4 484.4 490.5 484.7 482.7 485.4 483.4 479.7 486.1 482.9 482.3 486.1 486.1 480.2 479.2 480.8 482.2 482.6 478.7 481.8 480. 484.9 478.5 481.5 481.7] | [324.7 340. 343. 347.8 339.7 348.7 351.9 344.3 341.5 342.1 342.2 348.3 349. 345.5 343.6 346.4 352.4 350.8 340.7 339. 340.3 327.2 346.7 340.3 345.2 342.1 342.4 345. 348.6 348.8 343.9 346. 359.2 342.2 341.8 339.3 341.5 337.4 337.2 348.6 345.6 342.7 346.3 361.3 345.3 344.6 339.6 345.2 343.4 341.2 345.9 343.7 342. 341.2 345.2 345.5 347.1 342.4 340.1 340.5 342.6 345.7 346.7 344. 338.7 329.2 338.4 337.5 342.4 345.8 345.2 344.8 344.3 344.1 345.3 344.9 347.8 341.1 342.6 349.4 348.8 345. 341.5 344.1 342.3 346.6 341.8 345. 344.8 343.4 343.4 343.1 350. 347.1 341.1 338.7 338.9] | [240.6 246.7 246.6 244.3 248.4 248.8 242.7 246.3 249.4 244.2 240.2 245.6 248. 246.8 245.6 243.2 248.6 246.9 243.8 244.6 247.4 246.7 246. 245.4 244.2 244.5 245.6 243.3 241. 244.7 245.1 244.6 243.8 248.6 246.4 245.8 243.6 243.6 242.2 244.3 242.5 244.3 244.9 249.8 245.9 249.3 246.9 244. 243.7 247.2 246.6 249. 245.1 243.6 244.3 244.9 246.2 245. 244.8 247.2 249.3 252.7 247.9 246.4 245.1 237.9 244.6 247.4 245.5 246.2 248.6 247.7 246.7 248.1 241.6 243.7 244.9 241.5 242.2 245.1 246.7 250.8 242.4 245.2 249.6 251.4 249.5 245.7] | [174.8 175.9 178.1 180.4 178.3 178.3 179.1 294.1 298.9 176.7 180.6 177.8 182.1 175.9 178. 180.5 179.5 179.1 174.8 296.8 177.6 179.9 175.9 295.8 297.2 294.7 181. 179.5 178.7 178.7 295.4 180.3 179.7 177.5 303.6 290.4 298.3 180.2 306. 176.5 177.3 178.5 181.5 177. 299. 291. 300.3 300.7 304. 292.1 301.2 176.9 294.8 179.4 300.1 182.2 181.4 181.7 179.7 298.2 302.4 298.8 291.1 300.4 182.4 185.5 182.3 297.3 178.3 298.6 178.6 307.1 296. 316.5 296.2 176.6 181.5 179.3 299.6 296.5 180.7 177.1 291.4 183.2 182.1 178.5 180.5 178.1] | [141.5 147.4 125.8 129.8 146.7 147.1 124.8 115.1 141.3 119.7 116.3 114.7 138.6 141.2 141.1 146.6 142.1 144.5 148.7 139.9 144.7 123.4 144.7 140.9 145. 118.1 141.1 116.3 122.6 137. 140.6 115.2 141.3 135.9 146.3 139.7 149.8 145.7 150.7 145.2 150.9 145. 145.5 131.2 149.6 121.5 121.8 142. 125.3 121.1 119.3 118.6 151.6 142.4 124.9 149.8 147.6 151.9 150.2 147.4 150.9 149.1 153.2 145.9 146.4 143. 135.2 145.2 143.1 119.8 145.1 145.3 143.7 144.8 122.6 148.5 144.3 146.1 148.1 145.5 145.7 124.4 120.1 142.5 144.3] | [ 93.4 95.4 92.3 90.2 81.9 76. 91.6 80.7 87.6 86.1 74.9 72.3 86.9 101.5 98.7 79. 86.8 87.7 85.8 84. 87. 86.2 88.9 72.8 94.1 75.2 75.1 63.5 86.9 95.7 97.1 80.3 54.6 67.5 89.1 87.6 74.7 57.6 84.6 87. 85.5 69.5 59. 75.5 93. 88.3 77.4 68.3 62.1 61.9 72.1 74.9 69.8 71.4 91.1 82.5 100.1 90.1 76.9 88. 85.7 80.5 76.8 83.6 100.7 72.9 100.9 86.8 73.2 82.6 86.9 77.5 63.6 84.4 82.6 75.6 80.6 84.1 70.4 80.6 78.8 77. 86.1 85.2 88.3 63.3 75.2 89.3] | [ 3.1 2.4 6.1 2. 2.3 3.2 78.3 2.5 52.5 1.9 56.9 1.3 1.1 1. 61.3 6.1 78.3 1.6 10. 63.6 2.4 2.5 1.3 1.2 73.1 81.4 78.8 3.3 1.2 84.2 1.2 57.1 6.1 2.5 1.8 9.5 1.2 68.2 60.9 2. 94.5 2.8 57.2 2.5 84.5 79.6 1.3 2.1 73.9 88.9 53.9 77.5 2. 50.8 1.8 79.4 81.9 69.9 8. 6.5 85.1 2.5 78. 77.6 2.6 61.7 66.6 78.3 57. 54.1 2. 82.1 98.6 7.6 81.4 3.4 65. 1.1 50.7 2.1 65.2 74.7 1.8 45.1 1.8 5.5 2.5 69.8 2.5 60.7 2.8 88.4 86.4 84.8] | [109.4 3.4 8.4 8.9 5.3 42.9 91.1 9.5 4. 2.7 1. 3. 9. 92. 78. 1. 2. 1.2 3.1 1.3 1.8 57.7 1.8 2.5 3. 62.3 7. 7.2 7.2 3. 4. 1. 1.6 1.8 65.6 93.8 1.3 44.7 5.1 1.2 2.1 47.8 1.6 6.2 8.5 1.2 1.1 2. 43.7 4.7 6.1 60.5 67.9 2.5 1.2 59.1 6. 88.6 75.6 83.8 97. 97.1 92.4 59.9 102.1 2.1 56.3 3.2 1. 1.2 62.5 1.3 47.8 5.3 4. 2. 10. 1.4 73.5 9.5 3.4 1.3 4.6] | [ 4.1 70.5 2.8 68.4 52.7 2.5 3.1 70.2 5.3 4. 3.4 54.9 65.2 1.2 49.4 2.8 2. 9.5 62.5 2.8 64.3 1.3 2.7 50.5 4.1 1.3 1.8 3. 2. 74.2 1. 53.9 68.5 59.5 66.6 9.1 2. 1.1 1.3 2.1 58.3 2. 94.8 58.4 58.2 68.3 2.3 1. 2.7 66.6 1. 4. 2.7 88.1 3.3 66. 3. 2.4 2.4 64.5 1.2 4. 3.6 105.7 9.5 59.8 2.7 39.9 1. 74.5 4.2 53.6 4.2 66.8 3. 65.8 3.2 60.1 2.7 81. 73.6 5.8 2.4] | [112.9 104.8 6. 2.7 2.3 3.1 2.1 3. 59.5 1.9 80.2 2.6 1.3 1. 1.3 57.9 80.6 4.8 4. 2.7 68.8 7. 4.6 3.4 1.3 2. 8. 9. 90.1 63.3 4.3 63.2 62.6 48.5 9. 1. 2.5 8. 2. 1. 7. 3.2 1.2 2.5 4.6 1.4 3.3 44.2 1.8 4.5 4. 4.2 4. 1. 1.2 1.1 1.2 2.8 2.5 1.2 2.6 1.8 4.1 40.8 2.7 7. 85.4 3.7 1.4 3. 1.2 120.9 70.7 40.8 10. 4. 73.4 48.6 1.8 2.5 1.9 57.7 151.9 1.2 9.5 44.2 114.2 8.8 2.5 174.2 4.1 1.1 9.5 4.1]
4.004 | mrel_usc_NIST | https://osf.io/t5y28/download/ | 20200204_mrel_usc_NIST/20200209_mrel_usc_GE3T_MR1_TR10sec_NIST_Magnitude_T1map.nii.gz | Magnitude | Nam Lee | Keck Medical Center of University of Southern California GE 1 | GE | Signa HDxt | 3 | NIST | 130 | 46 | 20.9 |  |  | 2D Spin Echo, IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 10000 | 14 | [50, 400, 1100, 2500] | 244.141 |  |  |  |  | [1760.8 1753.9 1755.2 1773.4 1716.6 1750.1 1759.5 1777.3 1768.2 1756.5 1759.1 1737.1 1698.6 1694.3 1744.4 1791.3 1791.2 1757. 1728.6 1729.6 1758.4 1737.2 1763.6 1759.7 1762. 1790.4 1797.7 2151. 1732.8 1747.4 1806.9 1780.1 1733.8 1828.6 1765. 1726.1 1781.2 1775.4 1752.8 1782.5 1774.3 1760.3 1753. 1747.6 1764.1 1742.7 1750.5 1777.4 1771.7 1786.2 1814.4 1779.3 1770.8 1795.4 1744.8 1738.9 1766.8 1800.9 1799. 1789. 1763.8 1723. 1717.4 1771.6 1779. 1758.1 1760.3 1754.9 1757.6 1770.3 1739.4 1702.6 1729.8 1786.3 1750.2 1763.8 1738.8 1742.2 1755.4 1792.9 1767. 1738.9 1759.2 1751.8] | [1281.5 1301.4 1288.5 1294.5 1304. 1287.9 1311.1 1328.2 1300.2 1273.5 1276.5 1301.1 1296.5 1293.7 1309.7 1287.6 1321.7 1299.9 1281. 1292. 1298.1 1310.7 1301.6 1295. 1312.6 1290.5 1298.7 1290.4 1284.3 1305.8 1289.9 1284.1 1280.7 1274.3 1296.7 1281.1 1287.5 1296.9 1291.5 1309.8 1287.7 1279.2 1288.3 1294.1 1295.8 1315.3 1295.5 1290.4 1292.5 1312.9 1291.1 1289.9 1303.5 1296.8 1314.3 1331.2 1297. 1298.4 1303.7 1316.4 1312.2 1303.3 1305.9 1305.5 1295.1 1283.9 1295.2 1304.2 1292.4 1281.3 1315.7 1315.3 1290.8 1300.8 1278.2 1288.5 1291.9 1288.3 1271.7 1296.4 1305.6 1289.8 1297.7 1287.9 1295. 1281.1 1284.6 1287.9] | [918.3 927.9 928.4 931.1 916.3 875.1 940.1 933.7 938.2 934.6 937. 931.3 916.9 893.7 936.2 944.5 942.7 935.5 935.7 931.8 933.8 935.5 922. 885.3 929.6 938.1 943.4 930.2 935.2 920.5 921.9 925.8 931. 906.3 930.8 928.8 934.3 924.1 927.2 925.8 933.6 925.9 925.1 927.5 922.6 928.4 929.2 921.9 932.2 939.2 924.7 926. 928.9 933.7 934.3 940.1 923. 917. 924.6 945.1 917.2 917.1 922.1 934.7 923.7 941.4 923.9 931.7 922.3 932.9 923.3 935.7 931. 932.6 944.6 935. 944.4 937.6 943.6 930.3 928. 918.1 931.7 935.5 930.5 937.4 943.1 940.1] | [666.6 660.9 669. 671.8 655.6 648.6 670. 669.1 671.2 697.8 670.9 674. 674. 647.7 668.1 669.7 671. 669. 674.1 666.7 693.2 707.7 673.4 640.1 671.4 668.1 676.1 668. 672.2 673.7 669.4 670.4 689.4 656.8 676.7 666.3 669.2 670.5 672.8 668.3 663.4 663.2 671.6 676.6 672.3 663.1 663.5 671.3 670.3 663.6 669.3 671.5 667.9 702.9 672.9 669.2 669.5 675.6 674.9 668.7 666.2 672.8 668. 697.2 672.1 666.9 667.6 665.5 668.9 667.7 662. 694.3 673.9 701.3 667.8 672.1 673.1 668.2 672.5 673.5 698. 681.2 671. 673.6 673.1 673.2 673.5 675.1] | [489.8 482. 480.5 482.9 483.1 479.1 482.3 475.9 484.3 490.9 491.9 483.1 476. 482. 487.1 481.9 481.4 481.6 482.1 480.7 479.6 481.3 480.3 477.9 486.6 483.1 473.7 476.5 487.4 481.8 479.8 478.4 476. 477.9 478.2 481.9 481.6 483.1 482.1 478.3 479.9 478.7 478.7 478.8 477. 483.8 481.3 482. 484. 482.2 478.9 482.4 478.8 474.8 482.1 490.4 482.3 478.3 478.7 474.1 474.6 483.1 481.6 480. 477.4 481.9 479.3 478.8 480. 478.5 479.9 483.2 484.9 486.7 483.6 484. 484.7 486.5 485.9 484.7 483.7 483.2 481.7 485.5 483. 478.8 478.8 484.5] | [325.2 335.3 338.3 337.7 330.3 337.6 342.4 342.5 336.4 333. 334.5 336.7 351. 343. 337.9 340.5 338. 339.7 339.8 336.2 340. 334. 339.6 329.3 334.5 341. 340.6 338.8 337.8 336.4 342.3 342.2 356.2 341.5 337.8 334.3 338.5 343.2 341.5 339.4 340.1 340. 341.7 362. 338.3 339.2 336.7 337. 343.9 345.4 342.5 343.2 336. 335.2 349.3 337.4 339.2 336.9 337.8 339.1 338.9 339.7 340.8 338.1 339.3 339. 339.2 339. 342.7 342.9 343.2 343.2 344.3 344.1 343.6 338.8 340.2 340.8 340.9 337.3 339.9 339.6 340. 337.1 337.5 338.7] | [240.2 243.3 242.8 241.3 239.9 239.3 236.9 236.9 234.8 237.9 242. 239.8 238.8 240.3 242.9 238.5 239.4 237.6 238.9 240.7 240.4 239.5 237.1 237.3 240.8 243. 241.6 242.8 239.7 237.1 238.3 238.2 234. 237.7 238.9 242.2 237.7 240. 239.1 237.7 240.3 241. 237.3 240.7 235.4 242. 240.8 238.7 240.7 241.8 240.2 240.4 238.7 239.4 237.2 238.7 241.9 244.7 242.9 239.7 238.5 238.4 237.8 241.6 243. 240.2 240. 240.7 239.8 238. 242.1 240.3 237.5 242.6 238.5 239.8 239.3 239.2 240.2 239.6 239.6 240.6 231.8 236.1 240.1 241.6 240.1 239.8] | [172.3 168.9 171.2 171.5 171.8 170.9 173.1 171.2 168.9 172. 172.8 173.2 170.1 173.2 171.3 170.9 170. 172.8 172.7 170.2 174.1 176.3 172.5 172.7 172.4 173.6 172.7 173.1 169.5 174.1 174.7 174.6 166.9 172.2 171.9 174.3 174.2 176.7 171.5 174.3 173.1 173.7 171.6 174.6 277.6 171.1 171.2 176. 175.7 172.9 172.6 276.3 175.3 172.9 175.9 173.1 171.4 171.5 280.9 170.4 174.9 172. 178.4 172.5 170.9 172.8 172.9 172. 284.1 169.3 173.8 168.8 172.6 172.4 169.8 176.1 174.4 176.4 171.5 170.8 171.3 171.5 169.2 170. 174.4 174.9 173.5 173.5] | [112.7 108. 124.3 110.4 111.2 108.9 113.8 120.5 138.3 117.9 117.5 118.1 141.2 113.5 133.8 116.1 111.2 115.6 114.5 116.5 110. 114.5 108.9 116.2 120.6 117.3 117.2 145.3 138.5 141.1 114.1 141.2 136.3 121.2 119.5 115.6 118. 118. 112.4 135. 115.2 116. 115.1 144.4 151.7 116. 120.4 119.8 138.7 138.8 138.5 141.7 134.7 139.1 143.6 112.9 119. 114.6 118.9 118.9 115. 134.6 137.8 135.6 107.3 136.9 119. 114. 117.9 118.4 119.1 117.5 142. 139.2 120. 120.3 117.4 132.3 112.8 117. 114.2 113.1 114.5 116.4 136.7 113. 112.2 109.9] | [90. 87.4 78.8 72.3 2.8 52.4 80.2 57.3 71.8 71.3 3.3 3.1 54.8 70.9 83.5 53.7 58.5 4. 47.3 1.8 56.4 57.5 4.1 70.8 86.6 66.8 66.6 1.3 66.9 67.9 66.8 60.5 10.9 64.9 79. 73.7 73.4 67. 62.8 66.6 68.7 45.7 62.9 68.6 79.3 75.5 74.9 75.3 63.1 59.6 71.3 66.1 68. 67.9 80.2 76.1 84.8 81.2 82.6 81. 70.8 69.7 76.9 64.1 90.9 76.4 87.2 78.5 79.3 68.8 59.7 77.7 82.4 65.5 67.7 50.1 85.3 89.5 74.6 76. 72.7 70.8 77.3 82.2 70.4 1.1 1. 59.6] | [ 3.3 6.2 3.3 3. 1. 6.4 2. 1.9 49.8 1.3 3.1 1. 2. 1.1 1.8 1.3 2.8 2.4 7.6 61.9 1.2 9.5 2.3 1.1 1.2 3.7 2.7 4.2 1.9 60.1 2.8 1.2 2. 7.3 2. 1.4 4.3 1. 1.6 3.3 79.5 1.6 3.4 1.1 55.1 6.7 2.8 1.1 2.4 1.3 3.4 67.9 9.5 2.1 2.1 1. 56.3 3.4 1.2 1. 3.7 1.3 3.3 54. 7.3 1.2 6.7 1.9 1. 61.7 1.2 39.5 75.6 3.6 1.3 1.2 6.1 6.6 78.4 3.6 2. 56.9 1.8 9.5 3.1 66.1 3.4 9.5 3.4 45.5 9.5 4. 48.6 67.1] | [ 1. 1.1 1.2 1. 2.1 8.1 79.8 1.3 1.2 2.8 1.8 6. 2.5 2.7 2.5 1.8 8. 5.7 1.3 1.9 2. 1.3 75.6 2. 2.1 1.9 7.3 1.8 9.5 2.5 46.6 3.6 48.1 1.4 2.1 2.8 1.2 1.6 1.3 1.2 1. 2.5 1.3 1.2 1. 1.2 1. 3. 6. 6.4 60.3 2.1 1.3 1.1 2. 2.7 1.3 3.1 2.7 1.9 1.2 9.5 3. 1.3 3.6 5. 2.9 3.4 4.3 67.3 75.3 1.2 1.1 48. 1.5 1. 10. 8. 55.1 2.5 4. 2.1 63.6 1.8 1.4 9.5 63.1 1. 67.9 9.5 1.8 2. 4.8 2. ] | [ 53.4 8.3 2.1 2.7 3.4 3.7 1.9 2.7 96.3 69.5 75.1 5.2 9.5 9.5 1.2 1.8 1.3 48.9 1.2 2.8 4.8 76.4 4.4 8.4 3.2 2. 84.1 2.7 10. 4.6 2. 3.8 9.5 2. 9.5 78.8 10. 1.2 5.7 65. 1.9 1.2 9.5 1.2 3. 7. 89. 3. 3.3 3.3 45.8 1.8 1.2 3. 6. 3. 1.2 6.7 3.6 2.8 9. 1. 3.3 83.3 2. 59. 2. 9.5 9.5 9.5 8.8 2.7 4.2 107.3 8.1 4.5 9.5 1.2 1.8 3.6 47.5 2.7 2.8 2. 1.2 2.4 3. 10. 4. 3. 3.1 4. ] | [ 1. 2.5 7.2 4. 2.5 2.4 6.5 4.8 4.5 3.1 10. 1.1 1.8 3.2 62.6 3.3 3. 1.1 2. 1. 1.2 5. 1.5 1.2 5.5 2.4 1.9 2.7 2.1 2.7 7.8 4.2 1.2 2.7 1.4 1. 1.2 1.2 1.9 4.8 2.9 9.5 4.9 9.2 7.2 1.3 9.5 4.1 9.5 5. 2.4 4.5 1. 6.1 2.5 2.3 1.1 3. 1.3 1.2 1.3 66.8 4. 5. 106.9 1.3 1.1 10. 1.1 2.4 1.2 84. 58.6 10. 2.8 2.5 52.9 54. 3.4 8.8 71.7 2. 1.2 1.2 108. 53.8 91.3 7.5 2.7 6.6 3.8 3.2 5. 53. ]
4.005 | mrel_usc_NIST | https://osf.io/t5y28/download/ | 20200204_mrel_usc_NIST/20200209_mrel_usc_GE3T_MR2_NIST_Magnitude_T1map.nii.gz | Magnitude | Nam Lee | Keck Medical Center of University of Southern California GE 2 | GE | Signa HDxt | 3 | NIST | 130 | 46 | 20.9 |  |  | 2D Spin Echo, IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 244.141 |  |  |  |  | [1780.9 1885.2 1993.6 1985.8 1926.4 1872.9 1906.8 1920.9 1952.8 1961. 1887.7 1810.2 1910.5 1943.9 1929.3 1861.7 1866.8 1906. 1898.1 1888.6 1862. 1878.9 1933. 1953.9 1898.6 1845.4 1881.7 1902.7 1911.6 1907.4 1851.5 1800.6 1868.4 1933.1 1917.5 1864.1 1903.6 1942.6 1904.3 1887.3 1918.2 1936.5 1866.1 1887.3 1950.3 1898. 1843.9 1860.5 1886.8 1920.4 1931.4 1957.5 1964.3 1877.4 1883.4 1954.9 1897.2 1898. 1895.1 1893.3 1946.9 1933. 1903.6 1906.3 1844.3 1788. 1905.7 1886.8 1898.5 1895.8 1893.1 1927.5 1924.4 1858.3 1850.2 1856.9 1910.4 1944.5 1890.9 1872.4 1891.1 1898.7 1910.6 1879.3 1839. 1915.7 1891.1 1882.1 1881.4 1880.3 1918.7 1948.5 1868.4 1857.4 1894.3 1927.1 1895.6] | [1366.6 1379.8 1375.9 1369.2 1372.1 1367. 1402.1 1382.6 1372.1 1364.3 1353.5 1346.4 1375.7 1384. 1351.5 1359.4 1363.4 1374.5 1351.4 1361.6 1366. 1374.4 1370.6 1349.1 1357.2 1368.3 1379.8 1385. 1351.3 1362.1 1377.6 1366.9 1352.2 1359.4 1338.2 1355.1 1365.8 1371.4 1348.2 1342.5 1358.5 1372.8 1370.1 1351.8 1358.7 1357.3 1359.9 1362.6 1353.2 1351.6 1362.5 1369.1 1367.4 1349.8 1374.1 1377.2 1362.4 1338.8 1344.4 1367. 1373.3 1368.1 1360.5 1354.4 1344. 1368.4 1386.1 1355.5 1353.1 1356.9 1348.4 1365.4 1378.7 1370.8 1352.2 1382.2 1376.1 1361.2 1363.2 1355.5 1345.5 1358.7 1360.3 1354.8 1342.7 1355. 1365.8 1342.2] | [957.1 963.2 962.3 965.3 959.8 952.7 956. 956.1 954.5 950.1 951.8 961.6 952.1 963.8 958.8 961.2 961.7 953.5 967.2 954.2 959.8 954.9 955.6 973.2 952.5 970.2 960.9 949.1 965. 962.2 961.9 961.7 960.3 961.1 941.4 962.5 959. 962.7 958.4 953.5 956.9 954.2 956.9 959.9 941.6 954.1 953.2 967.6 964.3 951.7 944.3 945.3 947.8 954.2 943.9 972.4 954.2 956.6 951.1 954.5 951.8 956.3 959. 948.2 949.3 969.8 952.3 952. 938.6 944.1 954.1 960.5 967.5 962.6 939.9 954.1 970.7 973. 965.6 955.7 956.7 957.6 950.7 954.1 952.3 953.6 960.4 958.5] | [685.2 689.7 689.8 685.5 687.9 692.3 682.6 686. 695.5 693.6 690.3 692.5 692.5 691.1 694.8 694.7 678.4 681.1 687.8 693.8 690.8 689.7 689.3 693.8 688.2 689. 682.2 687.6 683.5 689.1 685.5 685.2 689.2 684.4 682.2 688.7 691.9 687.8 680.3 688.9 693. 687. 680.6 677.9 680. 683. 687.1 681.9 687.8 685.3 688.6 694.7 688. 690.7 679.1 681.7 688. 681.5 688.6 684. 677.2 688.2 692. 697.5 688.3 693. 686.6 681.2 687.3 695.8 687.8 690.3 679.1 688.4 688. 682.2 690.3 683.6 686.4 686.4 689.1 680.6 681.7 697.4 688.1 685.3 684.7 696.7] | [486.2 486. 480.2 478.1 481.6 485.8 487. 487.2 482.4 486.6 486.8 485.9 489.5 489.6 494.3 491.2 484.2 484.3 487.3 482.7 480.2 483.8 486.3 487. 482.4 489.1 486.6 487.3 486.8 485.4 484.2 482. 482.5 482.2 482.8 486.5 484.2 482.5 479.4 483.6 484.8 481. 482.2 489.1 487.3 486.6 482.2 485.3 486.5 487.7 488.7 488.2 482.7 485.9 484.8 492.1 484.1 483.6 484.8 487.6 485.2 483.9 483.5 488.2 475.1 492.4 487.7 488.2 486.3 486.5 487.4 487.9 488.7 488.2 482.4 482.9 484.7 480.2 484.1 488.5 488.7 485.2 485.3 487.3 491.5] | [337. 335.4 337.8 335.6 327.5 341.3 346.4 340.3 342.9 344.5 342. 341. 352.5 346.1 342. 345.2 347.4 350.1 341.6 334.3 339.4 317.8 343.8 336.9 343.4 342.7 340.7 343.4 346.9 343.2 340.7 339. 346.9 342.6 344.9 343.8 342.2 342.1 339.8 342.2 342.6 337.5 341.4 336.7 337.7 342.7 342.7 346.9 343.4 343.3 350. 348.4 345.8 347.2 329.6 338.8 340. 349. 349.6 342.8 346.4 352.6 347.1 343.3 346.3 332. 344.2 343.8 349.9 347.2 341.3 336.7 337.7 341.9 339.3 339.3 337.2 348.7 341.8 340.1 345.1 341.6 340.9 347.1 344.2 342.7 343.4 340.7 344. 342.6 340.5 339.6 341.7 341.7 341.1 342.4 344.3] | [246. 241.5 241.3 242.5 241.1 240.2 244.9 241.6 242.8 239.6 241.6 241.3 240.2 244.2 242.9 239.9 239. 240.2 240.3 241.8 242. 242.8 241.8 243. 243.6 239.8 243.9 242.8 242.4 240.7 240.6 241.8 240.9 243.2 245.1 242.1 244.5 242.7 242.4 241. 241.9 240.8 242.8 245.3 241.9 241.7 245.6 242.6 243.6 243.6 242.9 241.7 242. 243.3 243.2 243. 246.9 243.4 245.2 245.8 242. 241.9 241.2 243.3 241.2 245.3 245.3 243.3 246.7 250. 244.9 242.7 242.5 243.5 246.8 243.1 243.8 242.6 242. 242.2 244.9 246.4 244.2 239.1 239.8 240.4 242.9 242.8] | [176.4 174.6 290.1 178.2 179.6 292.5 176.1 178.9 177.5 175. 172.5 179.3 174.2 175. 178.8 175.4 176. 174.3 179.4 181.9 183.7 174.8 287.6 176. 175.3 176.3 179.5 176.4 180.2 175.7 173. 177.1 295.8 175.1 177.1 175.8 176.7 175.3 179.6 180.4 180.2 175.9 176.5 177.1 179.6 175.2 176. 173. 174.1 174.1 175.4 170.6 175.3 177.1 178.8 174.1 176.9 174.7 285.8 174.7 179.6 176.5 176.8 169.5 176.4 175.8 175.9 175.4 178.4 174.1 178.2 175.7 179.4 178.9 177. 176.4 177.6 179.3 179.3 180.4 179.4 178.9 174.4 178.9 178.3 292.4 179.5 179. ] | [118. 118.5 142.4 140.5 121.1 116.8 140.8 120.6 123.3 137.7 135. 122.6 118.8 122.2 128.3 115.2 136.4 144. 142.9 123.1 124.5 121.3 140.8 122.4 144.9 115.7 121.3 124.2 123. 118.8 137.1 144.6 124.1 122.2 127.1 117.6 141.8 122.6 123.1 141.2 121.8 125. 124.7 143. 150.5 124.6 125.7 119.7 119.9 118.3 120.7 120.9 122.7 123.7 126.3 121.3 117.4 121.3 122.9 119.5 122.9 139. 122.2 123.2 124.1 149.5 117. 122.4 122.3 121.3 123.4 122.8 143. 127. 122. 123.5 123.6 122.3 122.2 122.2 123.3 125. 122.9 120. 121.4] | [ 91.4 71.4 84.4 85.7 85.3 81.3 92.6 85.3 82.4 84.7 83.2 85. 85.7 84.7 120.6 84. 74.2 89.5 92.3 84.6 89.4 90.1 93.3 89.1 96.6 86.5 82. 92.3 91.8 81.7 86.9 90.1 91.3 93.8 80. 95.8 89.2 94.7 87.5 87.6 89.9 86.7 84.4 89.6 81.8 94.2 84.3 93.5 87.2 87.1 83.5 89.1 93.2 91.1 87.4 100.1 92.3 92. 92.4 91.1 92. 96.5 93.2 86. 83.9 89. 95.8 81.4 83. 89.5 87.6 86.1 92.4 92.3 80.7 100.8 97.4 92.6 93.3 92.2 97.5 97.5 81.2 87.6 96.2 98.1 97.1 93.4] | [ 2. 42. 63.4 69. 77.8 74.8 8.2 81.8 60.8 85.5 10. 86.3 1.3 68.8 1.1 83.4 66.4 61.4 75.3 61. 1.2 3.1 1. 66.1 3.4 2.1 1.8 1.4 54.1 3. 68.5 1.1 67.6 44.1 49.5 2. 1.1 2.8 4.8 1.2 9.5 77. 9.5 1.1 42. 59.9 5. 73.3 73.6 71.4 10. 2.4 76.5 4.6 9.5 2.5 40.9 63.4 70.6 67.6 58.9 53.8 40.5 52.5 1.8 3.4 1.1 3. 52. 1.1 3.3 52.8 1.8 74.8 4. 57.7 45.9 1.2 2.3 54.2 60.5 2.5 67.5 2.5 65.5 2.5 59.7 1. 62.7 3. 1.8 2.1 61. 1.8] | [ 1.9 1. 1.3 6.6 3.2 2.8 1. 61.5 3. 73.7 1.8 1.8 3. 3.2 62.3 1.8 2.7 2.7 2.1 3. 2. 81.7 3.4 73.6 7.3 1.3 1.8 1.1 3. 55. 1.3 56.6 3.4 82.7 1.3 7.6 1.8 3. 6.6 2.4 6.2 2.7 4. 1.3 78.3 2.4 10. 6.9 2. 1.2 70.7 1.4 3. 2. 4.2 75.9 3.6 64.8 9.5 9.5 8.8 5. 7. 77.5 1. 1.1 4. 2. 72. 3.7 5.6 3.3 3.7 3. 49.9 4. 2. 3. 4.2 9.5 7.6 1.8 1.8 9.5 3.1 73.5 2.1 3. 2.7 79.2 69.5 2.7 56.7 1.1] | [ 2.7 2. 3.3 90.7 1.8 1.8 2.7 1.3 4.2 57.5 1. 7. 2.1 2. 4. 1.9 9.1 10. 75.8 9.5 47.5 1.3 1.9 2. 1.3 2.5 9.5 1.9 2.7 2.1 5. 3.3 1. 6.2 7.2 1.8 1.3 4. 9.5 5. 1.3 3.8 3.3 8.4 9.1 1.2 3. 3.4 7. 2.1 1.3 3.4 9.5 6.9 5.2 2.3 54.4 1.2 1. 3.3 1.4 2. 9.5 3.1 2.4 2.9 80.6 1. 3.4 5. 5. 4. 10. 5. 2. 5.3 2.8 3.1 1. 2. 1.2 1.2 1. 5.5 6. 10. 1.1 1. 79.2 1.3 111.7 7.6 1. 4.5] | [118.2 3.3 56.4 65.1 1.3 2. 1.2 8.3 5. 39.9 58.1 66.3 2.3 4.1 69.4 49. 60.8 65.1 3. 54.6 3. 53.4 9.1 58.6 2.7 56.5 63. 1.3 78.1 47.6 69.7 1.8 73.3 57.5 55.2 3.2 1.4 2.7 71.2 1. 5. 56. 1.1 59.2 7.2 2.5 2. 8.9 1.3 76. 1.3 6.3 43.5 2.8 2. 5. 7. 1. 6.9 2.7 7.2 82.9 40.1 4.1 1.8 52.7 2.4 1.8 3.6 7.2 9.2 83.6 5. 2.3 3.3 1.8 6.8 1.3 9.5 64. 3.3 10. 2.7 3.3 4.4 5.3 67. 3.2 1.1 169.8 50. 3.5 2.1 79.9]
4.006 | mrel_usc_NIST | https://osf.io/t5y28/download/ | 20200204_mrel_usc_NIST/20200209_mrel_usc_GE3T_MR2_TR10sec_NIST_Magnitude_T1map.nii.gz | Magnitude | Nam Lee | Keck Medical Center of University of Southern California GE 2 | GE | Signa HDxt | 3 | NIST | 130 | 46 | 20.9 |  |  | 2D Spin Echo, IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 10000 | 14 | [50, 400, 1100, 2500] | 244.141 |  |  |  |  | [2008.4 2004.2 2004.6 1960.6 1986. 1958.7 1928.9 1946.8 1974.8 1988.2 1970. 1980.1 1956.9 1973.4 1994.4 1932.5 1905.5 1948. 1981.7 1965.8 1914.9 1966.7 1941.6 1909.5 1897.5 1934.5 1945.2 1957.9 1927.7 1948.6 1915.6 1950.2 1990.4 1996.6 1981.9 1986.5 1967.8 1940.3 1938. 1962. 1957.6 1956.4 1947.7 1959.2 1983.8 1955.1 1943.5 1930.7 1965.2 1957. 2020.8 1971.9 1962.2 1945.6 1945.4 1945.9 2017.5 1985.4 1920.4 1932. 1943.8 1980.6] | [1382.8 1368.2 1360.4 1369.5 1373.5 1370.6 1375. 1347.3 1357.1 1379.8 1373.4 1375.5 1379.7 1395.1 1377.4 1363.8 1376.4 1384.4 1403.7 1370.5 1350.6 1361.9 1372.2 1382.8 1363.8 1375.8 1356.7 1378.2 1375.6 1371.5 1366. 1368.4 1372. 1375.1 1373.1 1387.9 1364.1 1377.1 1366.9 1395.8 1400.1 1376.5 1367.3 1379.2 1370. 1395.5 1375.5 1379.6 1361.6 1379.8 1396.6 1385.8 1354.8 1362.2 1396.5 1411. 1355. 1366.2 1354.1 1363. 1379.2 1384.5 1377.1 1375.2 1354.4 1391.2 1363.5 1373.6 1369.5 1360.2 1365.4 1373.4 1377.1 1377.1 1342.1 1384.8 1373.5 1368.1 1363.9 1371.7 1373. 1353.9 1357.9 1393. 1386.8 1375.8 1384.6 1375.2] | [963.4 966.7 960.5 951.2 941.3 952.8 967.2 959.4 945.4 951.8 955.2 924.8 952.1 963.8 974.4 967.1 961.4 963.5 952.4 950.5 947.6 968.3 950.8 969.1 971.8 964.9 962.5 965.4 962.1 967.7 989.9 960.6 962.7 959.7 956.4 971.7 962.5 959.7 958.9 962.3 991.5 950. 968.7 968. 951.6 968.3 959.7 948.7 960.8 958. 975.8 958.4 962.5 967.3 960.1 960.3 966.1 957.1 964.7 958.1 979.2 954.1 964.5 964.9 971.1 962.4 982.1 964.6 950.1 955.4 958.6 954.1 942. 964.4 962.2 960.1 961.5 952.3 962.3 951.9 965.2 981.1 955.1 961.7] | [687.1 674. 705.1 704.2 692.6 701.1 688.3 697.1 690.5 683.4 690.1 694.6 696.2 694. 690.4 694.1 700.3 674. 691.4 693.8 699. 700.4 693.3 687. 689.9 698.5 699. 695.3 698.9 697.9 696.3 698. 691.3 673. 686.2 694.1 707.2 703.8 695.4 697.1 698.3 698.3 695. 680. 690.7 701.3 699.7 692.9 694.5 696.7 694.3 701.8 698.3 700.2 698. 697.3 685.5 693.5 700.2 698.9 693.1 697.9 695.1 694.5 693.8 696.1 693.8 699.6 700.2 700. 691.3 676.6 697. 678.1 697.6 696.6 690.5 692.8 695.2 690. 690. 707.1 696.9 694.1 686.6 673.6 693.6 693.1] | [491.1 487.7 481.8 486. 487. 489.6 485.4 483. 485. 487.8 486.9 487. 529.5 488.8 487.6 479.7 479.9 487.6 488.9 489.7 489.9 488.7 482.1 488. 483.9 481.8 489.7 492.9 488.7 490.7 492.1 484.3 483. 481.4 484. 486.6 488.4 485. 485.7 496.3 489.7 484.9 484.5 489.8 489.9 490.3 491.5 489.5 492.7 484.7 486.5 490.2 485.8 485.1 486.4 487.1 490.9 491.1 488.5 493.5 486.7 490.3 490.3 485.8 490.8 494.6 487.8 491.1 489. 486.1 488.2 490.3 487.7 487.9 486.7 489.5] | [347.1 347.5 342.4 342. 344.7 347. 355.3 346.1 340.1 341.3 341.3 343.5 348.6 348.4 341.7 350.8 340.1 343.3 343.6 345.2 341.7 340.5 343.5 348.1 331.6 348.9 348.1 344.4 341.9 344.5 341.4 340.6 342.4 346.8 342.5 343.9 345.1 342.3 342.4 345.1 344.1 342.7 341. 343.9 348.7 347.3 342.5 343.5 343.6 344. 342.2 343.4 343. 343. 349.3 349.6 344.5 349.4 348.9 346.7 343. 346. 345.2 341.2 354.5 353.2 346.9 344.8 344. 343.2 345.8 346.3 346.7 344.4 341.6 342.4 347.6 348.1 346.9 342.3 345.6] | [244. 240.7 240. 242.1 245.2 240.3 242.1 238.1 239.4 239.8 240. 239. 241.5 244.5 241.5 240.9 243.1 242.7 239.2 243.1 241.1 245.2 243.7 242.3 241.3 243.4 244.4 243.8 246.1 243.1 244.6 246.5 241.4 240.7 244.2 245.7 243.9 246.3 245.3 243.1 242.5 245.9 242.9 241.1 242.9 244.8 241.8 243.6 244.2 245.9 244.7 247.1 246.8 245. 243.1 244.8 245. 244.9 244.7 242.5 245.3 247.8 245.7 244.9 244.4 246.5 246.9 245.6 245. 244.6 248. 245.6 248.1 247.1 246.7 248. 247.9 242.9 244.7 251.3 250. 243.5] | [279.7 295.7 284.7 288.5 176.8 177.4 175.8 174.7 178.2 173. 174.8 171.6 286.4 279.8 172.6 283.9 173.7 179.1 176.1 179. 177.4 177.9 174.1 283.9 174.5 290.3 176. 177.7 176. 177.4 176. 175.4 177.3 176.6 177. 179.6 287.6 295.1 180.1 302.1 180.3 178.1 177.6 178.1 176.7 179.2 178.6 179.2 177.4 177.4 178. 179.2 178.9 304.2 176.6 180.6 178.7 304.1 180.1 181.7 178. 180.8 177.4 304.2 181.2 183.9 180.4 180.8 176.4 182. 177.8 181.9 181.6 315.9 304. 179.8 180.9 180.5 181.5 179.8 181. 182.9 183.9 186.3 306.1 179.4 180.7 180.8] | [135.8 136.6 130.6 138.8 134.7 137.9 139.1 135.5 146.4 122.4 139.2 135.1 138.1 119.8 116. 146.6 121.2 142.1 137.9 137.6 137.9 144.4 126.9 122.4 117.1 147.1 120.6 144.9 145.5 142.5 138.6 140.9 142.9 121.5 117.8 150.1 122.9 145.5 123.9 126.5 146.7 146.4 146.8 143.5 116. 148.4 124.7 149.5 125. 128.8 150.2 146.4 148.4 150.7 144.1 153.6 126.3 125.6 142.8 127.9 130.2 149.3 148.6 152.2 152.9 155.8 127.1 126.4 125.7 126.3 148.5 153.4 150.2 148.7 149.6 152.5 130.8 155.7 156.5 155.1 152.1 127.4 125.8 124.5 143.3 144. 147. 152.7] | [ 84.7 82.3 76.1 77.5 74.1 83.4 3.3 79.7 79. 76.4 80.1 84.5 84.8 83. 92.3 66.1 81.8 82.5 81.2 83.9 83.7 88.7 85.4 93.1 92.5 87.6 91.4 87.8 84.4 86.1 87.5 88.8 86.6 87.6 97.6 83.1 90.1 88.7 89.2 87.8 93.8 93.8 89.8 89.5 100.9 81.7 94.5 87.5 90.3 88.3 88.3 90.7 93. 95.1 96.9 92. 100.1 91. 92.4 97.5 95.7 96.5 100.3 98.1 88.3 99.9 94. 98.3 97.2 91.7 90.2 99.4 102.3 99.2 105.4 103.7 96.9 94.7 91. 96.6 105.5 101.9 93.7 98.9 102.4 101.2 94.9 93.4] | [ 3.2 3.3 3.7 5. 1.1 2. 9.5 1.2 70.4 3. 51.5 1.1 1.1 6.7 2.3 5. 2.7 4.1 72.7 5. 1.2 4. 10. 9.5 2.4 4. 3.3 60.2 53.7 75.2 2.5 54.3 3.2 4. 2. 40.8 1.7 10. 1. 59.6 89.4 9.5 73.6 5. 9.5 1. 68. 3.6 9.5 9.5 81.6 84.3 3.4 83.3 65. 57.7 1.2 71.7 61.7 72.9 82.3 83. 80.2 57. 81.1 69. 67.3 78.5 83.5 59.7 82.2 89.5 99.8 59.9 78.5 75.7 89. 72.1 62.6 74.3 78.1 85.2 91.5 91.7 78.7 82.9 69.4 93.5 78.2 89. 87.8 81.2 75.3 87.3] | [ 2. 1.1 1. 4. 1.3 4.6 9.5 1.9 3.1 2.7 1.3 1.1 3. 8. 3.3 1. 1.2 43.2 5. 3.4 1.5 3.1 1.2 92.9 3.2 95.7 3.7 2.1 2. 2. 2.1 2.7 3. 91.8 7. 93.2 1. 2. 3.3 5.3 1.2 64.5 10. 4. 107.2 10. 53.9 2.1 89.4 4. 80. 49. 1.7 1.3 2. 98.6 7.1 93.5 1.3 92.9 5. 86.8 66.3 81.5 2.5 6.3 97.8 1.3 96.5 1.3 87.6 64.7 93.2 1.3 90.6 68.6 62.7 2. 77.3 89.6 2. 85.7 60. 94.6 75.2 98.5 74.6 76.6 78.5 90.2 91.7 65.1 77.5 46.9] | [ 98.9 1.3 4.8 2. 2. 3. 3. 3. 7.7 1.9 1.6 1.2 1.1 3.1 6.9 8.6 79.6 7. 3. 1. 90.6 3.3 2. 1.2 1.8 8.1 7. 6.5 7.3 61.1 2.4 1.2 10. 1.3 1.3 3.4 3.4 2.7 3.3 1. 5.8 6.1 2.8 1.2 1.8 9.5 1. 8.5 53.8 63.5 74.8 1.2 10. 2.1 3. 1.2 77.9 84.5 9.5 66.9 58. 82.7 1.2 52.1 4.4 1.3 110.6 114.3 108.1 93.8 3. 105.9 109.3 75.7 7. 5. 103.9 55.7 66.6 58.6 90.7 77.7 108.2 56.6 71.9 74.7 57.4 9.5 69.3 4.6 64.2 117.9 79.6 105.7] | [ 1.1 9.5 6.3 7.2 3.4 10. 10. 2. 1. 1.8 54.8 9.5 5.4 1.8 8.2 4.3 1. 1.2 2.3 139.9 1. 8.8 5. 1.2 2.5 1.1 3.7 3.7 116.4 1.8 96.9 4. 3.3 2. 3.2 2. 5.7 3. 1.2 3. 3.6 9.5 5. 7. 7.7 79.6 4. 1.8 54.5 1.2 2.4 2. 69.9 79.8 108.9 2. 4.3 69.8 123.3 69.8 3. 91.9 96.8 2. 4.2 1.3 6.3 1. 99. 113.7 88.5 146.5 1.8 137.3 154.2 88. 69.1 9.5 118.4 149.9 104.5 6.9 3. 10. 165.2 142.1 87.1 4.4 92. 168.7 120.9 91.7 78.5 209.1]
5.001 | guillaumegilbert_muhc_NIST | https://osf.io/vw24u/download/ | 20200210_guillaumegilbert_muhc_NIST/20200210_guillaumegilbert_muhc_NIST_Magnitude_T1map.nii.gz | Magnitude | Guillaume Gilbert | McGill University Health Center | Philips | Ingenia R5.3.1 | 3 | NIST | 130 | 42 | 21.2 |  |  | RFE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 364 |  |  |  |  | [2139. 2094.5 2136.9 2175.2 2197.1 2126.2 2109.8 2076.7 2069. 2090.7 2080.7 2061.5 2130.7 2130.4 2185.5 2158.1 2136.1 2090.6 2109.4 2113.3 2052.9 2042.2 2152.3 2137.3 2176.2 2107.3 2109. 2086.8 2096.7 2098.3 2083.9 2080.9 2029.3 2124.2 2092.4 2073.7 2096.9 2062.7 2056. 2085.5 2133. 2082.4 2080.6 2082.5 2100.2 2081.8 2070. 2065.5 2025.2 2049. 2030.2 2078.4 2068.9 2054.1 2106.9 2098. 2085.1 2108.7 2079.9 2050.6 2089.7 2076.8 2091.8 2083.6 2050.4 2113.3 2107.3 2094.8 2040.8 2093.8 2042.8 2065.4 2055.2 2024. 2123.9 2097.4 2107.6 2064.9 2123.4 2082.1 2073.5 2086.1 2027.7 2041.8 2122.4 2112.4 2092.2 2077.7 2131. 2086.7 2073.4 2066.4 2141.6 2136.2 2120.2 2081.9 2078.7] | [1468.4 1482.4 1468.4 1457.5 1444.9 1456.4 1476.3 1441.6 1445.2 1422.6 1411.9 1413.8 1406.6 1397.6 1459.2 1465.6 1449.8 1433.9 1440.5 1420.8 1424.8 1422.9 1426.1 1423.8 1436.2 1448.6 1461.9 1454.7 1466.6 1439. 1442.3 1453.7 1457.7 1451.9 1450.6 1450.2 1462.9 1451.2 1447.7 1448.8 1441.6 1452.8 1448.1 1433.8 1430.7 1452.6 1443.3 1441.2 1442.1 1443.2 1446.7 1459.7 1426.7 1424.7 1400.6 1449.6 1435.2 1438.7 1457.6 1456.3 1447.1 1433.8 1450. 1433.5 1396.9 1443.9 1423.7 1419.1 1432.9 1451.9 1448.1 1459.3 1435.6 1437.2 1417.5 1419.9 1443.8 1437.7 1417.4 1435.6 1451.9 1430. 1435.7 1455.8 1439.4 1437.3 1454.1 1443.9] | [ 988. 1002.8 993.6 989.8 991.6 994.7 1014.9 987. 988.5 981.5 981.6 994.1 978.9 983.5 1035.6 993.6 1007.8 988.9 985.7 978.4 986.6 981.4 985.1 995.6 1028.9 992.7 999. 994.7 987.9 977.8 986.7 984.8 974.1 982.9 1010.6 997.1 999.2 991.5 977.6 979.4 979.2 981.3 985.6 994.8 1006.8 1002.4 1004.8 986.6 973.9 981.4 974.5 966.5 977.8 994. 1007.2 1004.1 1005.2 990.9 968.7 972.9 970.2 987.3 987.8 986.8 1016.5 1002. 1014.4 978.5 973. 976.5 976.5 975.7 984.6 974. 978.3 997. 1010.2 998.8 993.6 997.6 987.6 979. 970. 994.2 1003.6 1003.1 1005.3 989.7] | [715.8 707.1 713.7 707.8 701.8 726. 709.3 715.1 713.7 717.3 711.4 698. 691.1 709.5 700.8 723.6 709. 706.3 706.3 702.7 704.9 719.5 696. 726.3 714.2 703.4 700.2 708.6 711.3 699.6 702.2 703.1 691.7 714.6 698.1 715.8 705.8 713.8 718.2 706.5 698.4 711.9 702.8 693.8 721.2 697.2 708.6 701.8 711.6 712.1 705.8 711.5 712.5 704.9 707.7 716.3 698.9 711.5 710.6 710.1 712.6 714. 710.6 711.6 712.5 712.3 719.8 721.4 712.9 699.4 706.7 714.1 708. 715. 723.8 706.6 704.2 725.3 706.4 707.1 712.3 724.7 714.7 713.3 696.2 710.6 716.1 717.3 718.1 713. ] | [504.7 512. 506.8 506.8 510.7 506. 515.8 508.4 515.8 512.1 512.6 507.5 505.5 513.6 504.9 508.6 510.5 508.7 515.5 519.5 514.7 512.7 513.6 511.4 505.6 505.2 514.8 515.2 512.4 511.3 512.2 511.7 509.3 512.7 511. 504. 514.3 513.8 516.5 515.2 517. 516.4 515.2 507.2 518. 506.5 516.2 514.8 513.1 515.7 514.3 512.6 517. 511.3 508.4 499.8 511.3 507.9 511.9 517.7 520.5 513.3 515.4 510.9 508.4 499.9 514.2 522.5 508.7 508.9 508.8 505.5 509.3 520.8 525.4 492.9 517.3 515.6 513.7 510.9 510.6 514. 531.8 512.4 504.7 505.6 509.4 509.7] | [355.9 357.3 362.6 360.6 357.3 360.7 356.9 358.2 360.8 362.5 360.3 359.1 359.7 358.1 358. 361.2 360.2 359.3 359.7 359.7 359. 361. 362.9 359.6 358.6 361.1 363.5 359. 361.1 359.1 360.4 357.9 358.6 363.6 363.6 358.8 359.7 364.5 359.5 360.2 358.7 362.1 355.2 357.5 363.7 364.4 364.3 359.2 364.1 360.9 367.8 366.3 366. 357.9 358.8 360.6 361.1 364.5 359.3 360.7 358.2 364.8 361.1 362.2 361.6 365.1 362.4 360.9 367.8 363.1 360. 364.2 365.3 366. 366.7 371.5 368.2 368. 368.5 369.9 368.9 369.7 363.3 364.8 369. 368.9 370.1 370.1 368.9 365.7 369.4 368.5 365.4 367.8 369.2] | [254.9 255.6 254.5 256.3 258.9 258.9 251.7 251. 254.2 253.9 256.2 252.8 248.9 250.5 257.2 255.6 257.7 257.4 256.2 258.2 255. 256.5 248.1 261.2 256.9 256.5 262.2 261.5 258.8 257.7 256.2 254.3 255.4 255.4 259.6 258.3 261.2 260.8 260.2 256. 256.2 255.6 260.1 256.8 260.1 259.4 262.1 258.3 259.8 258.8 258.7 255.3 261.5 257.2 263. 261.2 262.7 260.2 257. 257.5 258.6 255.6 262.9 256.7 262.8 260.3 259.7 259.2 260.5 258.4 259.4 256.8 263.7 260.6 262.2 258.3 258.1 260.1 259.3 259.3 259.1 262.5 256.2 260.6 260.6 261.1 261. 263.3] | [465.1 456. 449.4 441.7 435.8 194.6 409. 420.1 436.7 453.5 445.5 432.4 412. 184.2 441.2 429.2 505.4 424.1 433.7 440.3 441.2 428.1 482.5 407.3 445.8 438.1 431.2 435.9 419. 436.4 438.8 436.9 427.8 423.9 443.5 445.7 445.9 439.3 437.3 426.9 430.2 441.2 432.7 435.4 439.7 460.9 448.3 447.2 443. 450.5 433.8 441.6 431.6 444. 439.1 456.4 458.9 448.7 436.3 430.3 435.8 433.2 455.2 468.4 457.6 458.3 442. 452.3 439.9 443. 443.6 441.9 450.6 475.7 432.6 506.5 444.1 448.8 442.3 461. 440.2 500.7 431.5 450.9 467.6 469.4 469.6 459.9] | [190.8 192. 191.2 189. 181.9 135.5 188.8 215.3 192.6 190.4 185.8 184.8 183.9 186.6 190.7 192.9 188.7 189.9 185.6 184.4 183.7 183.5 178.6 173.4 183. 192.7 187. 186.4 179.9 185.4 189.6 186.7 184.1 179.7 176.5 186. 185.6 196.5 191.9 193.4 185.2 186.8 183. 181. 178.8 187.4 191.6 194.6 188.4 187. 185.8 188. 194. 181.1 186.7 182.8 180.4 192.8 190.2 195.3 193.3 187.5 192.3 190.3 186.6 184.2 181.2 187.3 195.1 190.9 197.7 189.8 201.3 191.7 157.2 182.5 186.6 192.1 191.6 197.2 188.5 183.8 189.8 192. 193.3 188.6 184.7 196.9] | [119.7 119.5 117.8 116.9 119.1 124.4 120.6 123.8 123.5 125.3 122.4 117.2 125.5 117.6 115. 129. 123.8 116.7 118.1 117.4 115.6 113.7 121. 120.7 124.3 123. 125.3 124. 120.2 126.9 127.1 122.9 118.8 114.3 126.7 126.9 127.9 126.1 120.9 122.6 119.6 120. 123.7 126.1 118.1 125.3 126. 126.4 125.9 128. 130.8 125.8 125.7 123.9 125.7 119.7 130.7 132.5 137.2 130.2 123.4 122.4 116.5 125.1 124.8 131.1 138.8 136. 134. 124.5 124.5 127.4 124.8 122.6 121.1 130.5 118.1 143.7 128.9 123.7 124.8 129. 126.9 131.6 120.2 129.3 127.5 128.7 131.1 133.3] | [ 95. 82.1 100.2 106.2 83.4 101.2 136.6 115.9 107.1 105.1 111.1 109.9 121.5 132.2 80.3 105.2 108. 106.6 102.6 107.4 104.5 97.1 119.4 99.6 114.5 109.1 106.3 109.1 113.1 117.9 115.1 117.2 107. 98.4 103.2 110.9 112.1 112.2 117.8 115. 104.1 106.1 108.5 108.7 104.5 115.9 118.6 111.4 115.7 110.6 114.8 111.3 112.4 106.3 116.6 107. 115.7 114.8 112.8 119.5 111.5 111.2 118.1 114.1 106. 122.9 117.5 117.9 113. 106.9 108.7 108.7 108.9 114.6 116.2 121.7 117.3 115.3 111.8 114.8 114.8 108.8 112.5 115.8 123.7 114.9 113.4 113.8 109.5 110.3 111.4 114.5 115.5 104.7] | [129.6 109.9 118.3 141.8 126.5 126.4 127.4 130.1 147.4 162.4 163.9 123.9 303. 116. 136.1 120.8 128.6 137. 160.2 176.3 122.3 451.5 118.3 124.5 135.9 126.7 149.4 162.4 187.5 7. 239.8 704.5 97.7 120.3 128.8 142.8 132.1 160.6 124.3 196.5 95.8 65.1 333.1 84.9 108.8 120.9 141.2 141.4 160.9 91.3 161.6 75.9 8. 331.9 115.7 127.6 143.2 132.6 137.1 110.5 128.1 105.9 1.2 182.5 121.2 127.7 139.3 132.9 122.9 121.3 113. 119.8 102.1 133.9 160.1 136.7 134.1 126.1 132.7 118.5 142.6 107.5 134.2 139.4 141.4 143.5 136.7 96.2 149.2 141.2 140.6] | [ 96.1 89.7 89.6 99.1 120.4 124.6 137.9 144.4 152.8 147.9 121.1 111.4 101.8 113.6 113.1 121.9 138. 152.8 145.9 134.1 117.7 118.5 139.4 152.8 138.1 141.5 141.3 139.2 118.6 165.8 143.4 115. 106.5 81.9 134.3 154.6 145.8 134.7 130.6 126.9 168.2 146.8 133.6 148.1 146.8 149.3 139.9 135.2 127. 131.9 121.4 179.6 147.5 147.9 125.9 134.9 129.7 116.7 147.5 140.1 139.3 120.5 148.5 143.4 126.8 120.7 136.8 126. 146.4 137. 140.7 151.9 93.4 171.5 124.3 142.7 149.1 162.2 139.8 146.2 98.7 131.3 164.7 165.8 153.3 156. 164.6 166.9 140. 127.1 140.2 139.4 115.8 148.7] | [281.4 283.3 261.1 597.2 81.1 315.8 290.4 332.7 338.2 377.1 193.8 315.2 282.4 252.8 212.5 270.1 207.8 209. 177.6 219.9 246.2 251.6 239.7 202.9 192.8 239.5 268.2 197.1 176.4 174.7 235.1 198.9 209.5 160.9 176.8 167.3 86.6 64.5 156.5 169.6 136.9 224.5 153.4 176.3 169.9 208.4 205. 243.2 168.4 159.3 152.5 114.2 189.3 138.2 163.8 140. 148.3 86.2 86. 67.5 85.1 173.9 143.1 266.3 181.5 136.5 164.7 194.1 190.3 193.4 200.2 223.6 141.9 191.2 145.8 203. 190.4 214.2 169.7 227.7 238.9 102.4 769. 8.5 175.5 209.8 317.4 332.7 334.4 582.7 176.1 217.4 293.6 301.5]
5.002 | guillaumegilbert_muhc_NIST | https://osf.io/vw24u/download/ | 20200210_guillaumegilbert_muhc_NIST/20200210_guillaumegilbert_muhc_NIST_Complex_T1map.nii.gz | Complex | Guillaume Gilbert | McGill University Health Center | Philips | Ingenia R5.3.1 | 3 | NIST | 130 | 42 | 21.2 |  |  | RFE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 364 |  |  |  |  | [2131.8 2046.5 2133.7 2226.5 2169.6 2123.1 2088.3 2047. 2049.6 2092.9 2074.8 2055. 2139. 2120.5 2200.9 2196.8 2201.3 2113.2 2148.9 2183.1 2054.1 2055.3 2171.4 2120.7 2232.1 2096. 2151. 2126.8 2133.4 2105.3 2139.7 2043.8 2013.3 2162.2 2100.5 2067.8 2109.9 2077.3 2061.5 2133.4 2176.1 2117.9 2055.6 2087.6 2102.8 2054.8 2084.7 2069.7 2022.5 2061.4 2030.8 2102.2 2079.5 2028.9 2084. 2091.5 2098.8 2167.5 2114.2 2071.5 2147.4 2118.7 2114.9 2114. 2013.4 2088.2 2102.6 2119.6 2041.5 2142.2 2071.7 2104.1 2082.8 2019.5 2150.9 2137.9 2083.7 2029.4 2179.5 2112.8 2100.1 2121.1 2036.9 2032.4 2171.7 2092. 2081. 2049. 2183.2 2133.9 2116.7 2098. 2149.3 2150.9 2143.8 2066.1 2109.1] | [1459.6 1472.1 1458.9 1472.4 1465.2 1499.6 1491.9 1442.4 1464.4 1418.3 1408.4 1426.4 1417.1 1422.3 1456.6 1484.8 1473.8 1448.9 1452.2 1422.3 1438.5 1445.7 1443.4 1470.7 1429.7 1455.6 1480.5 1468.5 1493.4 1458. 1449.8 1480.1 1479.8 1498. 1426.8 1449.2 1482.6 1461.2 1453.8 1455. 1442.2 1469.9 1468.5 1442.4 1446.6 1461.8 1445.3 1438.4 1447.5 1443.1 1450.8 1467.5 1426.4 1435.9 1399.8 1474.7 1443.5 1446.2 1463.6 1464.3 1453.7 1437.6 1469. 1453.3 1400.3 1475.6 1445.5 1425.2 1445.3 1455.7 1456.6 1475. 1455.8 1450. 1435.1 1434.6 1483.3 1448.8 1409.4 1445.3 1474.6 1432.2 1460.5 1493.1 1428.7 1424.6 1448.9 1447.9] | [ 990.7 1010.7 989. 985.3 985.8 989. 1021.2 989.3 993. 981.2 983.4 1008.3 975.3 981.1 1027.8 991.5 1028.3 993.3 1000.1 983.7 1000.8 987.6 997.7 1006.6 1022.4 987.6 1002.3 1009.7 1007.6 998.8 1008. 1012.6 971.6 986.9 1005.2 995.7 1013.1 1008.5 994.3 1002.7 1001.3 999.9 999.7 1005.8 996.5 1004.6 1018.3 993.9 992.3 1012.8 1000.7 980.7 981.6 1010.2 995.3 999.3 1005.1 1014.6 990.8 995.2 989.6 1009. 999.2 988.6 1008.6 1003.6 1031.5 972.9 980.2 992.3 990.6 981.8 1000.9 979.3 970.8 1007.2 1034.3 1019. 996.4 1013.8 994. 991.3 965.5 1009. 1007.5 997.9 1014.3 992.2] | [713.5 705.1 716.6 707.7 697.5 731.8 714.8 722.4 716.3 727.8 719. 696.9 698.6 713.6 703.9 728. 711.2 703.5 705.8 704.9 713.1 737.1 708.6 728.8 726.7 704.4 697.6 714.7 718.9 704.7 700.8 709.8 688.3 716.5 699.5 727.8 709.3 722.4 725.4 709.3 696.5 723.4 706. 687.6 725.8 705.4 719.6 701.2 717.2 715.5 705.7 718.7 718.6 707.1 713.2 715. 697.2 711. 711.3 716.2 717.7 718.3 711.5 715.2 714.3 717.8 727.5 726.2 707.1 691.9 706.3 717.8 704.9 714.3 734.3 704.6 696.3 733.5 702.8 704.7 710.2 735.7 713.2 712.6 695.7 711.6 719.4 712.5 717.7 708.8] | [501.2 511.8 502.3 502.1 513.2 505.4 520.8 505. 519.6 506.9 508.8 501.4 500.4 514.6 509.9 511. 511. 500. 511.9 516.1 511. 509.3 515.2 508.9 508.3 498.6 516.7 512.8 506.3 502.9 507.2 508.6 509.5 513.2 515.1 496. 510. 509.5 513.6 510.2 516.4 517.3 517.9 504.5 522.3 504.1 518.6 513.3 507.9 511.2 510.9 509.6 520.3 511.7 513. 495.8 510.4 498.3 507.2 518.4 523.6 512.5 519.2 514.8 515.7 498.4 515. 523.5 501.2 503. 506.5 499.8 511.9 530.4 529.3 486.9 513.5 512.5 514. 507.8 509.8 522.2 529.4 506.8 503. 502.7 507.1 508.9] | [357.8 356.1 365.8 358. 355.4 367.8 358.4 359.6 363.1 364.6 362.5 364.8 360.1 358.2 360.4 364.7 362. 361.4 361. 362.7 360. 364.4 365.8 359.9 359.1 361.3 365.3 360.2 361.5 359. 361. 358.5 359.1 364.5 363.9 357. 357. 366.2 358.4 359.8 356.4 363.8 351.4 355.8 365.7 366.6 364. 357.4 364.2 359.5 369.8 367.5 368.2 358.8 360.9 358.7 360. 367.3 357.7 359.2 356.1 365.4 356.6 357.7 359.4 367.4 360.4 353.9 370.3 360.1 355.9 362.3 360. 365.2 362.8 371.4 365.4 364.5 366.7 370.5 366.9 368.4 354.4 357.5 363.9 362.6 367.6 366.4 362.8 358.8 362.9 365.5 357.9 365.5 365.6] | [255.1 255.8 251.1 256. 263.5 264.4 253.1 248.8 255.6 252.4 261.8 253.4 249.4 250. 259.1 259.3 257.5 257.5 254. 261.6 253.6 264.4 248.1 263.2 254.2 252.2 267.1 261.8 259.9 256.6 260.4 252.2 256.9 255.9 260.8 256.8 260.6 257.6 261.4 252.6 255.5 255.3 263.5 257.4 258.8 257.7 265. 256.2 261.5 257.8 263.2 252. 264.5 256.5 263.3 259.6 264. 257.6 255. 253.9 260.2 252.8 267.6 254.2 261.4 256.3 261.7 258.8 263.6 256.7 263.8 254.7 268.2 263.7 262.8 257. 256.7 258.6 258.3 258.3 259.7 263. 252.8 265.5 263. 262.3 261.3 263.9] | [192.2 193.3 192.2 193.5 195. 195.6 183.1 187.1 192.2 195.5 196. 192.3 181.4 183.2 189.6 188.2 195.3 183.6 189.7 190.1 194.4 189.4 195.5 187.3 191.7 189.8 187.3 190.8 188.6 192.6 198. 189.8 188.3 184.3 190.7 192.3 192.3 197.3 191.7 189.7 191.2 194.6 195.6 193.5 193.4 199.1 188.4 197. 195.5 193.4 191.3 191.7 188. 185.5 191.6 194.7 192.6 197. 188.4 188.4 188. 190.5 196.9 194.5 193.2 193.4 182.3 193.6 193.2 190.2 194.1 192. 192.3 194.5 188.3 206.3 194.9 192.7 188.9 197. 190.1 199.2 192.4 194.5 197.8 196.9 196.4 196.2] | [142.8 141.1 141.6 142.6 139.3 132.9 145.6 156. 140.1 143.1 141.6 143.5 141.8 147.3 141.4 145.3 138.2 143.8 141.3 138.8 141.2 140.1 137.8 134.5 136.9 147.2 140.9 139.7 134.6 139.7 146.9 145.2 144.9 142.8 132.1 142.9 135.5 144.6 145.4 145.1 141.2 141.1 140.5 142.9 136.6 142.7 148. 145.6 141.8 142.8 140.2 144.1 151.1 139.7 144.4 138.8 132. 144.6 143.6 147. 146.8 136. 143.5 144.9 143.2 145.5 138.9 140.7 149.1 140.3 151.9 141.6 146.9 143. 123.7 140.8 141.3 142.3 142.5 149. 142. 137.7 142.1 144.5 146.6 142.9 136.1 143.6] | [111. 111.7 106.9 106.9 113.2 118.5 111. 116.9 116. 124.9 115.8 111.5 122.5 106.8 106.4 120.8 116. 108.4 112. 110.8 108.3 103.6 118.4 117. 119.2 110.8 121.2 118.9 114.9 124.1 123.2 117.8 115. 102.8 123.1 121.1 122.3 120.7 114.2 118.3 114.6 112. 117.5 120.7 112.1 116.2 118.6 116.3 116.7 122.4 129.2 120.6 119.9 116.8 116.7 110.6 123.4 123.9 132.4 124.7 115. 116.2 109.3 122. 120. 127. 129.7 128.9 120.5 112.6 116.8 123.3 117.9 114.8 109.9 123.7 114.7 132.1 118.1 117.3 115. 122.4 119.6 124.3 106.5 125.1 121.1 119.7 120.8 126. ] | [ 97.2 90.5 111.3 117.3 96.4 110.2 136.4 116. 108.1 108.8 122.5 117.5 127.3 136.3 77.1 108.1 112.9 111.2 107.6 113.9 108.7 99.6 126.9 112.3 120.5 110.9 98.9 109.7 114.8 123.3 122.2 123.3 112.4 102.7 100.1 109.6 113.3 110.9 121.4 118.9 97.9 105.2 104.2 116.6 105.4 126.1 124.8 109.5 119.9 110.6 119.1 112.9 117.8 107.2 120.6 103.3 116. 117. 115.9 126.9 117.6 115.3 120.4 114.4 106.6 130.7 120.8 125.9 114.2 106.2 111.2 111.4 110.9 118.2 123.9 125.3 119.3 122.7 118.6 124.1 123.9 118.1 116.2 117.8 129.3 117.7 120.4 120.6 116.6 116.6 118.2 120.2 131.8 112.6] | [133.2 112.2 116.3 135.8 127.6 124.1 120.3 147.1 147.3 149.1 154. 120.5 347.4 127.6 153.1 103.8 121.6 136. 142.2 165.2 62.9 621.3 120.6 111.5 137. 123.3 153. 150.9 190.2 3.1 249.3 807.4 111.7 139.2 140.5 148.7 122.7 166. 117.8 201.3 111.4 65.4 340.6 94.6 121.3 118.6 144.3 144.1 161.8 92.5 160.8 79.1 1.9 336.1 134.6 137. 153.1 133.1 135.7 121.7 128.9 102.5 3. 182.6 124.4 117.5 143.8 130.5 121.1 122.4 113. 113.3 102. 129.5 174.2 129.7 136.5 115.3 135.6 123.3 141.6 105.4 132.5 140.3 133.8 142.3 126.1 85.5 144.9 136.5 137.3] | [129.5 125.4 114.9 119.8 165.2 166.4 186.4 192.8 179.8 162. 134.7 160.7 121.2 138.1 116.6 128. 143. 167.4 169.3 146.8 140. 150.3 199.9 191.6 166.1 159.9 149. 143.4 119.2 168.4 160.7 128.6 119.8 132.2 152.6 180.8 161.3 146.8 143.5 152.6 181.5 159.6 133.9 192.6 181.9 178.3 162. 144.5 133.8 149.4 143.4 189.7 152.3 183.2 144.2 153.9 147.3 102.2 170.7 156.9 160.6 141.2 154.3 160.7 139.6 151.5 157.7 140.5 166.4 139.4 153.7 171.1 152.4 193.8 113.2 178.7 161.1 198.9 128.8 166.5 182.6 147.2 184.8 191.3 170.8 174.1 183.2 192.3 162.4 119.6 155.1 157. 108.2 149.7] | [185.4 174.8 129. 270.6 3.5 419.4 224.3 307.9 268.2 528.4 119.4 281.3 308.7 203. 152. 276.8 111.1 184.6 119.3 305.3 282.2 242.1 213.5 154.6 173.7 292.4 446.9 229.5 195.9 143.4 231.9 242. 270.2 150.6 202.6 173.1 197.9 93.6 213.8 246. 158.7 241.5 190. 214.3 208.3 328. 288. 385.7 227.8 201.4 189.1 119.7 275. 178.8 260.9 160.9 161.9 243.3 279.1 127.3 246.4 272. 211.3 356.4 309. 144.6 187.2 260.4 248.7 243.9 273.2 391.1 181. 313.9 200.3 349.3 214.5 215.2 127.9 271.4 236.5 150. 805.3 7.5 250.8 196. 338.3 361.9 279.6 564.8 193.9 231.4 294. 224. ]
6.001 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Aachen_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Shuo Zhang | UKA RWTH Radiologie | Philips | Achieva 3T | 3 | NIST | 130 | 23 | 22.1 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1968. 2075.1 2100.1 1994. 2058. 1998.6 2040.5 2107. 2043.6 2140. 2154.4 2069.5 2077.8 1968.6 2153.9 2144.8 1995. 2110.5 2030.1 1960.4 2188.9 1940.1 2042.1 2113.9 2138.3 2080. 2071.4 2117.2 2169.5 2195.1 2076.5 2013.3 2091.8 2096.4 2157. 2169. 2101.3 2100.8 2226.8 2185.5 2139.6 2089.7 2154.3 2157.3 2220.9 2160.1 2191.5 2055.6 2044.4 2137.1 2039.7 2098.6 2127.6 2130.8 2085.5 2144.7 2184.9 2194.6 2085.6 2013.4 2046.8 2093.7 2094.4 2094.7 2076.6 2097.1 2172.2 2097.9 2161. 2172.1 2158.4 2185.5 2105.8 2245.1 2154.5 2137.2 2119.6 2154.2 2136.4 2159.7 2179.9 2165.8 2126.3 2041.3 2087.7 2115.2] | [1540.1 1556.4 1505. 1481.7 1536.9 1477.7 1491.4 1451.4 1517.1 1515.3 1490.2 1468.9 1425. 1506.9 1494.6 1528.9 1481.2 1553.2 1474.7 1447.6 1500.2 1492.9 1476. 1479.7 1451.6 1490.7 1525.4 1495.5 1443.7 1462.3 1481.9 1489.7 1435.9 1505.3 1472.2 1497.8 1494.4 1501.9 1519.1 1470.5 1466.5 1512. 1470.3 1511.7 1479. 1467.3 1516.9 1500. 1497.4 1477. 1481.1 1496.7 1519.5 1505.6 1480.9 1514.1 1492.8 1479.8 1551.2 1526.5 1494.9 1520.4 1572.6 1483.4 1483.2 1501.7 1498.7 1501. 1541. 1517.3 1505.4 1500.7 1482.1 1504.8 1530.5 1492. 1519.1 1503.3 1520. ] | [ 973.4 993.6 991.7 1000.9 992.3 986.8 967.4 1014.6 1024. 984.4 992.7 997.6 1010.5 1018.3 978.3 1033.8 1003.7 1005.4 975.7 986.4 996.7 981.8 1005. 989. 973.3 1009.5 988.4 994.6 997.4 1003. 1006.8 1001.9 1004.7 1012.6 1002.4 1004.6 1010.4 997.7 996.2 1005.3 999.1 1004.4 991.4 984.2 1012.9 982.7 992.9 1009.9 997.9 1031.6 996.4 985.7 998.6 990.4 1005. 993.2 996.2 994.2 991.5 1001. 988.9 983.4 1000.1 995.4 1024.8 989.7 988.6 1002.8 1012.3 987.4 1002. 1020.3 1008.5 1024.9 998.9 980.1 1006.4 999.8 991.8 1001.8 1021.5 998.7 986.2 1006.4 991.3 992.1 1004.3 1014.9] | [727.7 714.3 733.1 722. 707.1 715.1 703.7 703.3 717.4 724. 712.3 709.6 725.3 710.4 710. 713. 723.8 729.9 719.8 705.4 728.3 717. 723. 711.8 705.6 720.1 719.6 718.4 717.7 719.2 697.8 704.5 727.1 710.7 711.9 709.1 728.4 715.8 710. 717.1 722. 719.6 714.4 710. 713.6 725.6 701.3 720.5 715.8 714.4 708.5 720.1 722.6 717.8 719.2 728.1 712.2 718.1 708.7 713.1 722.9 716.3 710. 710.6 716.1 722.4 721.5 722.4 709. 724. 709.4 706.5] | [508.1 500.5 498.2 500.5 501.1 505.5 483.4 503. 501.3 509.2 501.1 486.3 500.9 503.6 490.1 501.5 502.4 503.5 506.1 505.7 500.2 507.9 501.2 500.4 495.3 499.8 492. 499.5 497.3 500. 494.8 502.6 508.3 507.2 497.5 495.8 497.5 508.4 500.2 503. 500.2 506.5 513.3 507.5 498.1 497.4 496.7 514.8 508. 489.2 496.2 497.4 498.8 497. 500.3 500.7 492.4 508.8 509.5 494.5 493.5 503.6 494.1 499.4 498.7 494.7 485.3 496.4 493. 504.8 505. 501. 503. 489.4 496.7 497.7 502.9 498.2 497.7 506.3 503.9 495.9 497.3 491.8 495.4 499. 489.7] | [324.1 313.6 322.9 325.2 325. 324.7 328.6 332. 318.8 315. 318.9 320.5 313.6 318.7 319.8 315.6 317.3 318.4 329.6 319.7 318.7 321.7 314.9 319.8 313.9 324.7 313.4 319.1 311.4 318.4 326.3 318.7 320.9 324.2 330.6 315.7 331.5 326.2 318.8 319.4 318.8 322.5 313.7 319.7 323.7 320.6 319.7 318.8 328.1 327.8 315. 317.6 322.7 307.5 316.5 315.3 313.8 319.6 320.7 312.7 322.4 322.2 319.2 325.8 316.8 318.9 317.5 313.6 316.6 322.8 311.8 316.1 312.5 320.8 315.1 315.2 313.7 308. 329.4 322.7 315.4 316.8 314.5 328.6 321.1 314.4 322.2 319.2 314.7 318.8] | [248.4 248.8 252.2 246.5 247. 255.3 250.8 248.1 245.4 244.3 252.2 249. 249.2 249.7 245. 251.3 253.5 250.2 250.2 253.9 250.4 252.2 248.2 254.1 250.9 255.6 248.4 249.4 246.8 248.8 248.3 251.9 251.7 249.1 247.4 252.9 247.8 251.9 248. 248.1 249.2 248.7 241.2 244. 246.2 245.3 243.6 245.2 247.5 248.4 247.9 251.2 241.7 251.4 243.5 256.2 245.4 241.8 243.8 249. 243.1 250.6 249. 250.3 243.6 241. 249.2 242.2 244.5 243.4 243.4 240.5 240.3 251.1 238.4 247.2 247.3 246.2 246.1 257.7 245. 254.9 239.8 252. 253.6 249.7 245.3 241.2] | [164. 170.6 165.1 167.2 166.8 168.9 172.5 166.9 172. 170.1 169.5 171.1 167. 162.6 170. 168.4 168.5 173.4 164.5 166.6 169.6 167.2 167.5 166.9 170.3 174.2 171.2 169. 165.4 360.9 168.2 349.7 170.6 167.9 165.9 169.3 170.6 337.6 168.5 173.2 167.8 165.3 169. 174.1 168.9 378.8 167.1 168.3 162.8 167.4 168.7 172.6 167.2 173.3 159.4 164.6 169.4 163.2 333.4 167. 163.8 163.2 160.6 395.4 403.7 164.7 174.9 164.4 318.8 165. 167.1 166.1 397. 163.4 354.6 334.1 334.9 162.7 162.2] | [129.7 169.7 167.8 129. 154.3 120.9 125.8 124.1 171. 156.8 120.1 127.2 167.8 124.8 128.2 128.9 125.8 130.9 157.5 163.7 129.6 173.4 130. 126.7 166.9 123.5 123.4 164.9 174.5 169.3 122.6 167.7 124.3 125.7 160.2 171. 157.6 173.1 172.4 172.5 129.2 169.5 153.5 168.3 158.6 175.8 170.1 123.2 159.2 119.9 112.9 179. 131.4 188.2 117. 185.2 180.9 131.9 174. 170.5 160.1 187.3 109.7 190.9 181.7 112.7 123.2 123.3 168.2 122.3 166.4 116.3 175.4 163.8 163.4 176. 173.9 181.7 188.1 184.6 162.6 118.9 106.1 111.3 113.5 109. 110.8] | [116.7 98.3 98. 116.6 107.4 84.4 97. 84.7 98.3 92.4 94. 96.1 88.5 85.2 90.8 81.5 104.7 79.4 90.3 100.3 97.2 93.3 99.7 94.5 77.9 90.1 84. 103.5 96. 87.3 85.6 76.5 101.8 95.5 80.6 93.4 107.3 107.6 98.7 89.5 61.4 95.3 103.8 85.1 75.1 91.3 84.6 94.2 1.2 104.8 76.7 112. 92.2 106.4 83.5 107.2 82.3 98.1 134.7 2.9 71.7 81.2 96.1 93.9 81. 83.4 134.3 2. 86.3 79.2 82.9 61.1] | [ 75.7 57.5 72.5 101. 58.1 69.8 72.7 66.3 78.6 72.6 94. 1. 84.4 64.2 83.6 3.4 64.2 87.3 69.7 76.1 94.2 4.6 90.5 72.1 84. 46.6 3.7 2.4 76.6 84.3 40.1 67.2 2.3 73.9 2.8 76.1 63.5 9.5 5. 85.8 87.2 6.8 81.7 3.1 79.8 3.4 48.6 1.1 51.4 86.8 51.8 3.7 61.4 58.9 3.1 93.6 3.4 1.2 4.6 1.3 5. 2. 86.2 3.6 4.7 94.5 80.8 69.7 57.2 2. 97. 1.8 104.1 3.3 59.3 9.5 9.5 1.2 5.6 1. 1.7 79.1 71.2 94.6 86.3] | [110.4 92.2 117.5 106.7 120.4 1. 158.4 83.7 60.4 80.4 61.4 59.8 4.9 73.1 93.3 2.5 96. 74.5 75.3 76.2 87.4 112.9 76. 59.4 3.7 77.2 93.3 76.1 82. 77.3 94.4 62.9 1.1 102.8 58.4 7. 3.1 94.7 79. 1.2 78.1 3. 53.7 4.2 104.4 2.7 1.1 71. 57.5 88.7 73.8 62.2 70.3 98.2 3.3 65. 2. 69.1 1.1 63.2 74.7 2.7 9.5 2.1 73.3 1.2 107.2 9.5 9.5 47. 91.7 3.4 2.7 5. 2.1 3.6 66.8 3.1 3.1 6.4 1.2 62.2 1.4 3.4 2. 1. 80.8 2.8 3. 1.8 5. 59.6 2.1 2.7] | [ 3.2 74.6 79.2 1. 74.7 1.8 146.4 87.2 87.7 70.9 69.8 116.4 2.6 66.7 73.7 83. 80.8 70.1 2.7 60.9 78.2 136.2 2.2 99.1 7. 62.3 159.9 171.8 100.8 84.6 1. 65.6 3.7 92.3 75.7 77.8 1.3 63. 83. 101. 100.1 1.1 87.9 2.9 71.9 64.1 119.5 58.3 67.9 76.5 1.3 1.3 3. 111.5 54.9 1.8 93.3 2.8 57.2 120.8 3. 70.3 58.3 54.3 3.1 70.3 4.9 1.4 4.1 1.2 1.3 2.3 3.1 117. 3.3 74.2 2.6 2.8 73.1 58.2 6.2 2.6 1.1 2. ] | [ 5. 60.3 6.1 153.9 104.9 60.1 7.7 84.6 4. 112.5 67.7 59.6 85.1 61.2 79.8 2. 117.8 144.9 7.7 178.7 2.5 78.1 90.2 159.4 1.4 78.8 7.2 202.9 150.6 72.8 70.2 3.3 108.9 63.8 65.9 6.7 161.7 2. 133.2 2.7 55.2 230.7 2.4 1. 2.1 64.8 3.3 10. 6.6 82.3 6.7 8.3 2. 3.1 1.2 3. 7.7 1.2 3. 7. 164. 3. 3.3 8. 118.2 102. 47.9 6.6 6.2 1.2 177. 2.1 2.3 2.1 1.9 3.3 2.2 1.2 4.6 192.5 69.8 4.1 3. 5.3 73. 150.2 173.1 3.1 2.5 2.9 4. 3.3 1.2]
6.002 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Aachen_MR1_NIST_Complex_T1map.nii.gz | Complex | Shuo Zhang | UKA RWTH Radiologie | Philips | Achieva 3T | 3 | NIST | 130 | 23 | 22.1 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1977.2 2083.1 2104.3 2009.9 2061.2 2003.4 2045.3 2122.1 2049.1 2145.4 2163.9 2072.9 2086.7 1978. 2154.2 2167.4 2010.6 2125.8 2047.3 1962.3 2195.1 1951.8 2053.4 2133.8 2162.4 2095.8 2072.2 2127.8 2171.3 2202.4 2086.2 2022.5 2095.6 2106.7 2160.7 2185.4 2115.1 2120.3 2229.6 2180.4 2140.2 2093.8 2158. 2174. 2226.6 2167.9 2203.6 2059.8 2046.2 2144. 2049.3 2101.7 2127.6 2128.7 2093.1 2142.5 2191. 2214.1 2090.7 2022.3 2052.5 2093.6 2093.8 2100.2 2083.8 2103.8 2178.1 2102.2 2177.2 2175.5 2157.2 2203.9 2124.9 2249.5 2161.7 2164.7 2128. 2170.1 2147.2 2158.6 2186.3 2161.5 2126.2 2048.4 2084. 2127.2] | [1546.8 1564.1 1506.9 1482.5 1540.5 1482.6 1493. 1457. 1521.8 1521.5 1492.6 1473.2 1431.1 1508.6 1500.4 1532.3 1487.5 1563. 1473.5 1447.8 1505.6 1493.7 1481.3 1481.4 1455.7 1495.9 1532.6 1500.3 1444.8 1466.4 1486.4 1495.3 1437. 1510.1 1481.6 1501.9 1503.6 1506.1 1526.3 1476.1 1466.5 1517.7 1474.2 1514.5 1483.6 1467.3 1519.2 1503.3 1500.3 1478.7 1486.4 1503.5 1529.5 1504.2 1486. 1518.5 1503. 1477.6 1558.3 1527.2 1498. 1520.6 1570.2 1485.4 1486. 1503.8 1501.9 1504.3 1551. 1521.9 1508.5 1504.4 1480.8 1506.4 1537.8 1498.3 1521.7 1501.7 1523.1] | [ 977.3 997.9 994.6 1006. 997.2 990.4 970.3 1018.9 1024.4 988.2 996.4 1001.2 1013.8 1021.1 981.5 1037.1 1007.9 1007.9 974.3 988.6 999.1 984. 1009.4 989.3 975.2 1011.1 992.9 998.5 998.1 1004.8 1009.5 1003.4 1007.5 1015.1 1006.3 1007.2 1011. 999.7 999.2 1010.2 1001.3 1008.3 994.5 984.2 1015.5 984.8 994.4 1013.6 1000. 1034.3 997.9 987.9 1001.1 991.6 1005.5 995.1 996.8 996.3 995.3 1000.6 990.2 983.8 999.9 998.2 1025.7 990.9 990.3 1004.2 1013.4 989.6 1002.9 1023.3 1010. 1026.4 998.8 981.2 1006.1 1000. 992.3 1004.3 1021.8 1001.2 987.3 1008.3 992.3 995.5 1007.1 1018.2] | [727.7 715.6 735.8 723.5 708.4 715.2 704. 703.9 718.2 725.3 712.1 711.3 724.6 712.8 711.8 714.2 725.6 731.9 722. 704.9 729.1 716.7 722.8 713.3 706.1 722. 721.3 719.4 719.2 720.3 699.3 706.5 729.7 711.7 711.3 710. 729. 718.5 711.5 717.7 721.9 721.1 714.9 712.1 714. 725.3 702.1 721.9 716.4 714.8 708.8 721.9 724.2 720.3 720.4 728.3 711.7 719.5 707.9 713. 723.2 716.5 709.1 711.4 718. 721.1 722.4 722. 710.4 724.9 710.2 706.5] | [508.9 502. 499.4 500.5 501.8 507.9 484.7 503.4 502.8 510.7 502.3 487.1 501.2 504. 491. 502.4 502.9 504.2 507. 506.7 501.1 509. 502.5 500.4 496. 500.5 492.2 500.5 499.1 500.7 496.6 504.9 509.7 506.8 498.3 496.2 497.5 508.8 501.2 503.8 501. 508.5 514.5 507.9 498.8 497.9 498.3 515.7 508.8 489.6 497.3 498.2 500.6 497.3 500.4 501.5 493.5 509.8 510.9 494.9 494.8 503.4 495.1 499.3 499.2 495.4 485.9 496.6 493.6 505.6 505.4 501.8 503.7 490. 497.1 498.5 503.1 498.7 499. 506.9 504.6 495.6 496.7 492.6 495.8 499.2 490.9] | [324.4 314.4 323.4 325.9 326.1 325. 328.5 332.6 319.7 315.1 319.5 321.3 314.2 319.7 319.9 316.8 317.6 318.8 330.7 319.9 319.4 322.2 315.6 320. 314.3 324.7 313.4 320.5 312.3 319.3 327. 319.8 322. 324.6 330.9 316.3 332. 327.6 319.1 319.7 319.5 323.3 314.3 320.3 324. 321.7 320.3 319.3 328.4 327.5 315.7 318.5 322.7 308. 317.6 316.3 314.3 320.3 321.4 313.1 323.1 322.8 319.7 326.4 317.3 319.5 318.2 314.6 317.5 323.3 312.2 317.4 312.8 320.8 316. 315.5 314.3 308.8 330.6 322.9 315.3 317.8 314.8 329.9 322.7 314.6 322.8 320.1 315.5 319.9] | [248.8 248.9 252.6 246.3 247.1 255.8 251.1 248.4 245.9 244.5 253. 249. 249.3 249.5 245.1 251.2 254. 250.5 250.4 254. 250.1 252.2 248.5 255. 250.9 255.7 248.5 249.4 246.5 249. 248.6 252.1 251.8 249.1 247.6 252.8 248.1 251.7 248. 248. 249.5 248.5 241.3 244.4 246. 245.6 244.4 245.1 248.1 248.8 248.3 250.9 242.3 251.6 243.7 256.5 246.4 242.4 244.1 249.1 243.2 250.7 249. 250.4 243.8 240.9 250. 242.6 244.9 243.6 243.8 240.4 240.8 251.7 238.7 247.8 247.3 246.2 246.3 257.4 245.1 255.8 240.3 252.2 253.6 250.5 246. 241.7] | [164.1 170.4 165.1 167. 166.7 168.5 172.6 167. 171.7 169.7 169.1 171.2 166.7 162.1 169.8 167.9 168.5 173.5 164.2 166.3 169.6 167.1 167.4 166.8 170. 174.2 171.1 168.9 165.3 173.1 168.1 169.9 170.3 167.8 166. 169. 170.2 167. 168.6 172.6 167.5 164.7 168.8 173.9 168.6 175.7 167. 167.9 162.3 166.8 168.6 172.4 167. 173.1 159.5 164.4 169.9 162.9 167.3 166.9 163.1 162.7 160.3 176.2 177.5 164.5 175.3 164.2 163.1 164.9 166.5 166.1 180.6 163.3 168.8 166.3 167.7 162.3 161.9] | [129.6 126.1 126.2 128.5 119.1 120.4 125.7 123.9 128.3 120.1 119.8 127. 126.8 124.4 128.5 128.7 125.6 130.9 121.1 124.2 129.3 129.1 130.2 126.2 125.6 123.3 123.4 124.6 129.7 127.8 122.3 125.1 123.9 125.7 123.9 128.5 120.6 129.7 128.5 128.3 128.9 126.8 118.3 127. 121. 130.8 127.6 123.1 122.1 119.6 112.6 132.4 131.3 135.9 116.2 134.8 133.7 131.3 129.6 126.4 122.1 136.3 109.6 137.3 133.4 111.9 122.9 123.7 126.4 121.9 125.3 115.9 128.6 123.9 123. 129.6 127.1 132. 136.3 133.9 122.7 118.5 105.3 111.2 112.6 109.1 109.9] | [106. 99. 97.3 108.3 99.6 84.2 96.8 83.9 92. 87.5 94.7 95.8 82.3 84.9 91. 81.3 97.5 74.5 85.8 93.2 91.2 86.8 99. 94.7 77.8 84.8 78.6 96.6 95.5 86.6 86.5 76.3 102.2 89.3 80.7 93.6 100.1 107.6 93.1 89.3 60. 89.1 95.5 83.9 75.3 84.9 83.4 88.5 2. 96.5 76.8 103.8 92.8 99.1 78. 99.6 82.3 91. 121.1 3.2 71.3 80.4 95.9 88.6 75.1 83.9 119.6 10. 79.1 78.3 84.1 58.2] | [ 79.1 52.7 76.2 101.5 59.9 74. 73.8 72.4 81.8 76.3 93.4 3.3 85.2 64.6 82.1 3.7 65.2 87.1 68.9 74.6 94.1 10. 90.8 75.3 89. 49.9 1.2 6.3 81.4 85.3 10.3 67.3 10. 79.1 4.5 79.1 67.9 2. 2.5 86.1 87. 2. 85.6 1.8 77.9 1. 1.8 5.7 50.9 86. 6. 10. 66.2 61.5 4.1 92.9 1.6 2.6 1.8 3. 7.2 3.4 91.6 2.8 2.5 93.5 83.9 73. 59.2 1. 97.7 9.5 103.2 2.1 60.1 9.5 2.5 10. 7.9 1.2 1.2 78.5 70.7 94.3 85.8] | [108.1 88. 113.6 105.6 119.8 9.5 155.8 102.3 76.1 101.4 72.8 64.2 10. 69.7 88.6 1. 95.7 95.6 95.1 78.5 86.4 111.5 96.7 77.8 6.3 98. 93. 97.2 80.9 79.9 97.4 77. 3.5 103.4 62.1 1.1 6.9 91.3 94.1 1.8 73.8 3. 67.1 6. 103.7 1.9 5.6 86.6 3.2 112.9 80.3 70.4 74. 99.1 6.3 64.4 10.2 78.6 9.5 73. 90. 3. 4.2 10.6 73. 1. 109.5 5. 4.6 66.4 89.8 1. 3.2 6.6 1.8 10. 72. 3.9 3. 67.1 9.5 54.9 9.5 10. 4.2 5.6 73.9 2.7 3.4 3.3 1.8 71.6 2.1 3.3] | [ 6.4 100.1 111.1 9.5 103. 7. 137.1 131.2 121.4 94.8 98.2 106. 1.2 70. 98.7 119.1 118.2 85.8 5. 67.8 120.3 121.4 1.2 96.3 3. 98.6 161.9 162.8 92.5 79.3 1. 100.4 9.5 123.1 106.6 87.8 3.4 98.6 113.6 50.9 98.5 7.6 129.2 68.6 99.8 61.4 118.6 74.9 55. 3.6 10. 3. 10.6 101.9 1.1 5.1 73.7 9.5 3. 115.6 6. 60.3 4.9 3.1 9.5 2.5 7.8 10. 2. 1.2 2.7 2.5 6. 112.5 10. 57.8 2.8 4. 88.3 1.2 6.2 77.1 1.9 3. ] | [ 10.6 145.2 10. 153.9 119.6 97.1 3.6 127.3 5.6 128.7 122. 131.6 143.9 8.1 98.8 3. 126.1 117.9 73.9 172. 85.3 157.5 124.6 124.4 9.8 90.5 3.2 157. 131. 94. 138.3 6.9 112.9 109. 89.1 10.6 89.2 3.3 106.5 3.7 91.9 150. 2.9 3.4 78.2 101.5 100.2 84.6 2.7 145.9 2.6 1. 10.3 5.6 9.5 8.1 6.4 7.5 76. 9.1 96.4 63.5 71.8 3.7 151.9 69.2 9. 3.9 2. 5.3 114.8 6.6 9.8 10. 63.5 2. 9.8 10. 2. 159.9 122.5 3.4 94.3 4.3 91.4 152.1 143.9 7.2 6.3 1.2 1.6 9.5 7.8]
6.003 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Berlin_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Christian Stehning | Deutsches Herzzentrum Berlin | Philips | Ingenia CDAS 3T | 3 | NIST | 130 | 23 | 23.2 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1947.1 1976.3 1991.6 1976.1 1988.3 2009.1 2055.6 2016.7 2000. 1978.4 2001.4 2021.5 1988.2 2040. 2075.8 2028.5 1994. 1969.6 1978.1 2014.7 2057.5 2146.8 2027.5 2030.7 2058.7 2042.7 2000.2 2002.4 1985.7 1991.3 2071. 2062.4 2151.2 2053.3 2033.2 2048.2 2046.7 2050.6 2046.3 2061.7 2057.5 2083.1 2056.9 2079.3 2087.8 2076.6 2059.1 2041.4 2024.9 2032.8 2049.1 2054. 2048. 2041.6 2034.8 2069.5 2096.1 2073.3 2028.5 2040.7 2050. 2019.9 2042.4 2021.6 2013.8 2090.2 2122.8 2059.1 2053.6 2020.6 2053.5 2084.8 2019.1 2018.4 2087.1 2081.8 2061.9 2020. 1996.8 2011.1 1981. 2066.1 2064.8 2043.1 2008.7 2017. ] | [1421.9 1418.7 1435.1 1453.6 1468.1 1469.2 1436.1 1500.3 1509.5 1492.7 1495. 1484.6 1477.1 1503.7 1479. 1527.9 1475.2 1482.9 1489.1 1469.7 1464. 1466.8 1470.8 1455.2 1509. 1485.9 1480.1 1461.2 1450. 1447.9 1462.8 1478.3 1473.6 1458.4 1482. 1475.2 1496.9 1477. 1443.3 1439.7 1462.4 1453.8 1478.3 1481.1 1465.4 1425. 1466.5 1474.8 1464.9 1468.8 1469.8 1458.9 1459.8 1470.9 1434.6 1463.7 1449.7 1460.1 1468.8 1455.3 1455.1 1448.6 1447.5 1421.3 1451.2 1459.5 1474.9 1452.2 1440.5 1456.5 1459.9 1452.8 1468.8 1478.7 1473.2 1472.3 1470.1 1458.1 1440. 1485.4 1458. 1464.2 1471.1 1450.3] | [1040.8 1030. 1000.5 996.8 982.3 1010.6 974. 978.8 986.5 981.2 977.9 968.2 954.3 933.9 965.5 996.7 989.4 977.5 984. 993.4 992.8 995.2 982.2 963.9 976.9 1000.1 981.5 968.8 974.4 975.2 963.6 974.4 976.7 969.9 977.8 979.4 982.4 974.8 973.7 978.3 968.8 964.2 973.1 963. 982.7 978.4 968.8 965.4 971.4 970.8 976.9 971.2 970.3 962.5 975.9 964.6 964.2 961.3 968.4 968.8 972.9 963.5 975.7 966.1 978.3 961.1 959.9 961.2 960.2 963.3 968.9 964.6 967.1 964.9 968.7 968.1 965.7 966.5 964.5 973.9 965.4 960.1 973. 962.3 952.6 969.1 969.4 971.6] | [754.9 721.5 730.9 719.1 712.9 727. 712.9 711. 695.2 710.3 715.1 709.7 715. 711.7 694.3 709.5 726.6 712.4 719.4 715.3 706.5 716.6 718.5 719.3 686.9 713.5 699.6 708.3 701.5 699.6 705.1 708.8 704.9 702.5 702.5 704.9 703.1 705.4 703.1 701.5 709.3 705.8 708.7 712.7 716.5 707. 703.2 696.1 700. 707.3 705.1 708. 704.8 705. 710.4 704.9 703.9 700.1 700.2 702.4 696.6 706.2 705.8 705.3 703.9 705.5 700.3 702.1 702.9 698.9 697.8 700.5 704.2 708.9 702.3 705.3 702.8 707. 711.4 703.6 707.6 702.6 703.9 697.1 700. 701.7 702.2 703.7] | [488.4 492.1 495.6 491.3 481. 486.4 473.9 500.1 498.6 501.7 500.9 500.9 502.2 496. 476.8 500. 490.4 492.8 495.4 497.2 498.5 497.6 497.5 490.3 485.3 487. 489.3 492.3 499.2 500.3 498.6 506.2 504.5 492.7 490.2 486.3 492.7 492.2 492.7 498.7 496.4 501.8 497. 497.9 496. 489.1 488.3 499.3 499.4 502.2 498.1 500.5 498.7 495.2 490.3 488.7 491.7 492.8 496.7 497.3 496.3 497. 495.4 495.5 487.1 485.7 493. 495. 494.6 494.1 495. 493.8 495.1 496.4 486.7 495.2 490.6 496.3 494.9 494. 490.1 490.3 495.2 498.6 496.2 501.9 496.6 492.2] | [315.9 319. 322.9 319.5 315.4 321.6 320.4 319. 317.4 313.2 310.1 307.2 311.9 313.8 313. 321.1 319.1 319.9 318. 310.2 311. 314.7 319.5 313.9 318. 323.7 320.6 324.2 322.6 317.4 317.7 313.4 312.2 318.4 316.5 318.7 320.7 324.5 323.8 320.9 325. 320.8 317.5 308.9 321. 317.7 319.8 320.7 324.2 326.3 323.2 324.5 323.2 319.7 310.8 322.3 316.6 318.1 319.5 320.9 321.5 321.3 325. 322.8 320.1 323.8 316.3 318.4 320.2 317.6 322.5 323.8 322.1 319.6 316.3 318.5 319.9 319.2 322.6 322.8 318.1 316.5 317.8 320.4 322.5 322.5] | [246.1 249.1 253.5 251.3 252.9 251.7 256.3 250.1 245.3 246.3 247.8 247.6 248.2 249.8 257.5 243.1 243.9 246.3 245.3 247.8 246.5 243.7 244.3 245.8 249.2 243.9 247.5 247.7 246.4 243.7 246.7 248.3 249.4 245.3 249.6 246.4 247.5 243.8 248.1 243.9 244.8 246.8 251.1 249. 251.4 246.2 246.6 242.8 248. 248.6 246.2 246.7 249.6 246.2 250.5 246. 246.4 243.5 245.3 248.1 246.2 247. 249.7 246.2 250.5 245.2 247.3 246. 245.6 246.3 247.3 246. 250. 247.1 243.9 247.2 245.9 244.5 243.2 247.9 245.6 246.7 248. 245.9 243.7 243.2 243.8 245.9] | [161.4 161.5 162.8 161.5 164.2 164.7 160.8 163.4 169.7 167.2 165.7 167.5 167.9 167. 167. 162.7 166.4 167. 165. 165.7 165.6 165.5 167.5 168.2 170.6 162.4 165.9 168.4 164.6 166.3 166.1 169.6 167. 169. 166.4 164.7 164.4 167.5 165.2 166.6 165.9 165. 163.9 167.4 169.9 167.7 164.4 165. 166.3 165.8 164.9 165.9 165.3 166.4 165.9 165. 165.6 167.6 167.3 165.4 166.3 167.6 168.4 164. 353.6 167.4 166.8 166.2 166.2 168.9 169.2 354.7 162.7 335.8 165.3 166.1 168.2 166.5 323.2 164.3 168.7 166. 164.6 163.4] | [111. 159.8 118.7 117.5 118.5 120. 117.4 169.9 121.6 122.3 116.1 117.1 120. 117.7 171. 112.4 120.7 119. 122.3 119.8 117.4 154.8 118. 118.4 128.3 114.8 163.8 118.6 121.2 123.1 120.7 153. 118.2 120.8 175. 118.1 161.7 121.5 122.6 162.6 123.8 120.4 156.9 122. 177.2 117. 120.9 118.9 121.3 123.7 161.3 158.3 156.6 122.7 120. 120.6 173.6 155.6 158.7 165.1 159.9 121. 121. 169.7 168.6 120. 118.3 170.1 120. 166.1 125.1 122.8 120.8 111.8 212.5 168.2 119.1 116.4 112.8 114.7 115.6 118.1 132.6 166.7 177. 171.7 169. 125.1] | [109.8 98.7 103.1 101.6 100.7 105.5 122.3 63.9 62.4 70.6 78.9 74.9 74.2 89.1 9.5 77.2 63.8 106.8 84.6 89.2 84.8 80.8 82.4 71.8 85.2 66.7 70.2 74.9 83.6 76.7 78.9 69.5 75.5 74.1 85.1 2.7 73.9 84.1 83. 92.2 93.8 83.7 72.3 80.4 78.2 68.5 84.3 90.9 77.8 82.8 85.8 81.8 76.7 86.1 92.1 66. 81.8 91.4 79.5 81.7 82.3 82. 85. 87.5 82.2 66. 73.4 59.5 82.5 80.6 84.9 87.4 84.3 84. 2.7 60.7 106.8 79.5 93.7 90.1 90.7 91.7 3.6 83.6 85.3 87.1 88.7 76.6] | [ 93.9 87.7 81. 121.4 2.7 6.1 4.7 2.7 6.6 9.5 93.2 1.1 1. 1.1 2.3 5.6 2.6 1. 2. 2. 75.8 65.7 67. 55.6 3.4 9.5 5. 1. 2.5 2.8 62.9 2.6 9.5 4. 5. 2.6 4.1 3.2 3.2 3.2 1.3 1.2 2. 2.5 2.7 1.3 1.8 2.4 1.2 2.6 2.4 2.7 9.5 9.5 56.1 1.3 3.6 2.6 2.1 4. 2.7 1.3 1.8 68.5 3.3 64.7 1.3 3.2 3.3 1.3 1.2 1.3 59.8 1.3 69.5 1.2 9.5 2.1 2.8 3.2 3.3 97.3 6. 3.5 2.5 2. 1.2 1.2 1.8 2.8 3. 2. 7. 2.8] | [ 8.2 3.3 1.1 2.4 2.7 9.1 64.5 73. 88.5 70.8 4.6 8. 47.9 97.6 1.1 1.6 3.2 2. 2.4 2.7 2.7 1.2 4.6 4. 1.1 2.5 4. 6.4 1.8 3. 1. 2.8 5.3 2. 9.5 3. 3. 3. 1.2 8. 2. 1.3 3.2 2. 2. 2.7 2.4 2. 1.4 9.1 1.6 9.1 9.5 1.2 3.3 1.1 2.5 3.1 1.3 2. 4.6 2. 3.2 3. 5.1 2. 2. 9.5 5.5 1. 2.5 3.1 2.4 2. 2.8 1. 1.8 5. 2.7 2.5 1.1 2.8 2. 9.5 6. 2.5 1.9 7.5 4.5 3.4 2. 2.7 9. 3.6] | [74.1 4.8 2.8 2. 1.2 1.8 4.4 9.5 95.6 52.6 1.9 6.3 3.4 90.1 85.2 5. 53.8 3.6 3.5 8. 4.3 5. 72.8 1.1 3.3 58.3 4.3 1.6 2. 7.6 2. 1.9 3.3 2.9 5. 57.9 4.1 4.3 3.3 9.1 2.1 1.3 3.2 6.1 2. 2.8 2.4 5. 1.3 2.5 9.1 9.5 1.9 2.4 9.5 1.8 3.3 9.5 2. 3.6 3.4 8.2 1.8 3.3 3. 3.3 2.7 1.7 9.1 4. 5.3 3. 6. 1.2 2. 4. 5.3 1.3 2. 8.5 3.1 9.1 5. 2.8 5.3 9.1 2.7 3. 2.5 6. 5. 3.4 2.6 4. ] | [ 1.2 1. 2.1 202.1 123.8 91.8 66.7 173.3 185.8 312. 8. 7. 1.3 1.8 6.4 4. 52.3 43.6 338.3 2.5 2.8 2.5 2.1 2.4 79.6 66.5 65.4 2.8 186.5 58.7 3.3 1. 2.4 9.5 3.1 9.5 3.7 1.3 1.8 61.9 2. 2.7 61.7 5.7 1.1 10. 2.7 2.7 1.3 3. 2.1 8. 1.2 4. 10. 1.8 2.1 2.7 1.3 2. 3.3 2. 2. 1.2 2. 3. 10. 52.6 1.2 2.7 4.7 2.6 5. 5. 1.8 2.8 2. 2.3 1.1 60.6 3.1 3.4 7. 1.2 4. 2.7 44.9 72.4 8. 1.3 4.1 9.5 3.1 1.2]
6.004 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Berlin_MR1_NIST_Complex_T1map.nii.gz | Complex | Christian Stehning | Deutsches Herzzentrum Berlin | Philips | Ingenia CDAS 3T | 3 | NIST | 130 | 23 | 23.2 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1976. 2030.9 2062.6 2047.3 2040.9 2067.6 2140. 2072.4 2079.2 2027.2 2065.8 2065.1 1948.4 2042.4 2118.3 2042.5 2029.8 1983.7 1998.4 2026.2 2037.8 2105.4 1981.7 2003. 2044.1 2061.3 1995. 2014.5 1985.2 1968.2 2086.8 2001.6 2129.4 2003.7 2001.1 2023.2 2020.7 2067.4 2070.4 2094.3 2067.8 2108.6 2028.7 2037.9 2074.6 2033. 2016.3 1999. 2024.1 2027.3 2049.9 2070. 2030.2 2013.6 1987. 1973.7 2033.4 2033. 1965.3 2016.5 2061.2 2016.8 2047.6 1997.3 1958.4 2034.7 2070.2 1960.2 2014.3 1992. 2051.3 2133.2 2003.7 2001.1 2019.3 2026.6 2013.1 1983. 1956.4 2018.8 1951.4 2016.9 2032.2 2037.4 2003.6 2010.3] | [1422.5 1424.6 1449.9 1463.3 1455.1 1460.3 1414. 1488.2 1515.4 1494.5 1505. 1500.4 1456.8 1502.3 1439.1 1492.2 1465.7 1475.2 1506.1 1460.9 1459.8 1467.1 1459.8 1432.7 1472.7 1458.1 1455. 1450. 1439.9 1443.7 1476.2 1492.8 1470.2 1440. 1446.9 1457.2 1517.8 1493.5 1425.5 1432.7 1465.1 1452. 1479.7 1488.6 1464.8 1436.9 1465. 1494.2 1473.9 1489.3 1490.6 1465.2 1463.2 1484.9 1455.6 1497.8 1450.9 1466.9 1487. 1461.7 1462.7 1451.5 1439.1 1455.1 1453.3 1468.8 1489.4 1456.3 1441. 1472.7 1486.8 1483. 1474.2 1489.6 1476.9 1489.5 1484.6 1474.3 1446. 1489.9 1433.8 1467.9 1491.1 1463.1] | [1024.5 1018.9 999.4 975.9 979.4 1013.4 965.8 971.4 980.7 983.1 977.4 972.4 956.4 937.9 955.1 986.3 985.2 973.9 986.1 1010.6 1010.7 1006.1 998.1 963.8 967.9 1005.5 978.5 966. 983.8 987.9 961.3 989.5 994.8 970.7 973.4 973.8 998.7 982.5 986.4 996.4 982.3 972.7 979.4 961.9 992.3 987.9 972. 970.8 989.6 985.6 1000.3 983.5 976.6 963.7 978.6 965.3 977.4 969.6 979.4 985.2 984.6 970.5 993.3 973.3 990.2 957.3 970. 970.7 969.5 976.3 982.8 971.8 974.8 976.6 980.2 981.2 977.1 977.6 972.9 985.6 973.6 962.4 985.2 966.7 949.9 981.4 977.5 974.4] | [751. 720.2 725.1 715.7 709.7 724.9 715.7 713.1 690.7 698.9 705. 695.9 709.5 704.3 700.4 710.1 728. 707.1 726.7 716.3 697.2 719.3 718.9 720.6 687.2 721.5 701.6 705.2 697.9 692.2 702.2 709. 699.3 697.5 703.2 704.7 703. 711.2 704.7 702.7 715. 705.4 710.2 716.8 719.3 708.2 703.6 692.2 701.8 715.7 709.6 713.5 703.2 706.6 712.7 700. 706.6 703.5 702.3 705.3 696.9 709. 708.6 711.6 704.1 712.1 699.6 703.6 709.7 700.8 703.2 702.5 707.4 715. 698. 711.3 705.6 714.1 720.7 709. 713.1 705.4 703. 695.3 703.4 704.7 704.9 711.9] | [484.6 490.4 495.6 493.3 482.7 490.6 473. 506.6 503.7 506.5 499.6 502.3 507.6 498.8 478.3 506.4 489.4 488.8 489.1 490.5 493.2 489.8 496. 484.7 489.4 491. 486.3 489.5 496.9 493. 487.9 505. 502.3 490.6 497.6 487.7 495.7 488.9 483.2 488.6 486.1 495.8 493.9 502. 500.9 487.8 483.5 499.6 494.1 495.7 490. 494. 496.1 493.4 495.6 490.3 490.3 489.1 490.5 489.4 487.8 492. 490.5 497.9 495.9 486.2 495.4 491.6 488.9 486.8 488.6 489. 492.4 500. 491.5 497.7 486.4 494.1 491.5 489.4 486. 487.6 496.3 494. 494.4 504.6 494.2 489.6] | [314.3 319.2 320.7 323.1 320.6 324.9 320.6 313.7 312.4 310.2 309.7 309.1 313.3 312.3 302.7 320.3 311.9 316.4 316.8 305.7 313.1 314.8 318.7 311.4 318.2 322.2 311.9 317.4 317.8 311.1 317.3 315.3 311.1 317.3 312.2 314.4 312. 317.1 313.2 310.4 319.8 315.5 313.4 309.7 319.9 313.3 317.5 313.9 316.3 317.6 313.3 317.3 316.8 317. 310.6 320. 311.6 314.5 315.2 315.3 314.5 313.4 320.2 318.5 317.5 324.2 310.4 315.6 318.6 311. 317.6 319.3 316.7 314.1 316.4 319.9 317. 314.4 318.8 319.7 310.3 309.8 315.6 317.8 317.3 316.1] | [245.1 248.1 255.1 248.1 252.6 250.5 256. 245.5 243.8 240.6 247.1 242.2 248. 248.2 255.4 237.2 242.9 246.4 243.4 249.5 246.8 242.7 242. 242.5 247.2 243. 248.7 248. 246.2 240.7 245.7 245.5 250.8 242.3 247.7 245.3 247.6 240.1 248.1 241.9 242.4 244.6 250.6 246. 248.8 245.3 245.7 239.4 248.1 248.6 244.6 244.7 248.2 241.7 246.1 245.4 246.2 240.8 243.6 248.1 244.9 246.5 249.3 242. 247.8 242.1 248. 244.4 243.9 244.8 247.6 244. 250.5 244.3 245. 244.5 244.8 242.8 242.2 249.3 245.3 245.7 248.2 247. 243.8 244.1 245. 248.9] | [160.5 159.6 160.3 157.1 163.5 162.1 160. 162.4 170.8 168.1 163.8 168.2 168.1 165.4 167.2 163. 167.7 165. 162.5 164.6 161.9 164.5 161.5 168.2 171.9 159.7 164.3 170.5 163.2 166.3 165.4 171. 167.2 170.2 165. 163.5 164.5 165.6 163.3 165.6 165. 163.2 159.4 166.4 170.8 167.3 164. 163.4 166.1 164.9 164.1 164.5 165.7 164.9 165.3 162. 164.7 167.7 167.2 164.4 166.2 166.9 169.5 163.8 171.4 167.1 167.1 164.9 164.2 170.1 167.2 168.5 157.8 164.7 163.7 165.4 171. 165.5 164.7 163.6 168.3 166.9 163.4 161.8] | [110.8 119.6 118. 117.4 119.1 121. 118.8 124.4 122.2 124.6 113.2 115.9 117.8 118.1 125. 112.3 120.4 117.4 122.6 118.4 114.4 120.2 116. 117.8 129.2 116.6 122.7 115.4 118.7 122.5 119.8 114.7 115.7 123.7 127. 119.4 124.4 121.8 122. 122.8 122.6 118.3 117.6 123.4 127.7 117.5 118.5 117.3 119.8 123.2 121.7 118. 119.3 122.1 120.4 121.8 127.9 114.3 120.3 123. 118. 118.4 118.8 124.3 124.5 119.1 113.9 125.8 118.4 126.1 126.1 122.6 120.9 106.8 144.3 123.1 116.5 115.1 111.6 114.9 115.8 119.4 131.3 123.1 128.1 126.5 125.7 126.1] | [100.2 85.7 94. 89.1 87.3 94.3 112.8 58.2 5. 49.3 78.2 59.3 2.6 79.4 70.4 79.7 69.3 96. 76.1 86.1 83.3 78.2 83.6 60.5 92.3 81.4 71.3 80.7 72.3 69.6 53.5 53.6 57.1 82.4 90.2 5.8 69.2 86.1 88.6 88.2 89.4 78.2 76.2 90.5 88.5 79.5 93.1 95.3 58. 64.4 87. 77.9 52.5 92.2 101.1 71. 83. 101.4 81.7 85.7 79.4 80.8 85.7 94.7 90.5 83.1 84.6 61.7 79.1 78.5 80.9 91.3 79.2 86.8 7. 63.8 95.9 75.7 94.2 87.8 93.5 82. 1.2 76.6 85.1 87.4 82.6 61.6] | [ 90.5 85.1 81. 111.5 9.5 9.5 1.2 10. 1.4 9.5 93.1 6.6 1.8 2.5 61.1 3.2 3.7 1.8 73.6 8.5 74. 69.5 91.1 3.5 2. 1.1 2. 2.6 61.2 5.3 81.1 3.4 4.2 3.6 1. 11. 1.3 10.6 56.6 3.6 1.2 3.4 4.1 79.7 1.3 2.8 66.3 7.2 2.5 1.3 10. 9.5 67.2 1.2 82.4 3.6 10.9 3.3 2.5 1.1 59.9 1. 9.5 92.2 2. 88.7 9.5 2.4 3.7 9.5 6.3 59.8 73.4 9.5 97.8 4.5 51.4 1.3 5.3 9.8 1.6 101.1 10. 1.6 74.3 54.4 62.7 54.5 64.9 1.4 52. 5. 8.9 1.2] | [ 8. 5. 2.2 10. 3.1 3.1 9.5 10.9 73.6 81.8 72.8 62. 81.8 78.9 9.5 1.5 5.3 1.1 1.8 3.9 1. 5.3 9.5 6.9 10. 3. 1. 71.8 2.7 9.5 2.6 2.7 6.2 4.6 5.6 9.5 5.3 44.7 2.6 2.8 10.6 102.5 5. 10. 9.1 2.6 2. 2.8 83.8 9.5 3. 4.9 1. 2.7 10. 2.1 1. 72.6 3. 2.7 1. 9.8 3.4 9.5 9.5 4.9 1.1 5.5 1.3 10.3 5. 1.7 2.5 4.9 3.6 4.1 1.1 3. 9.5 9.5 1.2 6.3 1. 38.9 7.3 2.1 1.2 1.5 3.3 1.4 3. 1.5 3. 68.5] | [ 79.3 5.6 2. 8.3 1.2 5.9 3. 2. 97.3 102.5 6.4 2.5 3.4 110.7 129.9 2. 82.1 1.3 3.7 1.3 89.4 125.8 112.3 3.9 2.8 80.3 10. 3. 1.4 5. 6.6 77. 2.7 10. 7. 124.2 9.5 3.3 10. 1.1 3.3 1.8 2. 94.6 92. 7.7 4. 2.1 2.8 3. 10.6 2.8 5. 1. 2.9 2. 8.8 114.4 3. 1. 2.5 3.3 3.2 3. 1. 3.3 2.7 105. 10. 2. 1.4 4.6 6.7 9.5 1. 1.4 4. 3.4 1.9 6. 2.5 8.7 2.5 95.1 2. 2. 1.4 8. 9.5 10. 1.2 10. 1.3 1. ] | [ 10. 10. 10. 133. 133.9 7.9 100. 180.9 159.6 177.6 5.9 1.5 6. 1.5 6. 7.9 2.5 5.7 301.5 3.8 80. 6.6 3.5 2.2 236.6 142. 178.9 10.2 144.8 79.8 2. 2. 8. 2. 2.4 89.1 10.1 3.4 4.5 103. 1. 3.3 170.9 10.9 96.8 4. 4. 8. 2. 6.5 124.2 108.7 10. 32.2 10. 1.2 1. 8. 3.9 134.1 2. 10. 5.2 2. 117.5 1.5 9. 146.2 10. 177.8 2. 153.8 7.1 1. 2. 4. 1.5 2.4 6.2 189.6 2.7 157.8 3.6 8. 9.3 6.9 112.3 162.8 10.3 87.8 74. 2.6 2. 10.7]
6.005 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Bonn_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Burkhard Maedler | University Hospital Bonn Radiology | Philips | Ingenia CDAS 3T | 3 | NIST | 130 | 23 | 22 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [2047.6 2071.9 1954.7 1975.9 2106.9 2113.1 2032.4 2017.9 2003.1 1968. 2005.6 2026.4 2060. 2088. 2062.6 2006.4 1990.9 2034.6 1954.9 1942.7 2025.3 1984.6 2141.3 2086.7 2009.9 2020. 2039.5 2003.7 2000.9 1994.3 2015.4 2056.8 1950.2 2081.7 1994.8 1919.1 1986.1 2012.5 1992.1 1975.3 2007.9 2036.1 2028. 1970.5 2053.1 2006.5 1952.3 2006.6 2018.3 1980.8 1901.4 1980.4 2084.2 2033.3 2036. 1975.8 2075.3 2056.3 2064.3 1977.8 1984.1 1946.1 1948.9 1967.8 1984. 2138.4 1966. 2097.3 2063.7 2073.1 2014.7 1971.5 1990.5 1974.8 1944.6 1942.1 2021.6 2123.3 2029.9 2072.6 2120.8 2037.9 2033. 2034.1 2024. 2009.2 2105.7 2037.5 2001.2 1997. 1998.6 1982.6 2007. 1978.2 1989.9 2009. 2004.3] | [1434.2 1441.8 1414.4 1432.5 1438.4 1410.4 1405.3 1477.9 1425.7 1413.6 1426.8 1432.2 1406.6 1414.6 1393.2 1467.4 1456.5 1401.9 1405.2 1406. 1423.2 1438.7 1413.5 1415.1 1399.9 1466.3 1446.4 1417.1 1411.4 1394.7 1406.4 1416.1 1410.3 1411.1 1359.3 1451.2 1459.2 1442.6 1432.1 1426.8 1410.3 1391.5 1405.4 1403.5 1342.5 1442.4 1443.2 1455.4 1451.7 1430.4 1425.7 1431. 1371.5 1411.1 1357. 1410.3 1425.6 1430.5 1442. 1440.1 1445.1 1439.9 1387.4 1441.4 1411.5 1428.9 1423.9 1410.6 1425.2 1430.7 1440.6 1446.6 1449.4 1432.2 1459.1 1446.8 1434.2 1427.1 1414.8 1433.4 1429.1 1439.6 1443.4 1434.2 1446.8 1429.2 1391.6 1398. ] | [1010.9 984.4 959.1 956.1 950.4 989.5 963.6 963.8 963.4 953.3 945. 951.1 998.8 972.2 949. 971.2 966.8 964.6 968.8 958.8 965.2 980.3 966.9 961.5 968.5 966.8 966.4 965. 949.6 958.5 959.8 961. 957.8 966.9 946.4 955.3 964.3 955.5 944.7 964.9 961.8 968. 967.8 943.8 945.7 964.3 952.1 954.7 954.9 940.7 950.4 954.4 965.5 969.4 951.5 950. 966.8 973.5 949.3 949.6 954.1 960.6 952.3 954.5 954. 960.4 961.8 976.3 962.6 955.7 964.9 955.4 965.1 953.9 958.6 965.5 928.8 972.2 967.7 970.3 957.5 941.3] | [699.4 694.4 706.2 714.1 693.7 700. 706.6 706.8 708.3 706.8 697.3 701. 702.8 701.2 700.7 704.6 702.4 695. 691.3 684.2 710.1 691.9 705.1 704.8 696.1 698.5 690. 683.6 694.8 675.5 707.2 690.5 702.7 699.2 694.7 695.9 688.5 690.2 685.1 698.4 703.9 690.8 702. 688.5 693.3 691.8 686.7 695. 701.4 696.1 708.7 694. 697.4 702.4 706.8 691.6 689.3 698.9 701.7 688.3 703.4 688.7 697.3 699.9 712. 702.9 697.9 703.6 705.5 704.3 694.8 694.7 690.8 698.6 702.1 693.9 697.6 700.8 697.3 696.5 706.1 708.3 699.8] | [475.5 498.4 498.4 494.4 493.4 483.4 473.4 493.9 487.5 488.7 490.1 496.5 488.1 485.9 480.2 491. 480.9 490.3 492.6 488.8 491.8 488.2 494. 500. 483.1 494.1 488.8 494.2 490.1 497.8 489.6 486.5 486.3 491. 476.2 488.7 483.9 489.7 497.3 502. 497.7 498.2 491.3 485.8 466.2 484.6 486.2 491.7 490.5 493.6 496.8 492.5 490.6 488.7 468.8 488.5 488.4 486.3 489.4 498.3 497.4 494.3 486.2 491.5 477.6 488. 479.5 483.1 485.3 486.3 491. 490.7 487.6 485.5 486. 480.6 482.2 488.4 486.6 487.6 488.3 488.8 485.4 481. 487.8 486.1 486.3 487. ] | [315.5 306.9 309.7 309.7 306.6 314.9 305.3 308.2 315.6 313.9 307.7 309.6 309. 305.7 308.4 308.7 310.9 315.8 312.8 308.7 304.2 307.6 317.2 305.9 308.7 311.2 312.3 316. 313.8 313.3 310.9 310.4 296.8 316.2 305. 305.3 314.4 319.3 317.2 314. 319.6 318.7 317.5 302.7 315. 309.6 307.3 310.4 311.4 309.9 307.7 312.9 315. 320. 307.6 320.7 306.4 307.2 311.1 310.9 311.2 311.9 313.6 317.2 317.9 303.9 314. 305. 304.9 313.7 314.9 317. 320.2 320.1 315.9 316.1 310.3 311.2 311. 314.6 316.8 316.9 318.8 314.2 310.2 323.9 315.2 314.9 313.3 314.8 318.8 314.7 322.5 323.8 326.8 323.5 323.2] | [239.2 236. 235.2 235.4 239.8 237.2 238.8 237.5 235.8 238.5 240.9 237. 236.1 247.8 242.8 240.9 234.2 240.1 239.8 236.8 241.5 243. 240.7 237.7 241.4 239.3 238.7 239.7 236.6 240.1 241.1 240. 239. 242. 238.8 234.1 241.5 238.4 238. 238.4 241.2 242.5 244.3 239.7 239.2 241.6 242.1 242.3 237.4 241.5 238.5 241.8 243.4 248.2 243.3 243.8 239. 236. 240. 242.6 236.5 238.6 244.2 247. 241.8 241.2 238.9 242.3 241. 237.1 236.4 239.3 241.1 238. 241.4 234.7 239.5 236.3 236.8 237.6 240.4 238.4] | [160.1 155.1 154.8 160.1 157.2 157.8 164.2 162.6 157.1 160.1 162. 162.2 157. 270. 155.9 158.6 159.7 157.1 159.2 156.1 158.4 160.5 161.2 157.6 154.4 156.6 160.2 158.5 157.5 157.7 156.1 159.6 162.5 153.8 155.4 155.9 157.7 159.1 158. 156.4 155.5 159.4 162.4 156.5 157.5 156.3 157. 158.6 158.4 159.5 158.5 161.1 165. 160.2 153.5 260.8 160.7 157.5 158.2 158.3 158. 161.3 165.4 161.6 148.1 154.7 163.3 161.4 160.4 159.5 159.2 166.1 293.5 153.8 147.6 161.3 159.4 157.6 158.8 155.6 159.2 155.2 165.1 163.8 160.5 161.5 164. 165.6] | [ 94.6 118.8 107.8 105.8 106.9 111.1 102.1 142.7 113.1 110.2 106.7 106.7 110.4 113.8 149.8 112.5 131.5 108.7 114.1 117.1 114.2 112. 112.7 112.3 163.4 114.9 146.7 118.1 120.4 117.3 118.4 118.5 114.5 109.8 170.3 147.1 157.8 121.4 117.3 119.7 113.3 115.3 118.1 112.6 177.7 125.6 148.9 118.1 113.6 117.4 118.5 117.4 116.6 111.4 110.1 116.8 150.7 118.5 113.3 112.1 114.7 118.4 113.4 142. 112.5 152.2 110.3 159.4 116.2 112.8 114.8 119.7 129.9 108.1 179. 115. 118.6 108.8 102.2 105.3 103.2 112.6 130.6 129.3 122.5 120.3 118.1 119.6] | [ 2.8 63.3 10. 3.2 61.3 74.9 45.7 8.4 3.1 66.2 72.6 1.3 100.8 9.5 68.8 55.8 81.8 75.3 74.5 1.3 72.3 7. 61.4 3. 83.6 74. 62.9 66.8 59.7 2. 1.8 1.2 53.7 62.9 85.5 50.1 4.6 2.8 9.5 3. 64.6 2. 58.9 70.1 75.2 2. 3.3 53.4 58.1 9.5 91.8 43.6 67.1 69.7 72.9 9.5 62.1 70.8 76.5 1.2 90.4 53.9 67.8 1.1 2.1 7. 79.5 67.9 80.6 89.5 4.2 9.5 1. 63.1 81.4 73.9 73.3 110.5 2.6 1.2 46.3 3.3 9.5] | [ 2.5 1.9 3.3 3.2 2.4 1.9 2.6 2.7 3. 1. 1.2 3.3 2.1 1.3 3. 5. 2.7 6.8 2.6 4. 4.1 2. 2.7 2.5 1.2 1.1 1.8 9.5 2.6 58.2 2.1 4.6 1.4 2.7 6.3 3. 4.9 1.9 3.6 2.9 1.8 2.5 1.3 4.1 1.8 2.8 2.5 9.2 2.7 1.2 1.1 79.6 1.2 2.9 3.7 2.6 1.9 1.3 1.2 2. 9.5 1. 2. 2.4 2.5 2.7 3.3 1.3 9.5 9.5 3.7 5. 78.7 2.8 2.5 2.6 1.3 1.1 3.4 1.2 3.7 1.3 95.7 2.4 1.8 2.1 1.2 1.2 1.4 1.8 84.5 10. 3. 2.4] | [ 9.5 2.7 1.8 4. 3.3 1.2 2.1 90.3 2.4 1.1 1.2 1.7 3. 1. 79.3 2. 2.5 3.3 6.9 2.7 9.5 1.2 93.2 3. 1.2 3.2 2. 1.9 2.8 5. 9.5 1.2 1.8 2.4 3.1 1.3 2. 2.6 1.8 2. 6.3 2.7 1.8 3.4 1.1 7. 4. 2. 2.1 2.4 3.2 1.6 1.6 5. 9.5 9.5 1.2 1.8 9.5 4.1 2.4 9.5 1.3 3.1 1.8 5. 4.7 2.7 2. 2.1 2.7 2. 7.7 1. 3.4 2.5 4.1 1.9 2.8 1.8 3.3 6.3 1.2 1.9 1.3 2. 6.1 3. 1.2 1.8 1.2 3.3 2.5 2.6] | [ 3. 2. 4.1 9.5 10. 1.3 5. 1.1 2. 1.8 1.1 6.1 1. 73.7 4.3 66.9 3.6 3.2 4. 3.2 1.8 1.8 2.8 2.1 1.1 2.1 6.8 4. 2. 7. 4. 2.5 1.2 2.9 3.6 5.6 2.9 1.1 1.9 4.9 3. 5. 2.7 2.1 2.4 4. 3.6 1.1 8.1 2.7 9.1 6.5 9. 1.2 7. 1.2 3.3 2.5 1.9 1.9 9.5 9.5 3.3 3.3 3. 3.3 1.2 4.1 2. 9.1 4.9 2.8 2.4 1.2 4.3 4.9 1.9 2.3 3.8 3.6 1.1 10. 2.7 4.8 1.2 3.3 7.2 6.2 4. 7.6 1.4 2.4 1.8 3.6] | [70.7 68.4 64.9 6. 3.3 3.1 3.4 1.8 3.1 5. 2.7 7.3 6.7 8. 1. 1.3 1. 1.2 9.5 9. 1.2 1. 4. 2. 10. 1.3 5. 3.4 5.5 3. 4.9 2.4 1.3 5.4 2. 1.6 1. 9.5 1.2 1.5 6.1 7. 1.1 2.7 1.2 4. 1.2 2.8 7.8 3. 6. 6.3 1.2 1.7 1.2 4.1 9.1 1.3 3.6 2. 1.2 2.8 3.3 2.7 1.4 3.4 1.1 1.7 2.6 1.8 2.2 1. 1.2 2. 3.2 4.5 1. 9.5 4.2 65.6 44.7 3.3 2.7 2.7 9. 66.4 2.8 1.3 1.8 2.8 4. 2.9]
6.006 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Bonn_MR1_NIST_Complex_T1map.nii.gz | Complex | Burkhard Maedler | University Hospital Bonn Radiology | Philips | Ingenia CDAS 3T | 3 | NIST | 130 | 23 | 22 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [2013.8 2074.5 1850. 1885.6 2169.2 2055.1 1953.2 1970.3 1996.6 1914.4 1972.4 2056.5 2011. 2028.9 2076.8 1945.1 1935.5 2049.7 1902.4 1900.7 2055.5 1931. 2128.7 2127.4 1970.9 2041.9 2051.6 2005.6 2019. 1998.7 2006.5 2082.5 1962.6 2053.8 1952.5 1863.7 1995.1 2021.1 1996.2 1959.9 2041.2 2050.1 2032.8 1987.3 2033.6 2004.8 1910.2 2010.3 2046.6 1983.2 1848.8 1977.6 2156. 2040.1 2051.9 1954.3 2096.5 2084.7 2095.2 1946.6 1980.5 1933.7 1926.4 1967.1 1981. 2060.5 1945.1 2066.9 2032.7 2062.4 1945.2 1937.6 1975.4 1967.3 1902.2 1881.5 1981.1 2089.3 1950. 2049.9 2179. 2001.1 2051.4 2044.9 2050.2 1960.2 2066.9 2002.9 1933.8 1969.6 1973.9 1930.6 1993.9 1935.1 1939.7 1991.2 1970. ] | [1430.9 1444.8 1413.5 1449.7 1462. 1409.8 1382.9 1489.7 1421.1 1419.4 1448. 1453.9 1399.2 1414.4 1367.7 1449.8 1453.9 1377.6 1423.9 1413.2 1447.2 1470.8 1409.5 1413.6 1396.8 1469.6 1441.2 1414.5 1412. 1393.8 1414.9 1433.8 1434.7 1413.5 1376.5 1460.5 1475. 1444.8 1433. 1451. 1405. 1399.1 1424.7 1413.2 1365.2 1467.3 1453.6 1473.9 1473.3 1430.1 1439.6 1452.2 1357.2 1435.5 1369.9 1413.2 1425.2 1435. 1450.1 1459.4 1468.4 1452.9 1371.3 1467.8 1399.8 1422.5 1418. 1390.8 1431.4 1431.8 1450.1 1454.3 1497.9 1427.8 1423.8 1432.2 1430.7 1423.6 1392.6 1459.6 1440.7 1450.7 1417.6 1400.2 1441. 1436.2 1388.9 1406.2] | [1001. 981.9 960.9 958.4 956.6 987.9 961. 956.9 976.5 962. 945.4 959.1 987.4 964.9 953. 992.4 975.4 978.4 993.9 974. 971.5 969.8 970.8 963.6 974. 981.4 990.7 977.5 960.4 965.7 952. 963. 952. 986. 948.4 965.9 990.4 964. 946. 971.6 984.6 991.4 981.8 946.5 948.4 981.4 964.6 963.3 959.3 942.2 956.7 954.9 985.4 988.1 961.5 957.6 987.2 988.4 962.3 966.8 964.3 972.2 961.7 963.1 968. 977.1 963.4 989.2 978.5 963.9 979.2 959.1 984.8 957.1 968.7 980. 914.2 971.2 974.4 979.5 969.6 936.4] | [699.9 684.3 708. 719.7 698.5 701.7 711. 708.9 709.9 707. 692.3 707. 698.1 702.3 701.3 712.5 704. 692.3 685.3 688.1 711.5 686.7 713.1 712.8 696.4 700.5 691.5 679.7 698.9 672.9 713.7 683.7 710. 700.4 698.5 699.8 690. 694.4 682.1 701.5 707.1 686.3 709.9 681.2 690.9 692.3 682.8 699.3 712.8 703. 722.7 693.7 696.8 704.2 715. 688.3 686.5 703.6 707.4 687.6 708.4 683.9 704. 708.3 721.4 705.2 698.6 707.6 711.8 707.7 694.8 698.3 683.5 697.7 701.7 689.1 697.8 705.6 702. 703.9 713.2 715. 697.2] | [475.6 500.2 496.9 494.9 495.7 481. 474.4 496.8 483.4 483.5 482.4 497.9 481.5 481.4 477. 492.8 477.1 491.6 485.8 479.2 487. 482.8 492.8 507.3 486.9 497.9 492.1 483.9 484.3 494.9 481.4 477.4 478. 485.9 478.2 494.1 476.2 483.7 494.4 499.4 489.2 494.9 486.7 479.1 467.3 483.3 485.3 487.8 480.1 483.6 489.9 484.4 484. 483.3 469.7 493.5 491.2 480.1 484.1 494.3 494.1 490. 480.4 495.7 483.3 493. 474. 478.5 477.9 477.8 483.1 486.7 482.9 481.9 498.4 480. 480. 484.6 479.6 483. 486.3 488.1 487.4 472. 486.5 482.6 483.7 486.8] | [315.3 303.5 311.2 310.6 304. 315.6 301.3 305.2 319.7 313.8 303. 310.9 308.2 302.8 310.2 308.3 306.6 317.9 314.3 309.2 302.2 305.8 318.6 303.8 310.6 309.1 309.7 315.3 310. 312. 309.9 310.9 294.7 316.8 299. 302.2 316.4 324.5 317.7 312.4 325. 318.4 317.1 302.5 313. 311.7 307.8 307.4 310.7 305.3 301.5 307.1 307.8 320.2 310.8 324.3 302.4 308.1 312. 309.4 308.6 306.6 309.3 314.5 315.8 303. 315.4 302.3 304.4 312.2 312. 316.3 317. 318.7 310. 316.1 310.2 308.1 308.8 316.7 316.2 312.1 314.2 308.7 306.7 321.5 309.1 309.2 304. 304.8 315.1 309. 317.1 318.9 323.6 312.8 315.1] | [237.5 232. 232.2 232.5 242.9 236.2 239.4 235.6 234.8 237.5 244.6 236.2 232.5 246.8 241.9 242.7 228.2 242.8 241.3 236.1 244.6 241.6 238.4 233.4 242.7 238.8 239.4 240.6 235. 242.8 237.1 239.8 235.4 241.7 239.1 229.8 244.7 238.4 235.8 235.2 241. 242.9 244.7 235.6 235.4 240.8 243.1 246.5 234.4 242.4 234.9 240.6 241.3 251.2 243.3 246.5 238.3 231.2 241.3 245.5 231.4 236.2 240.8 249.4 239.2 240.9 241.5 244.5 240.3 235.5 235.6 237.6 241.3 234.6 246.3 233.3 237.5 232.5 234.3 235.5 237.8 238.6] | [159.5 153.4 151.8 161.1 154.5 157.7 161.9 165.1 154.7 159.7 162.2 164.2 153.2 159.2 156.1 157.2 157.1 155.4 160.8 153.7 158.1 159.6 159.6 159.5 152.7 156.9 163.9 157.3 157.4 157.2 153.4 158.5 165. 151. 155.3 154.8 155.2 159.9 159.2 155.1 154.1 156.8 161. 154.2 158. 156.9 157.8 158.5 158. 160.7 158.5 161. 167.8 160.4 154.3 157.2 161.6 155.6 158.1 157. 155.6 156.6 162.7 161.4 148.7 151.4 167. 161.7 163. 158.6 158.2 169.6 169.5 152.9 144.5 161.2 157.7 156.3 162.7 151.1 156.8 152. 167. 163.5 160.2 160. 163.9 166.5] | [ 93.9 119.2 108.7 108.6 113.3 114.8 102.5 115.5 112.8 111. 105. 104.7 112.2 111.6 118.3 110.4 110.1 101.1 114.9 117.3 115. 111. 110.9 113.7 127. 112.6 115.2 118.5 121.4 117.6 118.4 119. 111.3 108.6 131.6 116.7 122.5 121. 114.4 120.9 107.9 110.5 120. 110.9 132.7 125.9 120.3 117.3 113.7 118.8 121. 118.6 115.5 112. 108.5 115.1 120.4 118. 111.9 111.1 113.7 120.3 113.7 111.3 111.1 119.1 105.1 121.5 118.9 116.2 115.2 120.1 126.5 107.8 134.4 114. 118. 112. 105.2 111.5 103.5 110.3 129.3 130. 125.8 124.4 121.5 122.9] | [ 2. 67.1 1.3 3.5 74.7 65. 4.6 4. 10. 4.4 71.4 9.5 87.5 6.5 57.9 2.8 87.7 73.1 74.4 2.6 78.2 10. 57.8 3.4 83.8 63.3 1.2 57.5 56.9 5. 4. 2.7 60.8 2. 92.8 1. 2. 2.4 2.7 9.5 82.3 2.7 4. 58.3 77. 9.5 1.8 1.2 70.2 55.3 95.4 3.3 79.3 68.7 79.3 2. 57.4 68.7 83.8 1.8 94.4 73.3 73.2 3.3 2.1 10. 89.5 61.1 82.7 97.6 10. 9.5 1.2 67.1 80.4 9.5 2.5 113.8 6.9 3.1 1.2 50.3 2.5] | [ 1.2 4. 10. 1.4 9.5 10.2 1.2 4.5 9.5 6.9 1.6 1.6 1.1 2.4 5. 3.6 3.7 1.9 10. 2.4 7.6 2.6 2.4 3.4 9.5 2. 1.8 2. 5. 72.6 3.2 2.6 1.1 1.1 6.6 10.2 2.3 6.9 3.3 9.5 47.2 9.5 3.2 5. 6.6 2. 5. 3.5 5.6 1.1 4.1 78. 4.1 4.4 2.7 10. 1.2 2. 9.5 1.3 6. 2.7 3.4 4.3 1.5 33.2 4. 3.2 5. 2.1 9.5 49. 76.8 2.5 1.9 3. 3.7 1.3 3. 3.3 3.2 9.5 98.5 2. 9.1 2. 3.1 9.5 5. 3. 101.5 61. 5. 63.4] | [ 9.5 3.1 4.7 10.7 4.3 9.5 10. 79.1 7. 1. 1.3 2.8 1.6 3.1 2.2 2.7 1. 78. 2. 2.4 7.7 1.2 95.1 2.6 4.9 9.5 6.9 5.7 1. 7. 3.4 7.2 9.5 1.3 1.7 9.6 5.1 9.5 2.3 1.1 6.4 1.2 4.9 9.6 4.6 1.3 2.6 4.9 10. 2.6 6.9 3.4 6.2 4.1 2.3 9.5 3. 2.7 2.7 5.9 4.9 5. 2.3 6.7 2. 1.2 9.5 70.8 9.5 2.6 4.2 57.4 3. 9.5 5. 10. 2.1 1.3 74.9 1. 9.5 2.8 2.3 2.6 5.7 2. 1.1 7.9 10. 66.2 5. 2.5 1.8 7.3] | [ 9.5 4. 5.3 9.5 10.9 10. 6. 3.4 6. 9.8 4.8 2. 6. 104.3 4.1 102.1 3.4 7.8 65.4 4.9 2.8 2.7 4.7 2.8 3.6 2.8 2. 8.1 1. 5. 1.1 1.4 4.2 7. 7.8 1. 5. 10. 4. 10. 5.1 7.6 1.2 7.8 1.1 1.2 2.8 3.8 7.7 1.3 1.7 1.2 7.8 7. 6.4 3.3 7. 96. 3. 8. 7. 6.6 2. 3.3 8.4 10. 10.3 63.8 8. 10.9 3. 65.6 2. 4.8 5. 7. 9.5 1.2 1.3 10. 3. 2.1 5. 72.1 5.3 9.5 10. 1.2 95.3 5.5 10. 10.9 1.1 8.4] | [147.7 109. 1.2 6.4 8.3 106.5 7. 9.5 6. 5. 4. 8. 1. 6.6 2.4 4.5 10. 9. 10. 3.1 3. 5.5 10. 10.3 134. 1.6 8. 7.9 2.2 2. 5.6 3.3 3. 1.6 8. 3. 2. 10.3 2.7 9. 96.1 8.9 176.8 86.3 3. 7.2 3. 1.2 5.9 178.2 5. 7. 6.6 133.4 8.9 1. 4. 6. 7. 2. 9. 73.5 3.6 10.3 1. 119.7 1. 4. 2.7 10.3 7. 3.6 1. 8. 2. 10.3 1.1 90.5 3.3 169.6 179.6 10. 2.9 9. 8. 206.3 5.6 4. 4.8 10. 4.1 1.2]
6.007 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Cologne_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Kilian Weiss | Universitaetsklinikum Koeln | Philips | Ingenia 3T | 3 | NIST | 130 | 23 | 21.7 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1962.9 1882.7 1897.8 1925.9 1885.9 1970.8 1943.5 1863.2 1918.1 1986.6 1889.1 1914.3 1979.2 1924.6 1872.3 1825.9 1927.2 1933.3 1927.8 1911.5 1922.3 1916.5 1953. 1966.4 1894.1 1863.7 1914.6 1896.2 1950.5 1916.9 1974.2 1924. 1912. 1934.6 1914.8 1888.8 1956.2 1956.4 1899.6 1898.9 1944.9 1981.4 1910.7 1943.8 1996.4 1915.9 1846.5 1908.9 1943.4 1855.5 1815.7 1902.9 1913.1 1909.9 2095.1 2080.7 2004.9 1872.8 1978.8 1982. 1900. 1848.7 1903.4 1935. 1913.8 1961.6 2053.2 2038.9 1931.6 1938.5 1963.9 1891.5 1866.9 1945.6 1921.7 2020.1 2043.1 1974.6 1884.4 1886.3 1913.6 1863.6 1874.6 2089.8 2027.2 1994.1 1914. 1981.3 2054. 2032.1 2069.8 2028.6 1994.3] | [1453.5 1473.7 1413.1 1376.7 1359.3 1371. 1445.4 1420.3 1439.9 1422.8 1399.2 1372.8 1397.5 1433.4 1505.6 1436.6 1414.3 1407.2 1432.8 1390.3 1365.9 1362.3 1378.2 1342.1 1408.5 1393.1 1379.4 1364.1 1389.2 1411.9 1377.1 1374.5 1368.6 1350.1 1455.7 1428.4 1371.7 1347.8 1382.5 1455.4 1390.8 1389. 1381.9 1408.3 1435.8 1416. 1376.9 1361.9 1401.2 1387.1 1411.7 1361.9 1378.2 1434.5 1355.3 1376.4 1404.6 1399.2 1408.6 1386.5 1454.5 1416.2 1416.2 1452.7 1399.4 1392.8 1416.8 1431.4 1408.7 1411.1 1412.2 1386.4 1416.2 1447.9 1402.7 1403.3 1433.2 1403.4 1417.9 1430.1 1405.7 1431.5 1402.6 1373.8 1346.8 1345. 1399.1 1439.3] | [944.1 967.7 971.4 965.3 969.9 964.8 937.2 935.1 957.9 986.3 967.5 957.1 950.4 955.6 947.7 938.2 959.8 966.9 961.7 972.6 959.1 946.5 945. 948.7 968.9 957. 952.2 967. 980.3 975.9 946.1 946. 957.3 948.7 978. 956.3 960.6 990.4 970.2 958.9 945. 946.1 957.8 952.4 950.9 968.6 953.6 966. 982.9 970.6 945.8 959.8 948.5 962.8 978.7 964.6 967.9 974.1 990.8 973.9 934.1 943.1 954. 945.1 949.6 952.9 988.2 983.4 974.6 944.8 936.6 941.8 928.3 947.3 956.2 993.6 987.2 942.9 954.7 938.4 946. 951.5 943.6 990.4 943.7 947.6] | [719.3 715.7 712.9 718.6 719. 717.2 728.1 727.6 721.9 724.5 730.6 721.2 723.5 716.3 721.8 723.4 721.6 725.2 725.8 705. 718.1 720.2 716.4 717.2 733.1 724.3 725.7 717.3 726.5 705.3 710.6 680.4 685.2 721.5 744.4 730.3 723.6 679.6 692.5 699.5 693. 685.2 688.5 687.2 693.2 720.7 687. 709. 719.6 692.3 689.2 693.6 692.8 691.7 695.8 693.1 696.6 722.7 720.2 688.2 693.4 705.1 721.7 714.6 700.2 729.3 735.8 705.1 701.7 718.8 695. 720.5 699.1 738.3 706.4 704.3 722.9 690. 698.7 688.8 692.6 727.4] | [488.6 502.1 510. 494.5 483.9 486.6 478.6 481.2 492.5 497. 491.8 492.1 477.7 482.3 486.2 484.1 485. 485.4 486.8 488.9 490.1 486.6 482.8 493.2 489.8 482.9 489.7 490. 489.3 490.9 490.1 484.5 489.5 491.9 485.4 491.5 494.7 498.7 499.7 490.8 488.4 486.2 499.8 488.8 492.5 491.1 496.5 494.1 487.3 494.3 494.1 500.8 490.5 499.3 498.6 489.5 494.5 486.5 487.3 491.7 494.2 486.6 497.2 497.5 494. 496.6 492.3 492.7 471. 496.4 500.2 494.6 485.5 493.6 499.8 499.2 477.6 494.5 489.8 482.9 485.4 492.1] | [317.5 319.3 312.6 318.6 328.8 308.9 307.4 309.6 314.9 308. 313.7 304.9 312.4 310.4 312.7 319. 314. 313.5 312.7 317.5 324.6 316.1 314.3 315.4 312.8 316.4 315. 315.1 315.9 314.4 315.4 310.4 311.1 313.1 320.4 316.7 313.1 314.6 308.1 314.3 314.5 309.9 310.1 312.9 320.7 317.5 315.7 318.6 314.9 313.7 319.7 316.5 312.7 318.8 319.9 313.1 314.2 314.7 315.9 319. 309.2 322.2 316.2 311.8 319. 315.8 315.4 314.4 317.3 317.2 312.3 306.8 318.3 317.1 315.7 312.4 311.1 314.9 317.4 321.3 311.2 318.3 314.3 316.9 318.4 319.1 315. 321.1 320.4 315.8 316.1 322.7] | [242.7 249.8 248.9 243.8 231.4 239.7 241.5 236.2 246. 245.7 241. 237.7 237.2 242. 246.2 247.1 241.7 248. 243.5 238.5 237.5 240.7 239.1 243.2 241.5 243. 248.5 245.8 245.5 239.3 239. 245.2 240.6 241.8 241.3 243. 243.7 248.4 243.9 244.3 248.5 251.3 246.6 241.3 239.6 244.2 243.8 248.2 241.5 249.5 256.6 249.6 248.3 246.6 240.7 241.7 243. 255.3 248.4 245.6 243.2 244.6 247.3 251.6 247. 245.5 243.8 264.7 243. 247. 245.9 245.3 249.6 250. 250.6 249.6 255.8 248.3 251.5 248.2 245.8 244.8 246.8 251.4 244.4 248.7] | [161.6 163.6 232.9 230. 254.1 160.1 243.3 226.8 161.3 229.3 234.2 234.7 227.7 229.1 228.2 225.6 205.7 214.1 231.9 236.2 164.2 237.8 167.5 232.5 165.1 227.6 235.3 231. 229.1 236.1 233.6 169.3 169.3 243.6 230.3 163.9 234.4 234.7 227.1 170.1 165.7 236.2 171.6 246.4 238.1 236.1 162.7 239.3 239.4 238.6 161.6 213.2 235.8 247.8 166.5 240.6 250.8 179.2 170.6 237.4 162.7 224.1 161.8 261.1 174.5 171.8 173.8 170.1 240.5 170. 164.7 240.1 234. 242.8 275.4 172. 166. 168.4 235.3 225.2 273.1 239.4 154.6 157.8 160.5 156.4 164.8 162.6] | [137.5 134.8 126.5 132.1 134.9 141.4 137.7 130.9 134. 141.3 124.1 132.8 140.9 122.9 119.1 120.3 139.7 136.6 116.4 140.7 118.9 119.8 117.4 135.9 118.4 120. 135.5 140.5 142.4 136.8 137.4 116. 140.9 139.2 143.9 147.8 143.5 127. 142.7 126.4 122.2 118.3 142.2 146.9 145.6 121.4 138.9 135.9 138.1 120.6 141.4 144.4 133.9 128.3 125. 142.3 112.8 139.7 137.8 141.1 147.6 148.2 149.8 128.2 123.9 130.2 147.3 109.8 149.5 124.6 118.8 113.7 115.7 101.8 127.2 123.7 133.1 139. 133.2 131.3 141.9 145.3] | [ 81.4 95.2 103.9 96. 89.5 78.1 105.5 88.8 97.3 81.8 75.3 81.1 86.9 93. 58.6 84.8 84. 67.8 62.6 75.1 73. 83.4 79.7 83.5 72. 86.1 92.3 81.7 87.4 99.7 98.2 95. 87.8 80.8 98.3 106.8 75. 81.4 82.8 83.9 83.2 94.4 100.1 99.1 92.5 75.8 86.8 89.4 92.7 90.4 95.5 105.1 107.2 100.3 84.8 92.3 90.2 91.5 91.3 87.2 96.8 98.3 106.4 118.5 96.4 99.1 94.1 92.8 115.9 88.1 86.5 86.4 96. 93.7 85.2 87. 88.8 110.1 108.3 85.7 95.3 91.8] | [ 1.1 2.7 1.7 101.3 42.7 2.8 58.3 73.9 4.6 1.1 101.4 78.6 57.1 54.1 2.5 2.7 2.7 58.2 1.3 76.8 50.6 4.1 1.3 3.3 4.6 2.7 10. 9.4 2. 2.7 2.7 7. 65.8 62.3 67.2 1.8 2.6 53.4 1.1 89.5 4. 69.6 59.5 70.7 74.4 69.5 2.7 2. 81. 72. 84.5 57.8 77.7 71.9 47.9 85.2 72.5 2.5 10. 72.6 77.2 94.2 100. 119.1 2.8 67.2 61. 76.8 2.5 1.2 90.2 68.8 74.9 115.5 80. 80.7 88.6 79.2 10.6 111.1 2.8 82.6 1.1 3.1 50.7 75.7 4. 2.5 6.8 94.5 98.5 90.9 65.5 100.9] | [ 54.7 5. 1.8 2.7 9.5 64.4 4. 93. 3.5 4. 1.3 1.2 105.7 2.4 120.9 74.5 69.3 9.5 8. 78.7 80.2 74.1 1.3 1.8 51.9 68.6 46.4 1.6 1. 52.3 54.8 1.4 8. 5. 41.1 1.1 8. 1.2 60.2 1. 41.3 103. 5.3 40.5 3.4 2.7 3. 66.2 70. 7. 5. 9.5 91.6 1.4 2.5 1.8 3. 54.6 82.4 48.4 4.5 2.1 67.4 85.3 5. 1.3 1.1 93.8 93.5 78.4 2. 1.3 9.5 2.9 98.9 4.4 129.8 91.5 9.5 91.3 89.8 1. 106.3 10. 2.5 3. 2.8 2.8 3.4 79.6 102.4] | [ 4.9 4.5 119.6 148.8 66.2 2.1 86.8 103.4 98.9 112.1 5.6 166.3 63.9 2.5 1.8 122.3 115.8 9.1 89.8 4. 1.9 1.8 1.8 4. 3.4 3.4 2.5 74.7 9.5 111.8 3.3 9.5 59.4 1.1 50.6 1.8 5. 87.2 2.7 7.7 68.4 69.1 101.1 77.1 54.6 1.3 9.5 2.1 68.9 61.9 9.5 69.3 10.3 79.5 3.3 2. 2.1 2.5 4.9 82.1 72.1 8.3 2. 5. 85.2 57.8 2.5 1.2 5. 3. 86.3 2.5 1. 114. 103.5 70.6 86.2 2. 5. 2.5 3. 1.4 79.4 1.2 101.8 70.5 50.3 2.5 95. 80.1 85.1] | [2.5580e+02 3.6340e+02 2.0000e+00 8.1000e+00 9.3000e+00 2.7000e+00 2.7000e+00 3.0000e+00 1.2410e+02 7.0200e+01 6.1200e+01 1.2000e+00 1.2000e+00 2.4000e+00 3.2000e+00 1.2000e+00 1.3350e+02 2.5000e+00 1.2000e+00 8.0000e+00 7.1000e+00 1.3000e+00 6.3000e+00 9.9600e+01 1.1523e+03 1.0000e+01 2.5000e+00 4.2000e+00 1.0000e+00 3.7000e+00 1.8000e+00 1.3000e+00 2.9000e+00 1.0220e+02 2.9610e+02 6.0690e+02 2.0000e+00 2.5000e+00 9.5100e+01 3.7000e+00 6.3900e+01 6.3000e+00 3.4000e+00 1.1700e+02 1.2860e+02 1.7670e+02 1.8000e+00 8.4000e+00 2.7000e+00 6.9100e+01 6.3200e+01 3.3000e+00 4.8300e+01 1.0980e+02 1.3000e+00 3.2000e+00 1.9210e+02 3.0000e+00 2.7000e+00 2.0000e+00 4.1000e+00 1.0000e+00 5.9800e+01 1.8000e+00 8.5000e+00 2.8000e+00 1.0000e+02 1.4870e+02 5.4200e+01 5.8100e+01 6.2500e+01 6.3900e+01 4.3800e+01 6.0100e+01 1.6480e+02 8.7800e+01 2.7000e+00 3.2000e+00 7.4100e+01 5.9800e+01 1.3000e+00 4.9000e+00 2.1000e+00 7.3160e+02 2.0000e+00 5.6500e+01 7.0800e+01 7.0000e+00 1.0060e+02]
6.008 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Cologne_MR1_NIST_Complex_T1map.nii.gz | Complex | Kilian Weiss | Universitaetsklinikum Koeln | Philips | Ingenia 3T | 3 | NIST | 130 | 23 | 21.7 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1909.4 1755.2 1853.3 1905.7 1722.9 1937.7 1964. 1803.8 1898.6 2019.8 1828. 1889. 1959.6 1868. 1833.7 1781.7 1924.7 1948.4 1907.1 1916.6 1905.4 1906.8 1914.2 1983.5 1855.5 1814.6 1896.6 1908.6 2013.9 1921.8 2066.3 1975.3 1848.7 1910.6 1891.7 1866.4 1961.9 1991.1 1875.7 1923.8 2016.6 2100.5 1940.9 1946.2 2011.2 1874.7 1790.1 1899.9 1947.2 1851.4 1823.5 1963.1 1909.7 1949.6 2145. 2123.5 2025.9 1839.4 2000.8 2014.9 1888.2 1839.6 1926.2 1987.3 1924.8 1917.9 2007.5 2072. 1901.8 1936. 2002.8 1866.6 1855.3 2032.7 1961.4 1970.3 2017.2 1959.3 1809.2 1855. 1911. 1819.2 1827.4 2078.6 1939.2 1972.8 1806.5 1966. 2060.8 1939. 2026.1 1972. 1955.5] | [1442.2 1488.9 1404.1 1364. 1351.6 1370.3 1428.6 1395.1 1462.2 1431.5 1427.8 1390.6 1428.6 1463.9 1484.5 1451.1 1421.6 1418.1 1468.1 1395.4 1369.6 1365.5 1402.5 1354.3 1373.3 1363.1 1379.7 1361.3 1407.2 1447.7 1396. 1405. 1395.6 1342.4 1472.2 1451.9 1388.1 1341.9 1404.3 1512.7 1410.5 1406.3 1403.3 1443.1 1461.5 1441.3 1379.3 1369.5 1430.4 1391.1 1431.1 1364.7 1369.1 1469. 1331.6 1354.2 1428.2 1411.8 1428.6 1384.1 1488.7 1434.5 1427.1 1467. 1391.5 1409.4 1436.8 1466.1 1417.7 1417.4 1411. 1368.3 1413.5 1444. 1390.4 1407.1 1494.4 1406.3 1462.8 1453.3 1388.6 1431.7 1411. 1405.7 1358.9 1326.2 1407. 1464. ] | [ 943.3 970.9 983.8 960.4 983.9 984. 925. 926.4 963.6 988. 966.4 989.9 945.3 970.3 961.3 931.6 981.5 979.5 944. 970. 962.3 952. 943.8 955.6 989.5 963.8 945.1 965.7 975.8 984.3 937.8 948.6 973.1 951.6 997.7 965.7 963.4 1017.7 958.4 946.7 940.4 944.9 968.3 959.5 949.4 981.1 948.1 960.2 987. 969.6 951.9 971.7 953.6 980.1 1002.6 968.9 973.9 983.3 1001.1 982.3 919.8 954.4 969.5 952.9 960. 954. 1009.2 991.1 994.6 947.6 934.1 949.2 917.8 955.6 962.1 1009.2 984.7 946.1 964.5 937.4 953. 970.9 928.3 977.6 938.1 955.7] | [706.6 713.1 694.2 684.8 695.4 701.3 694.6 695.5 693. 701.7 711.8 717. 719.8 682.5 710.8 698.2 711.5 693.4 696.8 700.5 709.4 698.4 701.8 687.8 702.2 692.4 723.2 683.9 689.6 713.5 696.4 671.7 683.2 702.8 711.7 699. 684.8 667.8 685.5 706.9 694.1 682.5 692.2 674.6 692.3 698.6 680.1 694.4 714.8 692.7 687.3 713.6 687.7 684.3 698.5 689.5 704.3 720. 703.4 684.9 699.7 719.4 715.8 692.8 700.7 719.1 704.9 706.7 705.5 692.8 696.4 689.5 704.1 718.1 716.9 714.6 694.5 691.1 705.9 695.6 690. 694.9] | [485.4 499.7 513.7 493.4 483.1 487.9 471.8 471.4 493.5 499.8 486.6 498.8 470.7 481.5 479.9 482.3 481.9 480.8 477.8 492.8 489.1 483.2 481.8 491.8 486.9 477.1 486.6 486.3 479.8 488.1 490.1 483.2 484.5 489.4 473.9 487.1 490.4 498.7 503.8 487.6 484.2 477.5 492.2 479.7 482.7 479.5 492.6 491. 480.5 496.3 486.9 501.2 481.9 497.4 496.4 481.2 493.9 479.9 479.6 488.2 486.7 471. 495.8 493.8 488.4 491.6 489. 488.6 464.7 500.4 495.2 490. 472.6 493.1 501.5 496.2 478. 495.4 484.7 475.4 479. 492.3] | [317.3 321.5 307.6 315.9 328.8 307.7 304.9 302.1 316.3 299.4 312.7 305. 315.3 307.8 312.4 323.1 311.1 313.8 310.8 317.8 330.3 314.6 314.6 316.7 304.4 316.4 315.8 314. 318.2 314.3 313.5 307.4 308.8 309.5 322.4 316.3 308.3 314.4 299.9 310.3 313.1 305.7 303.4 309. 321.3 316.7 313.1 319.8 309.5 312.3 320.5 315.8 309.3 320.9 321.1 308.7 313. 312. 310.2 322.5 308.4 323.9 313.7 303.3 319.3 314.7 314.1 314.3 317.3 317.4 307.2 302.7 316.6 315.2 316.8 308.8 306.3 315.1 316. 323.4 309.6 317. 310.6 320. 319.7 320.1 312.2 319.5 320.1 314.2 312.6 321.2] | [243.5 255.3 251.4 250.3 228.4 243.4 244.6 237.7 246.8 245.7 244.6 237.2 238.9 245.6 251.2 246.7 242.4 247.1 243.7 238.3 238.2 242.3 238.9 246. 241.7 243.1 249.7 245.7 246.4 237.8 237.3 247.2 239.6 244.8 242. 245.3 242.5 248.2 244.3 242.2 249.3 255.4 247.5 241.4 236.6 247.2 244.4 246.4 235.3 252.4 263. 248.6 247.4 246.4 236.3 242.4 242.6 254.2 252.7 245.8 239.9 243. 247.4 256.2 247.5 243.6 241.2 266. 239.5 247.9 244.9 241.8 248.7 250.6 253.4 251.8 255.1 250.9 250.3 249.3 243.5 245. 247.5 250.2 244.1 249.5] | [163.9 167.3 174.2 163.3 178.1 161.2 170. 164.8 162.6 166.3 166.8 167.1 158.5 168.9 163.2 164.9 148.7 151.3 165.8 170.2 160.1 170.5 168. 163.7 168.3 163.4 168.9 167.2 165.7 169.6 163.2 172.1 169.9 170.6 168. 160.4 168.4 164.5 155.8 174. 166.4 168.2 174.3 172.4 172.5 168.1 156.8 162.9 170.1 166.1 160.8 152.8 166. 172.1 165.4 166.1 173.6 184. 171.9 162.7 165.2 156.1 160.2 179.2 175.8 174.1 173.1 168.2 166.6 175.7 163.6 173.5 165.1 167.2 181. 173.3 165.5 173.2 165.7 158.5 179.1 168.9 157.1 159.9 164. 153.3 167.1 162.8] | [118.5 118. 131.5 115.8 120.9 126.1 125. 110.6 121.6 126.1 126.5 116.4 128.4 123.2 120.8 119.7 131.1 121.6 109.5 127. 119.7 121. 121. 118.4 115.7 117.3 117.1 125.6 122.6 119. 125.8 112.7 126.4 124.7 127.5 133. 129.7 127.7 127.6 131.5 128.7 114.4 123.8 131.5 130.1 116.1 124. 121.9 123.6 122.6 121.7 125. 132.8 127.4 123. 129.7 112.7 126.4 123.1 118.4 130.5 128. 131.2 133.2 120.8 130.9 123.8 109.4 128.2 127.7 117.5 116.1 124.9 98.8 121.6 126.9 135.7 141.7 136.2 134.2 150.7 129.7] | [ 78.5 103.6 107.4 105. 100.8 64.5 105.8 86.2 110. 86.7 70.4 84.6 93. 94.2 56.2 92.9 88.2 53.8 48.5 72.7 74.3 81.3 74.1 92. 60. 91.8 104.1 86.2 91.2 107.4 103.7 101. 90. 86.9 109.3 111.6 78. 80.5 77.6 81.5 70.8 98.2 105.2 96.3 93. 71.8 98.4 93.3 96. 96.6 99.5 110.4 114. 104. 93.3 96.7 86.7 94.1 89.6 91.1 95.1 99.8 114. 112. 99. 108.3 98.2 98.8 113.8 82.6 92.2 76.2 104.9 101.1 83.9 97.2 88.9 108.5 110.6 77. 103.2 93.7] | [ 6.7 1.1 1.3 90.9 1.5 4. 67.1 74.3 1. 10.6 86.5 76.8 1.8 81.3 10. 65.8 4. 44.8 6.1 91.6 11. 9.8 10. 53.7 1.4 62.3 65.9 9.5 1.8 1.2 1.8 1.3 73.9 65.3 79. 58.8 63.4 68.2 10. 85.4 55.5 64.8 3.4 67.1 87.7 88.1 2.1 1. 87.4 74.2 95. 74.9 57.2 73.8 4. 86. 96.1 2. 3. 78.8 80.3 95.9 106.1 107.8 44.3 9.8 61.8 58.7 9.5 2.4 99.1 88.1 1.3 96.7 84.6 83.6 96.9 86.5 70.5 96.8 1.2 74.5 2.4 4.5 3.1 88.3 3.4 2.4 2.7 99. 106.8 99.5 63.8 106.4] | [ 79.9 61.5 6.9 3. 10.9 90.1 5. 90.9 10. 6. 3.1 1.1 101.3 9.5 99.6 107.1 103.2 9.5 9.5 94.8 89.2 82.6 3.1 3.2 1.3 106.1 79.4 4.1 10.9 78.2 2.5 7. 60.3 1.6 91.8 1.2 1. 9.5 105.5 2.4 59.2 107.5 5. 105.7 79.6 5. 3.1 95.6 113.7 5.3 9.6 2.4 91.8 47.5 3.5 9.3 1.3 10. 124.2 1.3 7.7 77.8 92.3 84.2 10. 9.5 3.9 74.4 78.8 66.5 4.6 9.5 75.1 3.3 98.3 2.7 90.1 82.5 3.5 107.6 111.2 2.5 104.9 9.5 9.5 10. 76.9 68.5 1.3 85.5 109.6] | [ 81.7 2.7 129. 156.5 2.7 2.4 116.4 9.5 5. 76. 2.1 157.1 129.3 69.2 9.5 139.9 134.5 5.3 2.5 3.3 1.8 3.6 8.2 3.7 68.3 4.1 4. 108.1 3.2 136. 6.5 10.3 132.2 83.8 109.7 9.5 87.4 127. 4.6 10. 97.1 135.8 137.7 97. 116.5 70.8 4.5 87.6 127.9 108.2 3. 117.6 1.9 120.1 75.9 64.1 9.5 97.5 79.6 118.6 125. 2.4 9.5 8.6 66.7 98.3 2.1 2.5 103.5 4.8 94.2 9.5 1. 111.2 116.6 114.8 92.3 9.5 4. 9.5 117.9 8. 1.3 8.2 126.2 80.2 58.8 4. 89.1 83.9 89.1] | [2.7390e+02 3.7260e+02 1.5000e+00 1.8000e+00 4.9000e+00 1.0210e+02 8.0000e+00 2.4000e+00 1.9190e+02 2.2360e+02 2.6620e+02 7.2000e+00 1.6700e+02 8.0000e+00 6.0000e+00 6.0000e+00 1.9670e+02 6.0000e+00 9.4700e+01 2.4900e+02 3.0000e+00 4.9000e+00 9.2400e+01 2.0760e+02 1.4973e+03 6.0000e+00 3.0030e+02 1.6340e+02 1.7040e+02 7.5600e+01 3.2000e+00 9.0000e+00 2.1900e+02 2.5000e+00 2.6280e+02 4.8530e+02 5.7000e+00 1.2020e+02 2.2300e+02 8.9500e+01 1.8560e+02 2.0000e+00 9.3000e+00 1.2990e+02 2.0000e+00 1.5330e+02 8.2000e+00 1.0000e+00 8.0000e+00 2.7260e+02 2.1820e+02 6.6000e+00 2.1360e+02 1.7810e+02 2.1000e+00 9.0000e+00 1.5300e+02 1.9000e+00 1.9000e+00 6.5500e+01 2.4000e+00 1.2440e+02 1.5070e+02 2.0000e+00 2.6000e+00 8.3000e+00 6.1000e+00 3.6940e+02 2.8910e+02 2.3490e+02 1.3910e+02 1.8260e+02 7.2900e+01 1.3070e+02 2.4170e+02 2.9090e+02 4.0000e+00 1.0000e+00 3.9920e+02 8.2000e+00 9.0000e+00 2.0000e+00 6.5000e+00 6.8570e+02 6.0000e+00 1.4350e+02 1.9870e+02 9.0000e+00 1.1330e+02]
6.009 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Cologne_MR3_NIST_Magnitude_T1map.nii.gz | Magnitude | Kilian Weiss | Universitaetsklinikum Koeln | Philips | Ingenia 3T (tuned to slightly lower f0 due to proximity to other MR systems) | 3 | NIST | 130 | 23 | 22 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1803.1 1821. 1850.8 1859.7 1868.9 1838. 1784.2 1790.2 1817.8 1844.1 1845.6 1838.9 1847.5 1843.6 1874.5 1786.6 1775.9 1786.8 1805.6 1817.5 1799.3 1797.1 1773.7 1824.3 1864.6 1788.9 1880.3 1829.7 1806.4 1803.6 1773.3 1798.1 1783.3 1783.2 1802. 1784.3 1774.3 1852.4 1774.3 1794.3 1788.3 1768.2 1773.8 1956.4 1864.1 1794.3 1798.8 1792.9 1770.1 1778.1 1810. 1764.1 1742.6 1803.2 1854.2 1853.3 1800.9 1785.5 1773.2 1747.7 1837.6 1797. 1755.7 1795.9 1802.3 1816.9 1846.6 1830. 1837.7 1834.9 1814. 1839.8 1846.6 1835.3 1865.9 1847.9 1899.3 1846.7 1908.7 1935.6 1908.1 1847.6 1849.8 1854. 1836.3 1883. 1891.8 1859.9 1851.8 1828.9 1797.2 1810.7 1802.7 1856. 1821.4 1856.5 1806.4] | [1368. 1336. 1343.5 1327.9 1328.6 1338.5 1373.2 1337.1 1336. 1347.4 1331.4 1330.5 1346.2 1342.3 1347.1 1367.1 1340.8 1366.2 1351.3 1325.1 1300. 1333.1 1365.9 1342.5 1368.9 1357.2 1349.5 1347.4 1351.4 1308. 1289.8 1353.4 1368. 1370.4 1340.2 1355. 1348.3 1318.7 1357. 1336.9 1318.6 1337.2 1318.8 1354.3 1300.6 1341.2 1331.7 1326.4 1359.1 1372.5 1333. 1339.1 1348. 1334. 1289. 1316.8 1325.7 1330.3 1340.9 1342.7 1326.5 1350. 1367. 1341.2 1323.7 1308.1 1328. 1362. 1341.6 1351.9 1330.4 1320.9 1335.4 1346.1 1359.8 1350.4 1347.3 1335.5 1348.7 1355.5 1330.3 1360.6 1369.9 1386.6 1346.1 1324.9 1355.7 1350.2] | [917.2 916.9 910.5 901.1 904.4 903.6 921.7 909.5 905.1 906.4 907.3 918.1 905.8 906.2 935.2 916.2 917. 901.1 902.9 910.4 917.5 901.2 905.3 930. 922.3 914.6 918.4 905.9 912.8 916.6 913.9 900.6 902.5 930.6 919.7 913.2 927. 904.7 911.5 928.4 920.5 915.8 908.3 914.6 916.3 908.9 913.7 913.6 912.9 916.5 901.1 910.4 921.6 910. 921.7 919.4 917.5 912.1 912.8 909. 903.6 920.5 925.3 907.4 921.8 920.6 921. 908.5 916.8 926.6 920.6 915.1 919.4 917.9 914.2 912.8 927.9 932.7 932.8 925.3 908.7 921. 912.2 919.1 916.2 919.1 913. 915.9] | [678. 667.1 662.2 670.6 667.8 669. 672.8 662.2 677.7 665.6 672.7 670.9 666.7 672.6 676.6 664.7 666.3 669.4 670.3 668.3 672.6 666.2 675.4 662.2 677.4 670.5 669. 661.1 660.9 661.3 669. 674.3 670.2 665. 682.4 668.8 680.6 667.5 666.8 666.3 665.7 671.8 672.9 671.9 684.3 666. 668.6 670.6 666. 665.3 670.1 665.6 669.5 670.7 680.2 663.1 664.5 669.8 668. 666.2 666.1 672.2 674.8 669.6 675.2 667.3 669.3 664.3 669.9 668.4 662.7 667.3 668. 666.2 661.1 672.6 671.3 668.9 667.6 674.8 673.4 669.5 664.3 667.8 667.3 665.4 672.6 675.9] | [462.3 470.1 465.6 472.6 469.6 460.6 472.2 473.7 469.9 469.7 466.4 468.8 477.3 464.8 472.9 468.4 468.1 471.5 477.1 472.7 471.8 475.9 476.8 483.1 465.2 471.3 473.8 474.5 478.8 477.5 479.5 476.2 480.7 479.6 477. 478.9 476.6 478.3 484.3 476.3 480.5 474.6 475.1 473.8 474.5 477.9 479.8 479.3 480.1 479.3 477.7 479.2 478.5 472.7 471.4 476.4 479.6 470.2 472.7 478. 477. 481.4 480. 474.4 470.9 472.4 478.9 477.2 481.5 482.6 479.1 479. 478.3 468.1 471.7 471.1 477.6 475.3 473.5 482.3 473.7 474.3 480. 476.1 471.1 472.4 478.8 481.4] | [299.2 299.8 304.6 306.1 302.7 309. 307.1 313.6 314.5 315.9 321.7 312.8 308.9 305.7 304.8 305.9 309.4 308.1 310.7 309.6 306. 300.8 298.6 305.3 310.6 309.1 312.5 315.5 311.7 312. 307.4 309.9 306.6 305.3 304.5 308.7 311.5 310.9 314.8 311. 314.7 312.1 309.6 311.1 303.9 304.1 309.4 313. 311.5 311.6 307.2 311.4 312.8 310.5 313.5 305.4 303.3 309.8 314.6 311.1 313.2 309.8 310.1 315.3 312. 315.3 303.7 305.9 313.1 311.9 308.9 316. 314.5 311.9 314.5 311.7 297.9 301.9 309. 309.8 309.9 310.9 308.4 311.5 312.6 303.4 309.6 313.8 310.8 308. 307.9 310. 307.8 315.1 313.3 309.9 309.9] | [236.4 242. 241.4 240.4 238.9 241.1 241.6 245.3 246.3 242.7 244.4 243.8 247.7 238.5 242.8 239.8 240.2 239.2 238.8 238.5 241.5 234.6 248. 247.3 238.9 247.9 244.2 243.4 242.3 241.9 243.5 243.9 235.8 251.5 244.4 242.6 240.2 240.5 242.5 246.4 243.2 244. 240.6 244.6 242. 242. 240.7 242. 241.3 246.3 243.2 243.5 240.5 246. 242.2 241. 240.9 242.4 244. 240.9 240.5 241.8 240.3 244.5 246.6 242.4 238.9 246.7 245.7 239.7 241.4 244.5 238.1 242.2 242.7 236.8 240.4 243.1 242.2 243.1 243.5 240.2 236. 238.8 242.8 242.5 243.6 240.6] | [163.5 162. 160.4 162.1 296.5 164.5 154.1 304.9 297.7 293.3 304.3 301.8 300.5 163.4 164.7 311.4 316.3 309.7 164.5 293.2 307.2 294. 312.4 164.8 164.7 162. 162.2 161.8 164.4 161.4 162.5 160.5 160.8 165.4 163.2 289.9 160.9 161.3 296.2 168.3 166.2 164.6 168.1 164.4 161.2 164.5 161.5 159.6 296.9 295. 166.1 164.9 168.2 301.3 153.7 164.9 162.3 161.2 161.8 285.7 298.2 163.4 166.5 160.2 158.7 295.8 160. 161.1 161.2 162.7 162.9 164.6 165.4 160.4 288.9 326.7 166.7 167.4 162.7 164.8 166.7 167.1 161.7 294.7 303.6 291.4 161.7 310.5] | [115.7 141.3 148.1 150.6 118.8 118.5 118.4 150.8 144.2 154.5 153.6 158.2 161.3 161.3 121.4 151.3 146. 147.6 118.5 146.4 119.9 151.1 150.5 146.2 117.9 159. 156.9 160.2 156. 151.7 162.2 120.5 149.9 144.8 114.7 163.3 147.8 118.4 121.4 120. 154.5 154.8 157.3 151.4 107.3 158.9 148.2 150.5 148.8 148.5 160.8 146.7 155.2 118.4 116. 142.1 142.4 149.8 119.4 119.5 156.1 147.5 155.4 119.6 119.3 124.3 159.1 153.5 156.6 123.4 155.2 157.3 181.5 156.1 116.9 130.5 155.1 152.2 146.8 148.6 158.8 122.4 130.9 127. 119.6 118.4 117.7 126.9] | [ 3.1 91.8 94.4 100.2 86.1 66. 76. 76.6 81.9 88.1 104.9 93.8 89.1 99.1 85.1 84.2 69.6 76.3 87. 92.2 86. 88.2 81. 93.3 108.2 68.7 78.4 83.8 90.2 89.5 88.3 88.4 56.7 61.4 89.1 63.2 86.5 83.1 85.7 95.6 85.8 84.1 2.7 4.9 73.4 2.7 77.9 84.3 92.4 90.3 74.1 84.6 82.1 66.4 1.2 84.9 85.1 86.9 79.2 81.4 76.1 86.8 88.9 88.7 1.2 93.8 2.5 108.3 85.9 88.1 92.7 90.6 95.1 84.4 107.4 96.5 82.5 84.5 77.1 80.4 52.6 78.9 115.1 102.9 96.9 105.8 110.3 102.3] | [68.2 1. 60.2 2.8 2.8 1.9 1.2 1.9 1.2 8.1 3.3 1.8 2.8 6. 2.9 9.5 4.1 91.9 3.4 2.5 2.5 3.1 2.7 3.2 2.1 6.3 1. 46. 3.3 57.9 3.2 2.4 3. 62.7 10. 3. 5.6 1.8 3.6 1.3 30.7 2. 9.5 3.4 3. 2.5 2. 2.5 1.2 9.1 3.3 56.3 2. 40.8 1.8 1.8 3.6 3. 4.5 4.6 2.5 2. 85. 9.5 1.3 2. 2.6 3.2 2.7 2. 1.2 1.2 3. 59.7 1.2 3.6 2.7 3.1 1.8 3.4 1.2 2.1 3.3 8.2 5.1 1.8 2.4 3. 1.2 3.3 1.8 5.6 1. 2. ] | [ 79.8 71.9 72.2 67.3 90.1 157.8 1.2 1.8 4. 2. 1.2 2.7 5. 61.1 2. 92.4 64.5 2. 49.9 1.1 1.3 1. 1.1 53.6 64.6 3. 2. 3.2 1. 1. 1.2 2.5 3.1 10.6 4.1 3. 3. 3. 2.5 3.1 1.3 7.5 3.1 6.6 1.2 1.8 2.3 2.3 1.4 1.8 1. 2.5 1.8 5.3 3.4 2. 1.7 3.1 1.8 3.2 4. 1.2 3.4 3.4 4.1 10. 2.7 1.3 1. 2.1 1.9 1. 2. 1.1 2.4 10. 3. 2.8 4. 3.1 2.5 2.4 1.3 2.5 3.2 3.3 1.2 3.4 4. 2. 2.5 3.1 2.9 4. ] | [ 4.3 104. 97.9 83.6 58. 5. 1.3 2.3 1.2 5.3 8.4 4.7 3.9 1.2 3. 4.2 4.3 1. 65. 3.1 40.3 1.2 5. 1.4 2.1 9.5 2.4 7.9 2.1 52.2 105.6 3.2 2.5 1.3 6.5 1.8 1. 3. 3.3 1.3 63.4 113.9 1.3 6.7 2.8 1.2 6.4 1.3 6.8 1.7 3.3 53.4 75.7 8.6 3.1 5. 2. 5. 1.2 6.4 1. 9.5 1.4 69.4 5.6 2.5 1.2 6. 1. 4. 1.9 60.5 4.9 62. 9.5 2.5 2. 1.2 4.8 3.9 9.5 51.7 6.1 49. 2.5 1.1 1.2 1.2 1.3 3.6 8.7 2.7 3. 4.3] | [176.9 1.2 64.2 3.1 1.2 6.2 6.2 61.5 147.3 206.3 2.6 3.1 1. 1. 2. 1.6 2. 3.2 202.7 149.6 74.2 3.3 56.8 56.8 10.7 4.7 5.1 3.3 8.9 4. 2.5 10.6 5. 4.7 1.2 72.4 4. 1.8 3.3 7. 4. 1.3 2.4 5. 2. 1.9 2. 1.9 62.5 1.2 74.2 2.4 2.1 1.2 5. 2. 2. 1.1 6. 68.8 2. 1.1 1. 2.1 8.9 1. 1.2 3.4 5.8 4.1 9.5 1.2 2.4 5.9 7. 2.7 1. 7.9 1.8 5.8 2.5 98.8 2. 6.6 82.6 46.5 81.8 3. 2. 1.2 3.6 1.3 10.9 3.6]
6.01 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Cologne_MR3_NIST_Complex_T1map.nii.gz | Complex | Kilian Weiss | Universitaetsklinikum Koeln | Philips | Ingenia 3T (tuned to slightly lower f0 due to proximity to other MR systems) | 3 | NIST | 130 | 23 | 22 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1789.7 1807.6 1853.9 1782.5 1844.5 1838.6 1730.4 1779.6 1793.4 1843.7 1844.7 1829.5 1850.2 1832.2 1963.1 1778. 1748. 1766.5 1814.2 1837.6 1814. 1796.3 1732. 1844.1 1928. 1765.1 1927.5 1848.3 1826.1 1830.4 1778.6 1832.6 1725.7 1756.6 1820.5 1741.8 1755. 1896. 1761.6 1809.6 1817.2 1783.2 1773.3 1984.9 1878.3 1754.3 1796.5 1798. 1730.9 1759.3 1825.9 1750.8 1723.4 1830.1 1854.7 1827.9 1757.9 1745.1 1736.3 1702.2 1856.9 1792.3 1710. 1807.7 1763.8 1752.5 1811.3 1783.4 1786.6 1793.2 1776.6 1828.6 1849.4 1831.9 1906.7 1859.4 1886. 1822.8 1896.9 1944. 1910.4 1802. 1853.8 1841.5 1815.2 1866. 1885.2 1837.1 1898.3 1833. 1799.1 1789.8 1800.8 1839.2 1819.6 1904.4 1808.7] | [1351.5 1317.9 1355.9 1314.7 1322.5 1345.4 1369.7 1306.3 1315.7 1348.6 1319.8 1338. 1362.4 1330.8 1336.3 1354.2 1320.4 1386.9 1358. 1319.5 1287.7 1329.6 1384.1 1310. 1369.7 1359. 1355.7 1351.8 1357.1 1297. 1279. 1377. 1380.1 1378.5 1346.9 1377.7 1361.3 1296.7 1369.3 1350.7 1324.9 1347.3 1297.2 1350.9 1298.7 1360.9 1336.9 1325.6 1383.3 1395.6 1338.8 1350. 1368. 1323.1 1289.7 1315.6 1328. 1333.4 1321.2 1348.1 1318.5 1370.9 1386.6 1334.9 1323.3 1317.4 1319.8 1385. 1334. 1360.8 1330.3 1304.8 1326.3 1324.4 1355.9 1337.8 1349.2 1324. 1338. 1378.6 1306.6 1345.6 1355.3 1364.1 1316. 1297.2 1367.9 1331.8] | [914.6 926.7 919.2 896.8 912.1 906.5 921.7 905. 907.8 914. 917.5 941.7 913. 905.3 935. 913.8 923. 899.9 911.4 916.6 931.2 900.2 909.2 931.7 920.8 913.9 927.6 909.5 925.4 930.6 926.6 902.8 898.1 943.3 924.8 908.3 949.1 907.3 922.8 951.9 933.8 926.8 912.5 914. 917.8 905.4 918.8 925.4 918.4 923.6 902.7 913.6 934.2 909.4 931.1 923.7 923.2 913.3 921.3 908.5 902.8 932.5 938.3 905.6 929.6 925.8 933.9 898.4 915.7 935.4 928.6 915.6 920. 923.6 906.6 914.8 942.1 945.6 945. 936. 902.4 926.3 909.2 919.2 912.4 915.5 905.1 912.1] | [687.2 666. 659. 671.5 664.2 670.6 679.2 658.9 676.2 658.6 675.8 678.5 663.5 671.7 671.8 666.2 657.9 674.1 672.1 669.9 675.3 661.3 682.2 657.8 673.4 671.7 669.7 659.9 663.2 663. 672.8 681.4 667.4 664.2 679. 663.4 694.3 670.6 668.9 668. 666.2 675.4 674.4 674.9 683. 660.8 668.9 672.2 671.6 667.8 672.4 664.7 669.2 671.5 679.7 658.2 664. 677.2 674.1 667.5 669.1 675.6 680.3 668.1 666.6 667.6 671.2 665.9 675.4 673.2 660.4 666.9 667.9 662.6 656.3 671.4 674.3 672.9 669.1 680. 676.9 665.4 660.9 668.2 667. 661.2 673.3 678.1] | [468.8 479.1 470.8 484.4 480.4 464.8 477.7 480.9 470. 473.2 462.3 469.1 486.3 463.2 484.2 468.2 464.4 469.1 480. 469.5 468.4 473.5 468. 484. 462.6 472. 471.3 470. 474.4 475.7 478.4 471.3 481.6 473.6 481.1 481.6 471.8 473. 481.3 469.5 478.3 469.3 468.6 465.5 470.6 473.1 476.5 475. 475.4 473.2 472.7 475.8 478.4 467.9 466.8 473.7 478.1 459. 461.8 471.6 470.3 479.1 478.3 477.7 464.8 467. 477.1 473. 479.2 480.6 474.8 476.2 481.7 464. 470.1 467.4 478.6 471.2 466.9 481.9 474.1 468.7 478.7 478.1 467.3 467.3 476.9 482. ] | [297.3 297.3 300.3 302.7 301.8 309.8 304.8 310. 311.2 311.5 322.3 310.2 315.1 307.6 299.7 298.4 303.5 295.7 305.8 304.2 299.3 305.2 299.4 306.1 316.6 305.5 309. 315.2 304.9 311.4 302.4 307.6 310.5 311. 307.1 306.2 310.4 305.7 311.2 306. 313. 313.4 308.1 314.2 302.5 306.1 310.9 311.9 307.1 306.9 300.3 307.4 309.9 306.4 314.3 307.7 302.5 308.6 313.8 307.4 309.8 305.4 306.1 317.4 308.1 317.6 302.5 308.5 315.6 310.5 303.8 315.4 313.7 310. 314.9 309.5 293.2 301. 307.2 305.3 305.9 307.5 303.7 310.9 315.9 305.7 307.4 313. 305.5 301.6 304.8 311.9 305.7 313.4 309.1 307.2 309.2] | [236.1 239.6 238.9 238.7 238.7 241.6 239.7 244.6 245.9 242.4 245.2 244.2 247.9 239.3 241.6 237.4 237. 235.8 238.4 233.9 244.3 231.7 247.4 247.1 234.7 246.2 246.3 244.8 242.3 240. 244.3 244.5 232.5 252.3 244.1 239.4 236.3 238.6 242. 250.5 239.7 245.9 239.7 242.8 238.2 241.5 240.2 243.3 240.3 248.5 243.6 244.7 238. 245.6 237. 238.6 239.8 242.4 244.9 239.8 238.1 240.5 240.1 244.2 248. 241.8 236.7 249.2 248.3 237. 240.5 246.4 234.2 238.5 244.2 234.3 240. 243.1 241.3 241.3 243.8 239.6 233.7 236. 243.2 242.9 243.6 241.2] | [162.4 161.9 157.2 162.9 159.3 165.5 154.2 162.1 162.4 159.5 164. 160.1 162. 161.4 163.3 168.4 170.5 167.1 164.2 158.1 167.9 158.3 164.5 164.8 163.7 158.8 158.9 159.6 163.5 157.1 159.7 158. 155.1 164.3 159.7 158.6 160.7 159. 164.5 169.1 166. 162.8 171.2 161.8 160.2 165.9 159.9 156.2 164.5 166.3 166.1 163.4 168. 163. 150. 164.2 162.9 160.9 161.3 157.7 160.5 161.6 166.9 156.6 160.4 160.2 158.9 160. 159.8 161.6 163.3 163.2 165.2 160.2 161.3 169.8 166.1 171.2 162.3 165.2 168.5 166.1 160.3 160.3 165. 159.5 160.3 165.6] | [116. 110.2 115.6 118.9 116.3 117.2 116.9 121.8 112.6 125.7 118.4 125.7 128.9 126.6 122. 118.3 114. 116.3 116.6 110.7 119.4 112.3 117.1 117.4 117.1 126.5 125. 130.3 123.4 118.1 127.5 121.1 117.2 114.3 118.9 127.3 110.6 115.3 123.7 118.3 121.1 121.2 125.7 123.4 107.7 123.3 118. 120.6 116.3 115.6 127.1 111.9 122.9 116.1 117.8 109.3 108.7 122.7 118.2 117. 125.1 112.8 121.1 120.3 116.6 121.6 125.5 120.1 126.5 121.6 123.4 123.8 135.4 120.3 115.3 133.3 119.3 119.9 121.7 114.5 123.8 118.4 131.5 126.7 119. 118.6 117.4 130. ] | [ 10.6 83. 85.9 98.9 83. 2.4 71.7 66. 93.6 65. 110.2 88.3 89.6 98.8 90.8 92.2 71.6 68.4 82. 86.6 73.8 88. 68.6 103.1 108.2 68.9 80.2 91.2 95.7 80.3 78.5 89.2 3.4 75.6 93.4 72.3 99. 75.6 75.1 94.9 79.8 79.6 2.8 7.3 78.4 10.6 75.2 79.1 96.9 89.3 7.6 85.4 82.5 77.7 3.6 83.3 74.8 79.5 51.7 75.5 2. 80.2 79.6 94.6 3.4 91.1 3.7 104.7 71.9 84.2 88.3 85.4 90.4 86.5 101.7 88.4 76.1 71.1 64.2 81.9 4.6 80.5 106.6 94.4 88.5 99.9 108.3 94.6] | [72.8 2.5 69.1 10. 1. 4. 4.4 4. 2. 4.6 9.1 6.7 5.6 2.6 4. 74.6 3.5 85. 10. 1. 10.6 2.5 9.5 68.4 4. 1.6 2. 1.1 76.1 84.5 1.2 11. 1.9 91.3 9.5 2.1 69.3 63.4 1.4 1.2 76.2 3.3 59.6 10. 3.1 2.5 71.5 9.5 9.5 1.2 3.3 77.2 1.1 92.4 52.4 54.3 81.8 67.3 73.6 2. 1.2 5.7 90.6 6.2 2.6 5.7 5.7 9.5 1.2 5.2 5.3 69.5 7.5 44.6 2.6 3.1 2.4 69.4 4.1 2.8 51.4 4. 10.3 1.8 9.5 70.2 7. 2. 10. 2. 9.5 3.7 10.2 3.3] | [ 81.1 95.8 88.7 100.5 101.7 142.2 10. 4. 48.1 1. 5.2 3.6 3.4 88.1 3.3 98.1 99.5 63.8 76.4 79.4 10. 2.7 1.2 92.8 98. 3.2 2.4 3.8 1.1 1.8 1.2 4. 82.1 6. 7.3 74.5 63.3 63.2 8. 3.2 4.5 8. 2. 93.4 4. 54.3 8. 3. 2.9 1.5 2.7 3.2 87. 95.7 4.4 3.4 9.5 6. 3. 10. 2. 73.6 5. 6.7 2.4 9.5 5.7 2.4 7.1 2.1 76.2 4.6 41. 6.2 1.1 10. 1. 9.5 48.2 77.4 3. 42. 85.8 79.6 2.7 2. 6. 1.1 5. 4. 68.2 83.7 3. 10.2] | [ 7.5 133. 128.8 101.1 106. 3. 10. 1.3 2.8 9.5 4.6 3.4 4.5 4.5 3.4 3. 10.3 115.7 127.5 3. 101.3 4. 121.6 3.2 4.9 76.8 76.7 71.4 63.8 10. 131.2 2.7 86.7 5.6 5. 10.7 4.3 100.6 2.6 2.1 126.2 115.4 4.5 8.6 10. 3.7 2.6 9.5 10. 2.6 9.5 3.4 129. 9.6 121.2 2.1 9.5 97. 76.5 10.6 9.5 9.5 76.4 123.1 9.5 4.1 5. 10.7 94. 58.4 4.4 128. 80.1 103.2 9.5 112.8 7.1 9.5 2.8 9.5 4.1 130.6 5. 112.1 4. 72.7 4. 95.4 105.5 4. 9.1 3. 10.6 5. ] | [164.8 4.2 55.8 2.8 1.6 1.1 1.7 145.1 190.3 182. 9.5 5. 59.3 3. 4. 1.3 5. 4. 275.2 345.8 310.4 158.7 162.6 130.3 129.7 3.2 100.9 5. 88.6 7. 92. 141.3 1.9 1.2 1.6 276.6 2.1 126.7 2.3 108.9 8.4 3. 253.8 171.7 9.5 6.3 70.5 10.6 154.1 4.2 288. 183.5 89.5 9.5 9.5 7. 156.5 8. 71.1 325.3 1. 8.5 101. 2. 5. 188.2 120.6 7.7 10. 73.7 5.5 209.6 10. 10. 6.3 8.6 1.8 2.1 5. 9.5 2.3 338.1 5. 4.4 207.2 127.2 253.7 10. 2.7 6.6 2.1 119.2 3.7 5. ]
6.011 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Cologne_MR4_NIST_Magnitude_T1map.nii.gz | Magnitude | Kilian Weiss | Universitaetsklinikum Koeln | Philips | Ingenia 3T | 3 | NIST | 130 | 23 | 22 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [2041.9 2045.1 1956.4 1890.1 2035.3 2068.7 2103.7 2094.3 1982.4 1851.7 1982.1 2109.3 1933.8 2037.8 2054.9 1984.5 1970.1 1931.2 2002.6 2045. 1949.3 1926.9 1926. 1941.3 1929.6 1899.6 1935.2 1936.6 2025.6 1993.3 1930.8 1973.7 2048.1 2097. 1936.7 1857.1 1897.1 1993.7 1979.4 1972. 1940.7 1885. 2004.1 2030.5 2209.9 2078.7 2026.7 1959.4 1909.6 1977.6 1965.7 1904.4 1948.5 2008. 2042.4 2073.7 2079.8 2087.6 2029.7 1930.4 1915.8 1939.6 1895.2 1968.2 2033.6 1956.8 2024.9 1963.5 1905.3 1998.5 2087.5 2034.6 1965.8 1938.5 1915.2 1985.3 2143.3 2120.7 1882.4 1910.9 2057.9 2024.9 2063.1 2048.2 1891. 2022.3 2156.9 1927. 2003.1 1985.9 2016.1 2014.5 2049.8 1958.1 1974. 2042.8 2033.6] | [1421.7 1438.5 1447.5 1412.7 1399.6 1420.5 1437.9 1400.7 1421.7 1416.7 1413.2 1414.3 1381.2 1406.7 1379.3 1499.8 1414.8 1435.9 1454. 1464.9 1451.2 1377.9 1400. 1430.3 1359.4 1432.7 1425.5 1446.6 1432.1 1424.4 1395.9 1386.7 1447.3 1436. 1373.4 1426.6 1419.9 1428.3 1393.8 1393.6 1406.3 1402.8 1439.2 1459.1 1400.4 1424.4 1416.9 1436.7 1432.7 1389.6 1416.1 1387.2 1405.1 1439. 1360. 1419.1 1444.9 1474.1 1397.8 1386.7 1421. 1363. 1415.1 1415.3 1386.4 1427.2 1430.3 1478.5 1435.9 1436. 1437.8 1396.7 1444.1 1418. 1386.4 1370.5 1462.3 1437.7 1437.1 1416.6 1425.4 1443.2 1415.9 1457.9 1474.6 1429.7 1428.2 1448.9] | [955.9 973.4 978.3 947.7 923. 940.3 955.2 970. 969.4 967.8 974.8 964.4 938.6 949.3 953.9 964.7 954.8 953.9 973.4 987.3 965.9 982.3 945. 946. 947.9 942.5 953.7 960.8 964. 972.5 957.1 979.7 946.7 965.6 961.4 938.8 951.2 964.1 965.7 969.5 961.4 970.7 961.7 971.8 976.8 947.5 945.9 975.9 960. 963.3 972.8 971.8 946.9 938.6 977.4 954.7 948.1 962.4 953.2 937.7 980.4 968.8 946.9 955. 956.4 987.6 950.8 948.2 970.2 978.4 972.2 944.1 958.8 963.8 965.9 955.9 965.7 974.5 972.8 978.2 975.9 989.1 978.6] | [683.9 708.1 709.5 710.7 703.8 699.3 708.2 694. 712. 700. 701.1 696.4 691.4 703.9 689. 697.3 695.8 703.3 700.2 712.5 710.1 695. 699.6 697.5 697.4 681.6 701.7 701.4 696.1 705.1 700.2 685.2 701.3 697.6 707.8 689.6 699.9 689.6 690.3 689.7 691.5 687.5 690.2 695.3 702.1 699.3 696.1 699. 697.7 694.3 703.7 681.9 691.5 702.1 702.4 694.4 694.5 694. 707.9 711.1 714.3 689.1 703.3 696.6 709.1 691.9 697.2 695.1 693.3 709.2 703.2 690.5 705.8 702.5 694.6 702.5 703.2 695.8 695.6 695.7 695.7 691.5 694.3 708.1 708.8 702.3 703.2 702.1] | [492.7 487.2 489.7 493.9 498.2 487.5 493.5 495. 491.6 494.9 506.1 496.7 487.3 480.5 486.8 493.6 492.8 491.4 496.9 503.4 501.1 498.7 495.5 483.3 496.7 494.9 495.2 493.5 496. 494.7 498.8 498.1 499.3 492.9 492.6 490.3 489.8 486.6 491.7 489.3 493.2 494.3 495.8 486.2 489.7 492.1 487.3 485.7 486.8 493. 492.9 495.9 501.2 490.8 490. 497.6 490.5 493.4 492.3 490.5 495.2 495.4 495.6 484. 487. 498.5 484.4 484. 488.9 480.5 486.3 486.1 486.4 492.7 484.3 490.3 482.7 489. 486.6 489.6 485.8 493.6 487.7 485.8 487.9 491.5 488.6 493.9] | [317.9 327.2 325.6 316.8 319. 325.6 319.8 310.7 305.1 311. 311.8 314.4 323.3 311. 316. 309.5 308.6 313.4 315.9 318.6 319.5 304.5 316.2 310.1 315.4 314.7 310.5 309.9 312. 313.5 318. 310.8 308.5 315.2 306.9 309.2 309.6 308.5 309.1 306.2 309.1 312.3 309.5 314.5 316.3 309.7 307.2 307.1 308.8 313.3 312.5 307.8 308.4 307.8 309.6 312.2 308.3 311.8 316.8 309. 310.6 313.4 309.8 309.3 309.8 306.3 310.8 306.8 311.7 312. 305.9 311.4 315.5 312. 312.8 312.3 315.5 310.4 312. 307.8 306.8 311.3 314.3 311.2 312.3 307.2 308.8 310. 314.6 312. 309. 307.8 305.4 310.2 317. 314.1 311.7] | [246.5 245. 243. 243.9 240.9 238.3 237. 234.8 239.3 244.9 247.2 235.9 241.5 241.4 240.1 239.3 239.3 241.4 246.9 236.4 233.8 239.4 239.8 241.5 243.1 242.8 241.4 243.7 245.7 235.6 240.5 244.9 242.8 243.2 240.1 242.4 242.7 239.5 237.6 240.2 240.6 247.1 243.3 237.1 241.6 240.3 239.6 241.4 234. 238.4 238.1 241.4 240.5 241. 241.2 233.2 241.7 243.1 235.8 240.7 238.9 239. 239.3 243.9 241.7 235.8 239.9 240.6 235.8 238.7 239.8 242.9 239.4 243.7 244.2 238.9 239.3 241.2 247. 245.3 243.8 240.3 236.2] | [163.2 156.1 240.1 238.9 160.9 162.2 164. 155.9 250.2 242.9 249.2 244.2 242. 244.9 155.1 159.3 241.1 241.5 164.2 235.5 235. 241. 217.6 238.6 255. 160.2 238.6 242. 235. 238.7 239. 239.4 225.4 229.8 259.1 160.9 259.4 244.1 238.7 244.3 162.6 233.8 229.6 243.1 240.1 155. 256.9 244.2 244.1 161.3 226.4 238.1 159.6 236.7 238.3 154.5 236.1 241.9 240.2 233.4 157.6 161.9 233.5 237.3 228.3 157. 248.3 233. 232.5 231.5 231.8 239.8 163.7 164.4 151.5 157.3 261.1 237.5 236.5 231.5 238.5 238.3 235.8 232.8 156.6 156. 156.4 155.3] | [121.4 129.5 115.5 133.2 128.3 137.4 144.4 144.7 107.9 134.9 113. 113.5 118.3 131. 138.1 112.8 146.8 122.2 122.2 112.9 112.6 121.5 139.8 135. 140.8 117.5 152.6 119.1 118.5 116.7 113.2 142.6 144.4 116.2 159.7 107.1 141.5 133.2 115.7 134.6 140.2 142.8 146.5 121.6 135.3 113.9 134. 110.8 118.5 121. 121.4 123.5 120.3 138. 104.3 130.7 132. 131.2 134.7 140.4 136.8 139.1 140.8 148. 132.4 141.1 105.5 113.1 113.2 133.4 133.6 129.6 129. 104. 130.4 150.6 134. 119.4 142.7 139.5 134.8 133.8 98.6 106.9 127.6 146.3 135.7 120.5] | [ 87.7 1.2 72.1 80.5 91.6 92. 61.4 2. 54.5 80.5 90. 87.5 84.5 89.1 73.6 4. 59.7 77.3 9.5 57. 66.5 54.2 63.4 1.1 2.7 71.9 3.3 75.8 72.6 82.3 71.9 57.3 84.6 76.4 4.9 1.6 7. 74.4 67.9 79.5 70.1 81. 88.9 97.8 4. 1.2 3.3 72.5 65.8 75.3 79.1 98.7 95. 80.9 2.7 91. 7. 1.8 61.2 74.9 84.1 95.9 81.9 9.5 75.8 2.4 1.8 9.1 92.9 84.5 92.8 89.1 64.1 116.7 3.6 111.4 3.6 80.5 69.5 75.6 71.6 3.4 4.6 103.5 111.8 90.3 64.9 78.5] | [74.1 1.3 40.2 2.5 4.9 4.4 9.5 1.2 1.6 1.4 1.1 92.8 5.3 1.2 2.6 2.7 3.1 1.3 1.2 4. 3.7 2.6 3.2 10. 3.3 5. 2.7 1.8 2.6 2. 5. 2.1 1.4 9.5 2.5 2.6 2. 2.7 3.4 1.3 1.6 1.6 5.3 1.2 9.5 5. 9.5 2.3 2.5 87.8 2.1 4.5 5. 3.1 57.2 1.2 2. 1. 2.7 2.5 1.8 1.7 5. 9.5 9.5 2.6 2.8 9.5 1.3 1.8 10. 1.3 1.8 4.3 1.6 3.3 2.3 9.5 74.7 62. 2.1 3.6 1.3 9.5 1.2 3.2 2.7 1.8 10. 1.1 1.1 1.4] | [ 9.5 83.8 42.7 2. 3. 60.7 2.5 1.8 10. 2.3 4.1 9.1 9.5 9.5 2. 105.5 1.9 1.3 2.6 2. 4.4 2.4 2.8 1. 2.4 4.9 5. 1.5 3.3 1.6 3.3 1.9 69.7 2.5 2.8 5. 2.7 1.2 9.5 9.5 2.7 4.1 3. 3.1 2.1 1.4 4.4 2.8 3.1 1.2 7. 1.3 2.7 4.6 7.8 5. 3. 9.5 2.4 1.8 5. 1.3 3.7 1.1 2.1 1.1 9.5 5. 2. 3.3 1.2 4. 2. 9.5 8.3 9.5 2.4 2. 9.1 10. 4.2 2.4 2.7 1.6 1.3 1.8 5. 1. 3.7 1.2 4. 1.2 2.8 4.2] | [ 4.7 6.3 9.6 2.8 1.8 152.2 1.3 67.3 10.3 1.3 7. 1.1 79.2 163.1 2.8 9.5 1. 1.2 3.4 1.8 5.9 1.1 2.1 1.1 1. 5. 1.3 2.7 8.5 3.4 3.3 1.2 1.3 1.9 3. 2.6 4.3 2.5 9.5 6.3 1. 5. 2.6 4.9 7.7 7.3 2.5 9.3 4.3 5.7 1. 1. 1. 1.1 1.3 4.2 1.1 1.2 1.2 2.8 4.3 3.2 9.5 4.1 2.7 1. 1.8 3.8 3.3 1.1 6.1 2. 1.2 3.1 3.3 4.1 2.5 4.6 1.2 1.1 7.1 2. 5.7 1.3 8.1 69.6 1.9 63.5 4. 2.8 3.2 2.7 2.8 4.5] | [198.9 166.7 99.7 3.3 64.5 1.2 2.6 3.3 1.9 38.1 9. 100.7 41. 1.8 6. 7.3 7.2 180.3 1. 9.6 4.1 3.3 1. 1.2 8.6 1.2 6.3 1.8 10. 9.3 4. 69.4 76.3 2.1 9. 1.4 1.1 1.1 5. 5.2 9.8 2.2 3.1 10.6 1. 1.2 8.2 1.1 9.3 3.4 2.5 8.2 3.1 6. 4.4 5. 6.3 53.2 2.5 1.8 1.2 5.6 9. 10. 8. 5. 3.2 1.2 5. 2.5 4.2 1.3 66.1 1. 4.3 1.3 6.3 52.9 3.2 5. 1.3 1.2 3.3 2.1 3. 77.9 68. 79.3 4.7 5.6 1.3 2. 1.4 67.9]
6.012 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Cologne_MR4_NIST_Complex_T1map.nii.gz | Complex | Kilian Weiss | Universitaetsklinikum Koeln | Philips | Ingenia 3T | 3 | NIST | 130 | 23 | 22 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [2061.7 2128. 1970.8 1813.2 2080.7 2083.2 2110.8 2145.7 2014.2 1770. 1945.2 2174.3 1917.3 2091. 2057.2 1924.9 2003.7 1889.3 2011.6 2068.6 1906.6 1941.3 1883.9 1935.5 1909.4 1857. 1934.7 1909. 2079.1 2024.5 1905.7 1975.5 2065.7 2140.4 1864.1 1789.7 1838.7 2001.1 1997.3 1970.5 1930.9 1841.8 2001.4 2036.8 2331.6 2075.6 2043. 1940.5 1857.3 1958. 1965.3 1889.6 1969.3 2018.7 2076.3 2047.3 2121.1 2146.4 2019.2 1857.3 1859.8 1934.8 1856.8 1993.3 2071.4 1908. 2027.1 1903.1 1813.7 1974.1 2136. 2066.8 1931.9 1923.9 1887.5 1970. 2083. 2075.7 1820.7 1831.6 2098.7 1984. 2081. 2137.3 1769.4 2014.6 2092.2 1855.6 2023. 1926.6 1991.9 2001.9 1979.5 1924.3 1923.6 2036. 1988.3] | [1426.7 1434.2 1466.8 1413.4 1367.6 1427.3 1426.6 1385.1 1418.8 1400.9 1389.3 1426.5 1349.3 1392.7 1381.2 1506.8 1401.6 1436.6 1473.7 1478.5 1500. 1357.2 1392. 1429.5 1363.6 1419.9 1428.7 1464.3 1417.1 1432. 1372.2 1379.5 1489.4 1426.6 1370.4 1430.1 1411.4 1425.8 1372.5 1376.3 1404.4 1415.5 1467.4 1484.9 1420.5 1432. 1380.4 1437.6 1437.3 1374.4 1427.4 1382.5 1403.1 1459.2 1368.6 1409.1 1454.9 1503.7 1377. 1367.1 1419.7 1338.8 1422.4 1419.1 1388.3 1436.3 1418.8 1515. 1420.6 1461. 1435.1 1379. 1466.9 1423.7 1390.1 1342. 1455.5 1421.5 1442.8 1402.4 1416.2 1433.1 1388.4 1444.7 1469.2 1390.1 1406.9 1443.5] | [ 955.1 968.4 971.5 958.9 924.1 945.7 946.5 973.5 972.2 978. 983.2 972.5 951.4 954.8 955.5 971.1 954.4 947.7 993. 999.8 966.8 1019.7 938.3 939.3 947.7 935.4 962.4 961.7 964.9 978.5 949.8 996.9 940. 982.8 974.2 937.5 960.8 965.6 973.1 970.3 955.3 977.7 965. 981.7 995.3 933.7 943.2 996.6 958.1 976.1 969.4 980.8 944.1 924.9 1001.3 961. 947.8 974.2 955.8 931. 987.1 973.1 949.9 959.2 954.2 1006.6 958.5 933.5 984.1 996.4 983.1 933.1 964.9 965.3 974.3 946.7 977.3 993.1 973.5 969.5 973.3 990.2 984. ] | [688.8 709.1 707.7 716. 703. 690.1 722.8 693.3 714.3 697.2 695.3 689. 685. 694.4 693.4 704.8 694.1 702.4 697.1 721.3 720.7 690.9 701.3 693.7 693.5 668.9 703.3 702.7 696.4 712.4 697.6 677.3 709.6 696.4 713.4 683. 701.1 683.7 687.2 685. 684.9 683.2 689.3 694. 702.6 710.1 694.5 700.2 699.9 690.5 709.6 672.1 685.2 712.1 702. 695.4 694.4 694.6 712.5 718.2 726.8 683.7 712.2 693.4 720.1 692.2 700.3 698.3 687.7 718.1 707.4 685.6 714.7 700.7 693.9 706.4 712.5 697.1 694.3 691.9 693.6 684.9 688.3 706.4 714.1 702.9 704.9 705.2] | [491.5 481.6 486.3 491. 503.4 484.1 489.9 489.6 485.9 486.1 508.4 494.8 480.7 475.6 482.6 488.8 491.4 484.7 490.7 499. 493.7 494.4 496.2 481.1 493.1 488.4 489.7 484.2 492.2 485.4 493. 490.1 500.2 493.3 485.2 482.1 484.7 477.3 485.4 479.1 487.2 486.2 489.7 475. 485.2 489.5 478.7 473.4 475.8 485.6 484.5 486.1 501.9 489.2 488.1 499.6 488.2 494.6 489.1 485.8 493.4 490.7 494.6 467.1 488.5 503.6 476.3 478. 489. 466. 479.4 477.7 477.3 487.1 478.2 492.1 479.4 489.7 485.8 487.4 479. 489.3 489.6 484.1 487.2 491.9 481.5 492. ] | [314. 327.2 324.5 313.6 319. 324.1 319. 307.2 302.8 309.9 306.3 314.5 321.1 306. 318.3 306.1 306.3 316.2 315.9 320.3 318.1 302.5 315.2 306.7 317.7 317.2 311. 307.9 313.5 310.5 320.8 307.5 307.3 315.2 302.2 309. 308.8 308.9 307.7 303.4 307.2 313. 306.7 315.6 319.8 311.4 307.1 304.4 307.6 314.9 314. 306.7 308.1 305. 310.4 312.1 307.8 314.9 323.3 307.4 309.4 314.9 308.3 310.8 308.6 305.2 311.1 304. 313.5 314.7 304.9 311.7 317.1 311.6 313.4 314.6 317. 313.2 316.6 304.2 303.9 310.7 316.9 310.6 315.6 307.4 309.7 310.2 317.1 311. 307.8 307.5 304.7 307.8 318.8 310. 308.6] | [245.3 243.3 240.4 241.3 239.7 234.4 237.2 229. 235.7 244. 248. 233.8 245.1 240.3 240.8 239.7 237.8 241.7 245.9 236.4 233.7 237. 238.4 241.4 243.6 244. 238.7 244.5 245.5 234.6 243. 248.1 243.1 244.9 238.8 242.3 243.1 238. 237.1 242.8 241.7 252.2 244. 234. 240.7 240.6 239.1 245.7 232.2 238.9 237.7 241.8 239.9 242.5 243.1 229.2 241. 247.5 237.1 240.5 239.5 239.8 238.2 247. 241.6 231.4 243.8 240.7 236.1 239.5 235.9 243.3 234.5 242.7 249.8 236. 240.9 242.2 250.9 247. 244.2 240.2 234.8] | [164.9 154.4 163.4 162. 157.1 159.3 163.9 156.1 167.2 161.1 169.4 166.4 166. 162.6 155.8 159.1 164.2 167.7 161.9 159.7 156. 165.6 151.6 163.2 171.8 158.1 161.9 164.9 160.1 162.3 162.4 164.1 154.3 157.7 175.3 163.1 173.9 163.9 162.5 166.5 163.4 158.3 157. 168.8 161.8 153.5 169.5 163.3 164.5 158. 149.8 165. 158.5 158.7 160.4 153.2 157. 165.6 163.1 159.5 158. 163.9 158.2 160.1 157.8 158.2 164.8 162.8 155.6 156.5 157.4 163.4 162.1 167.7 150.5 156.4 167.2 159.8 160.7 157.4 164.3 161.7 162.1 158.6 156.7 154.4 155.1 152.5] | [119.6 115.7 113.6 121.3 111. 124. 127.2 121.8 112.8 120.1 111.3 114.7 118.3 114.1 117.6 115.4 125.4 127.6 128.7 112.9 109.7 123.8 120.9 119.9 121.4 111.5 129.3 120.9 120.3 120.4 110.7 126.1 124.3 113.9 135.9 110.5 122.3 117.9 116.2 115.7 124.1 124.2 126.9 124.7 119. 118.4 115.5 110.7 123. 122.8 120.4 123.8 115.5 115.7 103.9 114. 116.3 114.4 118.9 127.5 118.1 122.8 123.1 120.7 113. 120.5 97.7 110. 112.4 113.6 117.1 111. 110.7 103. 113.8 129. 115.8 121.3 125.1 120.4 117.1 115.2 98.3 96.6 113.9 127.5 119.2 103.4] | [ 89.9 3.7 64.1 73.5 89. 86.4 83.6 1.8 48.6 80.5 91.4 87. 84.2 87.8 82.2 67.3 77.5 82.5 1.1 1.2 10. 1.4 59.5 6.7 1.8 89.1 65.4 80.1 69.3 90.5 66.4 10.1 87.4 71.5 10. 69.7 56.1 78.1 57.2 73.3 1.2 64. 81. 100.7 2.6 71.3 6.4 81.9 10. 78.7 67.2 100.5 89.8 52.9 1.7 98.2 65.5 1.2 9.1 72.2 65.3 93.6 73.6 6. 75. 72.4 59.8 1.1 78.4 79.6 95.2 86.7 42.1 108.8 1. 114.7 3. 83.4 2.6 72.6 73.6 2. 10. 100.4 107.2 88.9 3.1 85.9] | [76.9 9.5 67.4 4.8 3. 7.1 2.4 6.3 8.8 3.4 10. 76.5 5. 10.6 3.6 8.6 55.5 9.5 71.5 9.5 6.9 1.2 10.7 70.3 47.6 6.1 1.8 3.3 6.6 10. 2.6 7. 4.5 7.5 79.6 3. 2. 3. 9. 10. 58.4 9.5 5.5 9.5 10. 4. 6.5 1.3 78.9 93.1 5. 10. 57.6 55.6 89.8 67.7 81.6 3. 5.4 5. 9.5 47.6 9.5 5.9 10. 3.3 3.4 9.5 2. 6.3 3.9 1. 2. 9.5 6.9 9.5 2.5 8. 92.1 9.5 5. 9.5 1.3 4. 10. 1.8 1.3 2.4 67.6 3.3 6.4 10. ] | [ 5. 99.8 82.3 10.3 81.5 2.1 3. 2.3 2.7 10. 1.4 4. 9.5 3.3 6.3 88.1 3.4 6.5 1.2 84.8 1.3 2. 7. 65.6 2. 5. 9.5 4.9 7. 2. 10.2 9.5 62.6 9.5 3.7 4.4 5. 1.2 6.4 86.3 4.7 3.3 2. 2.3 4.4 10. 6.9 7. 10. 7.6 10. 8.6 5. 2.8 1.2 1.4 1.1 10. 3.6 9.5 2.5 1. 10. 2.3 4.9 4.7 9.5 10. 1.8 3.5 7.7 6.3 7.2 3.9 1. 3.3 6.1 44.6 1.2 2.1 1.3 2.5 9.5 10. 1.7 9.5 1. 9.5 5. 3.1 4. 1.8 10.2 6.4] | [ 2. 9.5 1. 4.9 3. 148.6 8. 142.8 10. 6.1 2. 2.1 75.2 152.6 1.3 10. 9.5 10. 1.2 2. 9.5 6.9 1.5 2.7 116.1 7.3 73.8 1.9 3. 2.5 8. 1. 2. 8. 7.3 3. 3.2 3.6 4. 10. 7.1 5. 2. 7. 2.1 8.4 1.6 7. 9.6 2.8 49.5 96.6 10. 10. 3.3 9.2 1.2 4.7 7.7 4. 6.7 2.8 2.3 7. 7.6 2.1 7.2 3. 7. 7.8 2.8 10.3 5.4 10. 9.1 9.5 1.2 2. 4.9 9.5 3.3 8. 9. 10.3 8. 87.3 6.9 137.3 70.5 51.2 6.1 10. 4.9 10. ] | [196.4 171.7 80.2 4. 172.8 8.2 10.3 1.3 8.5 157.5 2.8 80. 101.2 2.5 5.4 102.9 2.1 307.3 7.7 78.4 72.1 9.6 5.6 2.6 8.8 7. 3.1 4.3 3.2 131.6 3.6 254. 250.7 120.6 1.3 5.3 5.7 2.6 9.4 2.8 91.1 5. 6.7 8.2 10.7 3.2 9.5 10. 6. 4. 2.6 113.1 4.3 155.7 10. 6.7 93.7 197.3 153.9 6. 10.6 5.6 10.6 124.9 8. 7.1 9.3 10.6 3.5 9.5 138.7 8.6 179.6 152.7 10. 8.5 8. 164.2 2.4 7. 4. 5.3 127.7 7.6 3. 309.7 144.4 126.3 1.6 5.4 8. 2.7 3.2 140. ]
6.013 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Hamburg_MR1_NIST_Magnitude_T1map.nii.gz | Magnitude | Christian Stehning | Philips Research Hamburg | Philips | Ingenia DDAS 3T | 3 | NIST | 130 | 23 | 20.8 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1946.2 1880.9 1872.1 1875. 1864.7 1933.9 1894.3 1883.8 1882.8 1877.1 1893.7 1904.6 1848. 1897.6 1867.7 1887.2 1876.3 1882.2 1877.3 1882.9 1884.5 1825.9 1895.7 1887. 1851.4 1835.7 1845.3 1903.8 1857.9 1875. 1891. 1850.4 1818.6 1863.8 1875. 1868.5 1859.1 1849.8 1869. 1838.9 1852.8 1870.2 1853.6 1835.4 1850.4 1864.5 1850.5 1840.2 1856.2 1842.9 1851.3 1857. 1870. 1877. 1838.9 1868.6 1883.3 1850.8 1857.1 1853.4 1821.4 1848.3 1829.5 1852.3 1886.2 1850.2 1814. 1867.9 1878.9 1855.2 1839.2 1853.1 1847. 1816.6 1839.3 1876.4 1826.2 1868.6 1881.1 1840. 1856.9 1861.1 1847.4 1856.4 1883. 1852.9 1854.9 1921.4 1916.3 1877.9 1873.3 1869.3 1867.4 1888.5 1867.1 1869.1 1878.3] | [1332.3 1377.7 1386. 1372.5 1393.7 1412.3 1346. 1385. 1372.2 1361. 1361.8 1361.5 1384. 1412.9 1354.7 1377.9 1367.5 1356.7 1377.4 1370.6 1364.4 1382.1 1402. 1402.7 1324.1 1374.4 1346.6 1369.4 1383.1 1360.6 1365.4 1369.5 1379. 1388.6 1376.6 1355.3 1353.6 1372.6 1373.6 1355. 1362. 1372. 1373. 1370.2 1379. 1381.3 1362.9 1368.6 1386.9 1364.3 1354.1 1361.9 1364.6 1371.3 1352.8 1390. 1361.2 1362.5 1382.7 1369.1 1375.7 1361.1 1348.8 1362.6 1366.7 1392.9 1369.2 1358.9 1362.4 1364.7 1372.2 1364.1 1340.5 1345.8 1419.4 1377.3 1370.4 1371.5 1377.5 1379.1 1368.8 1348.9 1386.1 1377.5 1387.4 1392.5 1398.8 1379.1] | [953. 967.3 936.1 925.9 937.8 966. 949.6 945.3 926.4 920.2 933.7 944.2 950.9 926.8 929.6 929. 926.8 926.3 924.4 929.4 932.6 928. 920.4 932. 923.4 923.3 920.6 920.1 919.7 923.4 940.6 920.8 926. 926.3 924.3 922.2 925.6 923.3 917.6 934.7 933.6 925.6 926.3 936. 927.5 935.8 922. 932.2 935.8 932.3 928.4 928.8 931.9 920.9 923.6 933.2 937.4 943.1 940.8 930.8 938.1 939. 928.4 930.7 929.9 922.1 940.1 942.9 926.6 942.4 934.4 935.5 926.3 929.6 928.1 942.2 940.6 943.5 944.8 932.1 926.5 932. ] | [677.9 680.7 679. 686. 689.5 685.3 680.3 704.8 693.4 692.6 686.9 690.6 688.4 681.7 659.7 698.6 682.3 685.6 684.9 680.7 684.6 682.5 683.7 679.9 683.8 681.3 687.1 688.6 685.1 685.8 681.7 677.9 678.9 676. 691. 681.6 680.1 677.5 678.7 677.4 672.6 676.8 674.5 677.9 694.5 676.5 679.7 685. 693.7 676.1 672.8 692. 681.2 678.7 692.8 680. 692.6 687.5 688.3 679.8 680.8 688.7 680.2 683.7 692.6 679.3 685.4 688.8 691.1 679.2 676.5 682.4 686.1 692.8 680. 681.5 680.8 681.1 678.8 677.8 684.4 685.7 688.4 684.1 681. 676.9 681.8 688. ] | [484.4 483.7 483.1 483.6 483.4 486. 484.5 489.4 485.6 481.7 480.5 482.2 483.1 483.3 487. 489.8 483.8 487.3 490.2 488.1 489.3 488.6 492.9 486.5 480.8 492.1 488.6 482.6 490.6 482.6 486. 485.7 486.8 492.4 486.1 486.5 485.7 488.8 490.2 488.8 488.5 490.3 485. 484.3 492.2 488.5 489.6 488.5 489.4 489.1 493.3 493.9 481.6 484.6 492.9 489. 488. 488.1 494.9 488.8 487.4 490.1 484.6 483.2 491.5 488.2 485.6 484.2 487.1 490.3 490.2 485.1 482.2 481.7 484.4 489.1 482.2 480.4 486. 487.3 482.3 476.4 480.5 484. 481.5 483.3 483.8 482.5] | [313.1 311. 305.9 307.6 313.6 310.8 307.8 306.4 302.3 307.2 311.4 311.2 310.6 310.2 314.1 307.6 306.2 310.8 317.2 315.9 316.6 299.6 310.8 309.1 310.5 307.2 309.9 310.3 313.7 314. 314.1 312.8 307.6 309.1 308.3 313.1 310.3 312.4 313.1 313.8 314.4 312.3 310.2 314.1 312.8 311.3 309.7 308.9 313.3 309.8 310.9 316.8 313.3 310.9 313.6 314.2 311.3 311.4 309.5 312.9 310.6 313.5 314.8 314.9 311.2 310.3 316.1 312.2 313.9 311.2 312.4 310.4 313.7 311.5 313.5 313.4 317.2 310.2 314.7 312. 312.8 311.2 311.5 311.9 312.1 309.9 308.7 309.8 314.5 312.8 316.1] | [243.8 243.6 243.6 243.8 242.3 239.8 240.8 239.5 239. 240.4 240.7 241.1 241.7 246.7 240.9 241.1 243.3 242.6 241.3 244.4 239. 241.1 243.8 240.4 240.1 240.8 241.1 242. 243.9 241.6 240.3 243.2 243.4 242.1 241.9 243.5 242.1 244.6 243.7 243.2 241.2 242.3 243.9 241.1 241.5 243.9 243.2 242.2 240.4 241.3 242.9 242.9 243. 242.9 244.6 246.8 241.7 248.6 242.5 244.3 240.2 240.7 241.7 242.5 243. 247.2 243. 250.1 242.2 242.9 244.1 239.5 244.1 246.4 245.7 243.1 243.5 243.3 239.7 242. 246.3 244. ] | [159.6 161. 161.1 161. 161.4 162.9 296.7 317.2 161.6 165.5 330. 332.2 339.3 164.7 163.5 161.7 161.6 160.8 163.8 163.6 164. 161.3 328.1 160.1 156.6 161.7 161. 161.1 161.8 162.1 163.8 162.6 159.8 156.9 160.8 161.7 160.7 163.2 160.6 160.8 162.7 163.3 161.4 157.9 159.1 163. 161.1 163.5 163.6 162.6 161.8 162.2 163.7 159.2 159.2 160.9 160.4 161. 162.1 162.2 161.4 161.4 162.6 158.7 160.4 331. 161.1 158.9 159.2 161.1 162.1 164.2 349.4 159. 154.8 322.8 159.8 315.3 313.7 161.1 338.9 157. 162.8 161.5 157.1 158.8 161.8 164. ] | [152. 171.5 123.3 121.2 114.2 113.5 116.7 121. 114.5 115.8 117.6 114.4 113.7 115.3 188.2 114.7 121.9 120.3 117.4 115.6 115.6 116.5 114.6 121. 179.2 117.4 124.3 123. 120.2 119.2 113.8 116.1 116.3 118.6 178.4 112.6 123.7 119.5 117.5 118.6 117.3 116.4 116.8 116.5 173.1 118.7 116.2 112. 115.5 116.7 115.6 113.9 112.9 115.7 169.8 119.9 161.2 157.6 116.5 118. 113.5 114.2 116.8 121.8 117.2 170.1 108.9 160.3 117.2 118.9 118.7 115. 118.3 102.4 165.6 175.5 117. 118.1 116.3 119.6 121.1 120.9 102.7 113.1 121.2 164.5 128.3 127.4] | [110.7 1.3 4.2 55.9 44.1 43.4 112.9 5. 78.9 49.8 64.1 65.3 62.5 90.8 5. 85.4 76.2 80.1 63. 70.3 73.4 73. 70.2 101.4 49.2 79.3 65.2 58.8 45.9 74.6 70.8 77.6 70.3 75.6 61.4 58.9 70.5 70.6 70.2 83. 76.6 77.1 72.2 94. 57.3 68.6 75.1 82.6 74.9 84.6 84.2 66.1 62.4 79.4 73.4 78.6 64.8 80. 73.1 90.4 79.6 75.3 69.9 74.3 1.3 86.6 1. 95.7 75.4 80.2 67.3 81.8 68.9 95. 93. 2.1 81.5 80.8 89.1 86.2 77.6 82.5 104.9 3.4 76. 59.3 82. 75.5] | [54.8 2.7 2. 1.8 2.8 3.2 2.7 3.3 1.2 1.2 1.3 1.2 2. 1.2 1. 2.7 1.3 69.5 1.3 2.5 1.3 1.2 1.2 2. 1.3 1.2 1.4 74.9 2.1 3.3 1. 2.5 2. 1.4 3.4 2.7 2.5 1.3 1. 2. 3.3 3.4 1. 3.2 1.8 2. 2.8 1.3 4.6 88.8 1.1 4.9 1.1 1. 3.4 1.1 2.8 2. 2.7 1.1 83.9 1.8 5. 1.1 1.1 9.5 4. 4.9 2.7 2.7 1.2 1. 4. 4.9 1.3 1.3 3.4 1.2 1.2 5. 1.3 5. 4.9 2.4 1.2 1.1 1.3 3. 2. 6.3 2.7 1.2] | [91.5 1.2 45.7 3.1 1.8 1.3 86.6 4. 3.5 3. 1.3 2.7 3.6 3.1 3. 7.2 5. 2.8 1. 1.1 2. 67.2 1.2 1.8 9.5 1.3 2.4 2.5 3.3 9.5 3.4 3.6 2.6 2.7 4.6 1. 1.1 1.2 5. 1.3 8.1 4. 1.2 2.4 2.8 1.3 2. 3. 1.3 2.7 2.7 3.3 2.7 1. 8. 60.6 7.3 4. 9.5 2. 3. 2.5 2.7 1.1 2.7 1.3 95.3 2.4 4.5 3.7 1.2 2.5 1.3 1.1 3.3 3.1 6.3 3.6 4.3 2.7 9.5 1.2 1.9 70.1 1.8 1.3 2.6 2.4] | [ 2. 7. 9.5 5. 1. 55.6 7. 4.3 3.3 8.1 9.5 1.2 141. 66.5 2.9 10. 1.1 9.5 1.2 2.5 10. 1.8 1.7 112.2 3.3 1.1 3. 1.1 2.1 5.8 10. 4.8 78.4 9.5 2.7 2.4 1.7 2. 1. 6.4 9.5 1.9 4.2 63.7 1.8 3.4 1.2 3.2 1.9 5. 3.4 1. 58.8 2.5 45. 2.9 1.1 3. 2. 2.6 3.2 6.7 2.7 50.3 1.3 1.3 2.5 1.8 5.6 9.5 2.5 1.1 1.2 2.6 1.2 4.3 2.4 5.6 9.1 1.6 2.9 2.5 1. 67.6 60.4 9.1 54.3 58.1 2.9 2.7 1.2 2.9 9.5 1.3] | [ 2. 9.1 8.6 2. 3.1 2.7 3.4 1.2 4.8 1.8 9.5 1.8 9. 2.2 8. 3. 2.1 8.9 3.1 3. 6. 2. 2.4 1.4 9.5 2.7 6.3 1.8 5. 3. 4.1 1.2 6.3 1.3 3.8 1.8 3.6 2.7 1. 3. 2. 2.5 5. 5.6 1.8 5.3 1.3 2.7 3.2 1.8 3. 1.9 2. 2. 2.8 10. 3. 1.1 2.4 4.8 6.5 3.6 2. 9.4 1.8 1.9 7.5 1.2 3.1 9.3 2. 3.8 9.3 1.9 6. 1.9 3. 1.9 4. 3. 1.2 1.9 2.8 5. 8. 2.8 2.6 9.1 72. 6.7 7.9 7.9 3. 1. ]
6.014 | CStehningPhilipsClinicalScienceGermany_NIST | https://osf.io/sjy5g/download/ | 20200225_CStehningPhilipsClinicalScienceGermany_NIST/20200225_CStehningPhilipsClinicalScienceGermany_Hamburg_MR1_NIST_Complex_T1map.nii.gz | Complex | Christian Stehning | Philips Research Hamburg | Philips | Ingenia DDAS 3T | 3 | NIST | 130 | 23 | 20.8 |  |  | IR SE | Spin-echo inversion recovery | ['192 192 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2500] | [5 empty] | [1933.9 1817.8 1819.1 1851.1 1808.9 1914.5 1843.2 1856.1 1874.7 1828.6 1866.8 1880.6 1851.4 1889.9 1840.7 1885.9 1845.5 1858.7 1860. 1844.9 1855.8 1822.7 1887. 1876.2 1834.7 1795. 1811.2 1903.9 1834.2 1860.2 1898.5 1812.4 1795.6 1880.4 1869. 1879.2 1854.3 1827.6 1854.4 1819.7 1843.9 1883.3 1852.4 1824.3 1839.4 1840.2 1826.1 1820.4 1852.5 1831. 1856.3 1870.5 1883. 1896.1 1841.8 1884.2 1892.6 1824.5 1856.3 1854.1 1816.3 1863. 1825.5 1863. 1897.9 1847.7 1822.1 1835. 1890.3 1852.7 1822.5 1871.1 1864.4 1792.8 1836. 1875.6 1811.3 1854.3 1861.3 1786.5 1845.1 1861.6 1834. 1858.3 1880.7 1828.4 1820.9 1916.2 1902. 1844.4 1863.3 1862.9 1827.5 1875.2 1831.6 1840.2 1870.3] | [1343.7 1377.8 1388.5 1370.9 1387. 1396. 1362.5 1399.6 1382.4 1359. 1365.6 1351.4 1375.1 1402.8 1359.5 1383.6 1375.1 1363.2 1404.3 1387.1 1377.1 1390.8 1413.4 1396.6 1324.8 1368.3 1346.6 1392.6 1409.2 1367.2 1383.6 1373.2 1376.9 1401.2 1367.7 1342.3 1355. 1395.1 1389.8 1372.9 1383.6 1389.8 1396.7 1375. 1378.2 1401.8 1373.8 1379.9 1413.8 1379.4 1359. 1378.9 1377.5 1388.7 1352. 1389.6 1360.1 1367.3 1403.3 1383.1 1387.3 1373.6 1356.8 1383.6 1372.2 1397.8 1365.4 1358.2 1355.4 1371.7 1379.8 1386.3 1346.2 1350. 1417.4 1363.5 1369.6 1362.8 1376.7 1383.9 1381.8 1346.7 1358.8 1359.4 1378.9 1387. 1417.8 1384.9] | [943.1 958.3 931.4 924.6 948.4 964.7 945.5 952. 926.4 916.4 941.6 963.2 951.5 928.5 929.5 932.9 936.1 941.4 932.6 935. 940.9 935.3 923.2 947.8 921.1 928.7 926.1 929.7 927.6 935.2 953.8 921.1 931.1 938.1 933.6 929.3 939.8 935. 919.7 943.3 930.5 930.3 937.1 944.3 940.5 950.5 928.9 940.2 931.6 938. 926. 937. 942.7 927.7 931.6 947.1 940.4 939.3 951.3 925.6 947.7 949.7 931.2 936.8 938.3 923.4 945.5 940.7 933.9 951.4 936.1 942.9 934.5 937.1 931.2 951.8 951.6 944.9 954.6 928.6 924.1 943. ] | [680.3 676. 672.2 688.9 688. 677.6 674.7 702.8 693.2 697.8 687. 698.1 695.6 684.6 665.3 707.3 675.1 680.7 684.4 678.3 688.2 684.5 688.6 689.7 689. 680.4 689. 695. 685.6 692.9 684.8 677.3 684.7 676.2 696.6 683.4 677.3 672.4 678.7 679.7 672.5 681. 675.8 682. 699.4 669.6 679.2 683.8 701.9 674.4 669.2 700. 684.2 677.4 691.8 675.9 698.3 683.9 693.6 678.3 683.2 691.2 676. 683.9 693.2 676.8 690.4 688. 693.9 679.1 673.2 682.4 686.7 696.1 679.4 682.8 676.6 683.6 678.3 676.1 685.3 687.4 687.1 683.3 679.5 676.9 683.9 689.9] | [487.7 483.7 482.5 481.4 481.5 490.9 498.4 487.9 483.5 472.6 472.7 475. 476.1 479.5 489.7 488.4 480.1 481.2 489.7 483. 487.2 486. 495. 483.2 478.5 492.2 484.5 477.5 487.5 472.7 477.8 479.1 481.4 494.8 482.3 475.2 477.9 482.5 483.7 480.7 481.4 486.1 480.8 480.2 490.6 479.8 482.5 482.8 482.3 479.7 487.2 492. 473.6 483.2 492.1 481.7 484.2 483.2 493.2 482.8 481.7 487.8 481.2 481.7 493.1 484.1 480.9 479.4 483.5 487.8 488.2 482.9 483.2 483.9 482.8 491.5 481.2 475.2 482.6 488.3 480.9 475.8 477.2 483. 479.5 481.4 481.4 485.1] | [309.2 316.5 308.1 308.7 312.4 305.8 301.8 305.9 297.4 305.5 308.8 303.5 311.4 309.2 322.6 305.6 303.5 311.6 317.9 314.4 319.9 301.8 316.3 308.1 308.1 303.1 307.6 306.8 311.3 310.5 310.7 311.7 310.5 309.8 305.2 313.5 307.9 310. 311.5 312.2 311.2 309.5 305.7 315.5 313.5 311.3 307.6 304.7 311.1 306.4 307.4 317.6 311.4 309.4 313.1 311.6 307.7 308. 305. 310.1 307.9 310.9 313.7 314. 306.8 308.1 311.9 306.4 311.3 306.2 310.2 308.7 313.8 308.6 310.4 313. 310.7 301.7 317.2 308.2 312.2 309.1 309.4 309.1 307. 305.9 305.3 306.2 315.6 309.8 313.3] | [245. 241.2 239.8 242.3 242.7 237.4 241.4 238.1 237.8 239.5 239.1 239.5 240.9 253. 237.4 242.2 245.4 244.2 238.3 245.8 234.9 238.9 244.2 238.8 238.4 240.6 240.1 240.6 243. 238.5 237.7 243.4 243.8 241.4 240.9 244.1 240.4 246.1 240.7 246.2 241. 242.1 244.8 239.8 239.9 243.7 239.1 238.5 236. 240.2 242.4 242.1 241.4 241.7 243.5 248.3 235.3 241.6 242.8 244.8 238.3 238.6 239.7 240. 236.8 243.2 239.6 247.3 249.2 240.3 244.5 236.3 242.1 242.1 246.6 237.3 244.3 243. 235.5 238.2 241.3 238. ] | [158.8 159.6 156.9 157.7 157.8 162.8 155.6 161.7 159.7 166.9 162. 165. 167.8 162.8 165.5 164. 158.8 157.8 164.1 162.2 164.3 158.8 163.3 162.7 153.4 163.5 163. 159.7 160.8 161.1 164.4 163.6 160.8 156.4 161.5 161.3 157.9 163.2 158.3 158.6 162.4 163.8 161.4 158.7 158.6 165.8 159.9 165.1 163.3 162.8 160.2 161.1 166.8 159.8 156.8 161.7 159.3 161.7 160.8 161.4 160.4 160.1 162. 158.3 160.6 162.8 163.9 157.8 157.9 159.7 160.9 161.8 164.5 156.7 151.3 160.4 160.3 163.9 159.4 161.8 163.6 151.3 164.6 164.6 156.1 158.4 162.6 162.4] | [114.6 127.5 127. 124.5 114.4 114.5 116.5 122.5 113.7 115.4 117.1 115.6 113. 116.5 133.3 110.7 122.9 117.6 118. 114.4 116.5 118.9 114.9 124.5 128.5 113.6 127.3 121.7 122.3 120.1 110.7 116.1 117.4 120.1 127.1 111.4 126.3 119.6 116.2 119.4 118.1 117.2 120.5 118.7 124.2 118.4 114.9 108.3 117. 115.6 115.8 115.7 111. 115.2 126.2 120.9 120.8 121.1 118. 119.6 112.9 116. 118.1 122.9 117.2 125.3 107.4 120.1 116.8 121.5 121.6 113.2 116.6 103.3 121.3 131.1 117.5 118.3 114.4 120.3 120.6 120.8 106.4 116. 120.5 119.9 127.2 125.3] | [ 99. 2.5 6.5 57.4 2. 50.9 101.6 2.7 78.3 2.5 71.9 71.4 57.7 91.8 1. 87.2 69.7 85.5 57.1 72.8 76.2 75.9 52.2 88.1 65.1 88.3 64. 3.5 3. 69.5 68.4 76. 62.7 72.6 69. 67.4 77.3 71.7 68.4 80.2 68.5 71.7 67.1 89.6 67.1 77.8 79.3 87.9 68. 83.7 78.7 5. 2.7 85.4 78.6 79. 58.3 80.5 62.6 86.3 68.4 74. 65.1 61.9 4.6 81.3 9.5 92.2 66.7 71. 10. 82.2 45.2 89.6 87.3 2. 85.1 80.7 84.5 89.2 63.9 78.2 96.8 2.2 81.3 2.7 81.4 70.1] | [68.9 9.5 9.5 9.5 3.1 9.5 10.3 4.8 2.5 9.5 43.8 3.1 1. 2. 5.3 62.5 47.5 75.2 9.5 1.3 1.2 1. 2.4 3.3 9.5 10. 10. 83.5 2.7 9.5 2.5 1.2 3.1 2.7 1. 1.9 4.1 2. 5. 1. 1.1 8. 3.7 1.3 1.2 2.1 8. 1.2 10. 97.5 1.6 9.5 2.6 1.6 7. 1. 1.3 2.3 62.3 10. 90.9 1.2 1.2 3.4 10.6 2. 10. 1.9 1.3 3.1 10. 8. 6.9 2. 1.2 6.2 1. 6.9 1. 2.6 5.2 1. 2.7 2.4 9.8 2.1 2.8 7.6 9.5 10. 1.3 1.8] | [ 86.2 9.5 75.6 75.6 39.4 7.5 83.1 2.7 3. 2.6 10.9 2. 3.1 2. 7. 10.6 6. 9.5 69.5 1.8 4.9 10.2 7.5 10.2 1.9 1.2 3.2 2.7 1.8 9.6 3.9 10.3 3. 79.6 1.2 81.1 10.9 59.2 3.2 1.1 2.7 1. 1.4 10. 70. 5.8 6.4 3.2 3.2 3.3 3.6 1.1 1.2 4.6 2.7 85.4 4.4 4.1 2. 4.6 3.2 2.7 2.6 10. 9.5 10. 113.2 5. 9.5 1.9 1.2 4.5 9.5 1.3 3.2 10. 7. 7. 1.4 2.5 9.5 9.5 1.3 102.1 3.5 9.5 1.1 2.3] | [ 5.6 3.7 67.4 74.2 5. 2. 6.9 4. 4. 78.6 9.5 3.6 173.3 10.3 4.1 7.6 3.6 10.2 6.9 3.8 3.2 2.4 5. 105.9 2.3 10.3 9.5 3. 9.1 74.4 4. 3.7 107. 2.5 4.9 6.3 9.5 10. 3.2 10. 4. 9.3 7.4 117.2 4. 3.8 5. 5. 9.5 2. 7.6 5.5 101. 5.9 1.3 2.4 2. 1.1 2. 4. 73.3 10. 9.5 88.3 3.3 1.9 103.4 1.3 4.1 5. 4. 65.5 8.5 2. 1.9 1.3 6. 3.2 73.3 40.4 4. 1.2 2. 121.1 72.3 94.7 104.9 117.2 8.8 3.4 107.8 2.8 5.6 7.6] | [116.2 71.4 92.7 1. 10.6 2. 10. 6. 1.6 5. 1. 2.7 9. 4.8 10. 79.1 6. 94.1 3.8 3.1 1. 1.4 8.1 10. 6. 1. 9.4 3.6 2.6 10. 10. 175.4 10. 1.4 10. 9.5 5.2 1.8 10. 5.7 10. 2.9 5.1 2.7 131. 2. 7.9 2. 9.4 1. 4.7 1. 1.6 4.8 9.6 1.9 4.8 5.1 9.2 4.8 6. 2.2 4. 4. 4. 3.3 9. 2. 4. 8. 3.1 1.9 6.8 2.5 9. 2. 10. 1.2 5.1 8.1 1. 4.8 9. 2.4 7.4 3. 5.4 2. 4.9 76.7 1.5 3.2 9.5 2.3]
7.001 | iveslevesque_muhc_glen_NIST | https://osf.io/nax9z/download/ | 20200226_iveslevesque_muhc_glen_NIST/PowerError/20200226_iveslevesque_muhc_glen_PowerError_NIST_Magnitude_T1map.nii.gz | Magnitude | Ives Levesque | McGill University Health Center | Philips | Ingenia R5.3.1 | 3 | NIST | 130 | 42 | 21.2 |  |  | RFE | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2700 | 14 | [50, 400, 1100, 2500] | 364 |  |  |  |  | [2097.5 2003.5 2046.2 2015.5 2003. 2072.3 2032. 1956.6 1984.7 2011.7 1976.7 1937.5 1954.5 1909.9 1989.8 1953.3 1952.1 2026.6 1901.8 1907.5 1948.5 1948.2 1912.8 1832.1 1960.5 2010.4 1910.3 1954.9 1907.5 1897.3 1954.6 1902.4 1841. 1922. 1915.5 1900.5 1945.1 1956.4 1942.4 1905.4 1879.4 1947.3 1928.2 1815.6 1930.2 1985. 1946. 1943. 1904. 1902.6 1916.6 1922.4 1955.3 1924.4 1866.8 1938.6 2014.5 1959. 1867.8 1904.8 1906.2 1940. 2001.7 1949.1 1891.2 1892.4 2007.5 1979.3 1897.3 1873.8 1915.1 1937.1 1961.7 1981.8 1907.2 2050.8 1970.6 1891.7 1884.7 1913.6 1991.1 1924.6 1917.3 1981. 1942.3 1888. 1920.9 2007. 2014.2 1901. 2023.2 2022.5 2013.1] | [1365.4 1352.4 1336.2 1350.8 1359.3 1394.9 1379.5 1320.8 1322.5 1320.7 1333.1 1362.6 1388.4 1351.6 1547. 1340.8 1315.6 1321.2 1343.5 1377.2 1357.8 1360.1 1316.8 1321.1 1541.2 1329.2 1281. 1324.8 1397.9 1411.2 1375.7 1353.9 1326.8 1337.2 1289.1 1322.9 1292.7 1349. 1418.2 1372.9 1324.5 1328. 1311.5 1305.3 1312.4 1359.7 1319.8 1332.6 1357.2 1324.6 1305.6 1320.6 1348.6 1319.4 1378.2 1383.5 1330.4 1321.9 1346.4 1356.9 1347.8 1331.2 1383.9 1383.3 1383.1 1365.9 1364.7 1361.6 1354.6 1357.1 1341.7 1343.4 1349.8 1354.9 1390.8 1378.6 1357.2 1371.2 1353.2 1355.9 1369.3 1355.4 1372.2 1348.1 1378.8 1384.3 1360.4 1364.4] | [958.7 938.1 952.3 958.1 952. 967.8 978.9 954.3 948.1 970.3 969.5 948.4 952.7 955.5 981. 952.9 950.1 962. 972.8 976.3 961.7 934.7 937.9 942.9 953.7 961.7 934.3 961.1 973.3 962.7 976.7 962.3 940.2 944.6 929.5 947.8 950.5 954.9 965.1 966.1 970.1 954.5 930.9 930.2 939.6 954.1 961.6 950.8 964.9 963.7 963.8 957.8 950.8 947.8 956.9 957.2 962.2 966.5 978. 953. 951.2 966.9 960. 964.2 976.5 947.7 951.4 947.4 974.9 960.8 940.9 955.3 947.9 957.5 934.4 955.2 937.7 940.6 947.6 940.8 939.4 937.7 961.6 965.1 943.6 931.2 931.5 911.2] | [704.7 712.8 723.5 695.2 674.9 678.3 679.8 705.4 715.9 707. 685.5 688.9 685.3 706.9 693.6 676.9 670.9 679.3 701. 691.5 699.7 694.3 664.3 699.5 692.1 674.9 681.7 690. 693.2 672.8 675.4 681.9 674.6 678. 683.1 680.8 698.4 703.3 693. 693.7 682.4 685.8 683. 685.4 676.7 684. 682.3 677.2 675.6 681.3 680.1 682.9 685.7 686.5 689.2 690.9 676.2 670.2 670.7 675.7 676.5 686.5 681.2 683.6 687.5 684. 683.3 674. 685.8 699.2 687.1 677.8 669.6 684.9 684. 678.7 679.6 685.2 688.8 689.2 690.6 686.7 678.8] | [493.4 492.2 492.3 487. 486.8 480.1 502.5 500.9 496.2 491.3 495.1 488.9 496.8 480.8 498.6 498.6 495.9 491.7 486.8 501.5 494.7 498.3 484.6 495.9 489.1 492.9 491.2 486. 504.3 500.1 488.2 480.4 498.1 493.9 506.4 501. 490.6 505.9 494.3 484.5 482.2 493. 501.3 507.7 498.5 492.6 497.8 494. 487.4 482.8 488.8 487.7 500.5 493.4 494.1 496.5 489.6 486.5 488.1 496.8 490.6 492.9 485.3 487.4 489. 485.4 485.6 478.5 502. 494.9 483.9 478.7 476.9 478.8 486.3 475.9 485.5 483.9 481.4 482.7 482.3 482.2] | [358.6 353.2 352.9 359.3 352.2 357.4 358.3 354.7 362.5 357.7 355.3 357.4 353. 346.4 347.1 354.2 349.7 349.8 353.4 355.6 356.9 356.4 349.1 345.1 355.1 358.7 348.9 344.5 348.4 357.2 353.6 351.5 350.9 345.8 353.7 359.5 354.7 349.5 350.2 348. 350.6 354.5 353.8 346.6 346.6 355.8 352.3 346.8 352. 356.3 346.9 351.9 351.7 353.7 349.9 344. 350.3 352.1 343. 353.8 355.2 348.9 353.5 346.1 350.7 352.1 350.3 347. 349.8 349.8 359.8 337.8 342.4 346.8 341.5 346. 349.3 348.3 345.4 362.6 349.9 351.9 354.2 357.6 348.5 349.5 365.3 339.3 348. 355. 342. 344.7] | [243.2 240.3 243.1 246.1 248. 245.6 241.5 245.7 247.3 247.3 247.2 246.4 241.9 238.4 243.9 246.2 245.9 245.7 245.7 246. 244.4 239.2 236.6 248.4 244.4 245. 242.3 245.2 245.3 245.9 242.9 245.3 244.2 249.5 239.8 243.5 244.4 245.4 243.9 244.9 245.7 249.7 244.8 245.1 240.2 244.6 241.6 241.3 240.1 245. 243.8 246.2 238.1 238.3 240.1 244.6 241. 242.4 237.7 241.1 241.8 242.8 243.1 247.1 241.2 240.9 242.5 247. 243.8 252.4 252.8 245. 251.2 253.3 239.1 239. 245.1 246.9 247. 249.3 247.4 233.4 241.5 241.7 238.5 232.5 241.3 239.2] | [176.3 176. 176. 173.9 171.2 178.7 176.2 177.4 176.8 178.4 177. 176.3 176.8 181.3 176.2 177.4 381.7 178.9 176.3 174.7 179.5 175.7 179.3 372.3 184.7 186. 379.1 181.5 379.5 181.5 178.2 174.7 182.9 394.7 180.6 181.7 380.9 382.1 388.5 181. 175.4 173.9 181. 388.8 387.6 374.5 182.1 389.4 368.4 173.4 386.1 173.2 179.4 395. 385.9 177. 176.8 382.1 358.5 172.5 180.7 364.1 180.7 177.7 399.5 383.6 367. 176.3 177.7 174.7 421.9 173.9 408.7 169.4 359.4 168.4 178.2 180.6 173.2 172.8 175.3 189.6 171.4 177.6 180.3 180.3 386.6 434.4] | [137.7 116.7 123.1 119.4 159.5 183.1 151.2 163.3 161.8 126.1 161.9 110.8 103.8 121.3 137.9 133.2 130. 124.2 125.3 124.2 137.1 107.7 171.3 155.4 162. 159.5 127.4 129.2 119.5 110.8 110.9 180.6 158.5 161.4 165.1 170.6 175.4 123.7 147.2 107.5 123.3 159.1 165.4 157. 158.2 170.3 161.6 153.5 140.5 169.5 132.6 169.9 157.4 154.4 160.8 152. 153.3 117.3 164. 137.7 161.5 154.9 149.8 146. 152.6 154. 118.3 143.7 163.4 120.3 133.4 153.3 152.7 154.4 117.6 120.3 122. 117.6 117.8 116.5 144.5 115.7] | [ 89.6 97.9 86.9 91.1 88.6 110.5 96.1 7.2 91.3 82.6 91.2 89.1 65.5 117.1 93.4 93.3 60.9 90.6 92. 101.3 98.5 89.8 93.8 80.7 99.3 93.3 66.2 101.6 112.2 98.2 87.1 86.7 90.3 90. 88.2 63.2 86.9 103.4 101.7 92.1 92.7 82.7 70. 84.4 84.9 78.7 77.2 97.9 100.4 93.5 84.9 79.9 79.1 79.4 92.7 91.6 87.1 79.2 88.3 77.9 70.1 82. 81.7 85.2 90.7 82.1 69.3 74.9 87.4 67.2 1.2 89.7 76. 100. 109.6 79.9 86.8 6.3 90.5 86.4 84.6 88.1 91.8] | [ 55. 2.5 1.9 9.5 72.4 72.3 2.1 3.2 1.3 6.3 16.3 73.6 58.8 1.3 3. 1.2 2.6 68.5 5.3 9.5 4.3 9.5 3.3 70.8 5. 1.1 3. 73.6 1.2 1.2 1.3 14.1 58.2 68.2 51.8 1.3 6.3 2.5 5. 68.1 3.1 3.1 1.8 3.5 1.2 5. 1.3 3.2 1.3 52.1 86. 6.3 9.5 65.1 2.5 2.5 1.3 1.1 3.3 2.6 64.3 94.6 3.1 9.5 73. 85. 3.3 3. 3. 9.5 66.9 49.7 2. 4.4 93.6 6.1 9.5 9.5 59. 68.9 1.2 85.6 62.3 93.9 9.5 6.4 58. 81. 3. 107. 98.7 112.4 114.2 117.5] | [ 1.8 2. 2.4 2.9 2.5 354.9 10. 1.8 1.2 4.5 2.8 1.2 1.9 2.5 100.9 102.8 3.6 2.8 1. 10. 2.8 1.1 100.8 102.2 80.2 71.3 3. 74. 96.4 59.7 1.1 10.2 109.9 77.4 1. 84.5 3.3 3.4 81.9 4.2 76.3 82.4 1.3 94.3 86.8 87.1 70.9 5. 1.2 4. 4.6 2.4 1.8 64.7 87.6 81.8 91.3 63.2 75.4 1.8 2. 1.2 1.1 65. 67.1 77.9 77.7 79.3 10. 3.6 1.8 2.4 77.6 77.9 2.7 2. 67.2 1.8 78.7 83.9 3.6 2. 56.9] | [ 5. 1.2 2. 55.5 2. 1.8 72.9 20.2 73. 81.7 4.2 100.5 75.9 1.1 80.8 72.7 85. 86.6 9.6 9.5 59.5 74.4 4.9 62.5 100. 97.2 69.4 73.3 85.2 94.2 13.2 74.6 84.9 65.1 88. 94.9 80.4 59.4 92.8 90.3 82.2 2.5 1. 81.6 73.1 75. 82.2 79.2 67.3 83.3 84.1 59.5 76.3 76.2 78.1 75.2 95.4 96.6 69.5 93.3 87.6 76.5 86.8 50. 1.3 59.6 83.2 82. 76.6 145.1 3.4 77.8 84.3 6.6 1.2 1.2 75.3 83.9 75.7 7.7] | [170.8 241.2 234.1 72.1 63.4 97.6 105.2 224.9 208.7 3. 3. 2.1 82. 94.3 74.3 56.4 1. 86.7 2. 5.3 4.1 80.7 82.1 87.4 97. 82.7 77.6 80.6 1.3 56.5 60.3 63.5 86.1 77.8 79.9 2.1 1.2 3.2 1.3 56.7 85.1 78. 9.5 1.1 3.3 78.6 186. 72.5 82.2 3.3 3.2 79.5 2.1 76.7 76.6 103.3 66. 144.5 97.8 85.6 144.5 82.3 87.8 87.7 3. 4. 82.1 95. 243.7 66.2 92.5 692.7 191.1 98.9 181. 1. 71.4]
7.002 | iveslevesque_muhc_glen_NIST | https://osf.io/nax9z/download/ | 20200226_iveslevesque_muhc_glen_NIST/NoPowerError/20200226_iveslevesque_muhc_glen_NoPowerError_NIST_Magnitude_T1map.nii.gz | Magnitude | Ives Levesque | McGill University Health Centre - Glen Site - Cedars | Philips | MR 53.1 | 3 | NIST | 130 | 42 | 21.65 |  |  | WIP IR_SE_TI_ISMRM | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 397 |  |  |  |  | [2004.3 2072.1 2095.2 2093.3 2037.7 2041.4 2085.6 2077.7 2005.8 1996.7 2040.5 2089.7 2044.4 2023.3 2039.3 2046.4 2016.9 2135.9 2036.1 2044.5 2032.9 2035.6 2030.5 2044.5 1974.9 2067.3 2080.8 2060.7 2064. 2072.4 2034.4 2042.6 2039.4 1989.2 2022.4 1976. 2126.1 2071.7 2043.2 2050.1 2029.7 2026.2 2063.8 2057.6 2033.9 2072.5 1969. 2152.4 2094.4 2042.2 2025.2 1993. 2018.6 2049.8 2043.8 2035.6 2063.2 1982.5 2100.1 2086.3 2013.8 1991.8 1996.3 2013.5 2030.1 2036.9 2043.8 2067.7 2010.2 2045.4 1976.4 1976.1 1982.6 2007.7 2006.4 2029.3 2059.6 2076. 2019.4 2080.5 2009.3 2033.2 2055.2 2030.9 2050.9 2075.3 2063.7 2007.3 2050.5 2031.7 2056.8 2082.8 2074.4 2034.4] | [1446.1 1430.8 1438.9 1454.1 1469.3 1421.6 1391.8 1391.7 1396.1 1409.2 1439.5 1417.9 1399.8 1418.5 1404.1 1414.5 1413.7 1420. 1421.8 1429.7 1433.8 1410.5 1401. 1392.6 1418.3 1420.8 1424.1 1429. 1442.6 1430. 1416.6 1449.9 1422.9 1402.1 1406.2 1418.9 1424.2 1444.3 1464.1 1452.4 1411.2 1420.8 1415.9 1431.2 1410.3 1411.6 1421.8 1424.9 1437.3 1432. 1415.5 1427.9 1413.8 1410.7 1394.4 1400. 1418.5 1411. 1427.2 1439.5 1426.7 1435.6 1431.8 1387.4 1400.9 1411.6 1408.4 1427.3 1445.1 1428.6 1432.4 1434.3 1404.6 1417.8 1423.2 1420.8 1434.3 1435.5 1442.3 1423.5 1424.8 1441.2 1425.8] | [ 975. 972.5 986.4 969.3 968.7 980.9 988. 983. 988. 994.8 988.5 993.2 989.5 1009.4 998.6 981.4 966.9 966.8 965.4 966.9 971. 962.4 967.5 998.6 980.4 966.1 978.8 976.5 966.5 973.1 976.7 986. 960.7 980.7 972.8 967.3 973.5 974.2 963.5 974.5 978.1 987.1 986.3 978.3 974.9 984.3 982.2 984. 966.9 974.2 985.7 986.3 987.4 974.9 978.7 996.2 972.8 979.1 973. 974. 981.4 988. 986.8 980.1 987.9 982. 971.4 981. 978.9 986.2 990. 991.2 990.6 981.6 986.5 980.9 986.7 975.3 989.4 993.8 994.1 989.2 994.7 983.2 983.5 990.8 1001.2 1012.5] | [702.1 717.6 718.9 712.4 709.9 721.2 717.1 717.7 695.3 695. 703. 701.5 709.6 707.7 700.8 685.1 712.1 697.7 712.6 712.4 714.2 715.1 709. 700.1 713.3 695.9 714.7 702.5 704.5 705. 711.7 708.3 710.9 704.3 709.6 706.7 717.5 696.7 707.3 710.7 715.7 706.6 707.3 705.8 707.9 706.2 716.3 704.6 712. 707.7 712. 699.6 700.8 706. 710.2 702.3 718.9 703.9 711.3 705.5 704.3 701.8 712. 709.2 708.4 696. 720. 726.9 707.4 710.7 706.5 711.9 713.1 712.8 697.2 711.1 707.6 707.2 706.5 703.9 712.7 706.7 703. 701.7 703.7] | [499. 502.6 508. 510.1 504. 504. 495.7 509.4 496.3 501.7 497.2 497.5 496.1 498.1 497.9 495. 487.9 504.1 500. 497. 497. 498.7 502.4 504.3 497.7 496.2 500.1 505.5 498.1 498.8 497.6 501.1 499.3 500.9 499.7 501.1 500. 502.7 497.6 498.5 500.3 501.9 496.8 502.5 500.4 503.4 499.2 500.5 501.1 498.2 497. 497.3 493.1 498.7 504. 499.6 499.1 501.2 499.8 497.2 498.1 499.9 495.3 496. 503.2 501.3 501.5 500. 497.9 499.4 498.5 499.2 492.5 494.3 489.7 501.3 511.1 497.1 503.5 502.1 498.1 491.8 487.8 506. 501.4 499.5 503.4 500.9] | [342. 344.7 353.3 351.8 355.9 355.3 355.7 354.4 353.3 352.9 355.1 356.2 360.5 354.8 351. 350.8 351.8 352.6 353.8 354.6 346.6 353. 350.3 354.7 353.1 356. 354.5 354.2 353.5 356.9 342. 354.8 350. 351.4 356. 356.9 353.5 352.8 348.5 351. 351.8 353.7 357.1 349.6 351.4 355.7 354.3 352. 351.6 349.4 350.5 350.7 355. 355.3 353.1 351.6 349.5 351.9 352.7 352.7 349.1 349.8 349.8 352.3 354.3 353.8 353.2 352. 353.4 353.4 355.2 351.8 350.5 350. 356.5 354.4 351.5 352.5 351.5 349.7 351.1 349.2 351.5 354.2 353.3 354.2 353.4 353. 350.6 355. 352.3 352.1 357. 354. 351.5 353.4] | [253.1 254.4 254.3 254.9 256.8 255.1 250.7 250. 247.7 248.7 249.5 247.4 246.9 248.7 245.4 246.1 249.6 248.5 250.1 251.5 248.6 249.9 248.9 249.3 248.1 248.5 248.2 247.2 248.5 249.5 249.6 250. 250.5 248.9 248.7 248.1 248.4 250.7 248. 248.8 252.1 249.6 249.2 245.7 247.8 248.5 248.7 250.1 247.6 247.7 250.1 249.4 251.5 245. 247.5 248.5 249.1 249.8 248.6 246.4 247.4 248. 250.3 244.7 247.8 248.9 247. 246.2 246.4 244.2 246.2 246.3 246.7 244.8 249.5 247. 244.3 245.2 247. 247.9 246.2 246.1 250.2 244.8 246. 247.6 246.2 246. ] | [178.8 178.7 179.4 179.4 178. 181.8 180.2 183.2 181.7 181.4 181.7 180.8 422.3 180.2 182.1 183. 179.5 179.4 181.3 408.2 180.9 404.5 405.2 407.5 178.7 182.1 180.4 181.9 184.9 415.6 414.3 181.1 401.1 411.6 179.8 180.9 181.5 181.8 182.4 405.4 404.3 400. 423.4 434.8 178.2 178.3 179.6 180. 179. 398.8 411.8 402.1 407.6 173.3 178.3 393.4 176.9 177.6 392.1 391.1 396.8 178.7 417.2 172.4 176.9 423.4 399.2 418.2 412.5 395.2 421.7 177.1 171.4 378.3 176.4 178.8 178.2 176.7 179.7 436.9 181.5 414.4 181.6] | [161.6 125.9 126.9 150. 159.6 165.1 161. 192.9 172.3 168.4 165.8 163.6 169.8 158.3 165.1 183.7 166.5 165.8 165.8 166.8 165.5 172.8 172.9 166.8 168.4 166.4 161.3 169.3 167.3 167.6 168. 172.5 169.3 159.1 167.1 166.4 165.3 167.4 165.4 168.9 167.2 167.2 165.3 157.5 166.3 167.3 162.9 164.7 168.1 166.5 168.6 167.4 158.2 155.2 168.1 163.6 158.9 166.2 169.2 164.5 165.1 166.3 156.4 167.5 161.6 176.6 166.7 163. 164.1 166.2 164.3 167.3 159.9 170.4 154.8 171. 159.6 157.8 161.1 161.1 158.1 149.1 170.1 166. 160.2 160.6 170. 169.8] | [107.6 105.9 99.7 98.5 106.7 98.1 95.6 101.8 103.3 113.5 105.4 100.8 99.7 100.9 96.7 112.7 90.2 101.6 105.9 100.2 99.7 96.3 100. 97.7 104.1 90.3 107.2 97.2 99.9 96.7 99.7 101.6 105.6 94.8 99. 94. 109.6 96.5 95.7 97.8 104.7 103.7 103.6 89.8 94.4 94.2 109. 100.6 97.5 101.2 104.8 101.4 101.2 92.7 97. 88.8 110.7 105.9 97. 89.2 94.2 97.9 98.3 93.3 103.4 99.6 118.3 98.7 95.6 95.4 97.1 89.8 86.2 106.5 80.9 119.3 90.5 76.4 82.7 78.1 77.2 4. 123.8 107.3 107.7] | [ 2.7 74.2 69. 4. 2.4 78.7 110.6 99.6 102.8 98.5 101.8 109.3 111.8 2.7 74.8 71. 81.2 83.2 90.7 67.2 89.9 3.1 100.2 88.7 83.7 91. 89.1 92.1 106. 2.9 91.9 88. 93.2 93.8 86.6 89. 84.5 79.2 96.1 92.4 90.2 62. 91.8 89.1 86.3 88.8 91.6 85.8 82.2 57.3 65.9 77.9 73. 90.1 80.3 66.1 74. 89.1 85. 76.6 73.4 4.5 95.4 84.4 96.6 83.6 76.9 76.8 90.9 84.1 72.8 76.3 62.5 81.2 90.4 85.7 77.5 73.2 74.8 75.2 78. 74.8 92.5 92.8 84.3 63.9 69.8 73.8 80.8 79.4 77.3 79.6] | [ 93.3 1.2 60. 93.9 132.9 61.2 99.9 71.5 106.5 106.7 123.8 133.2 104. 102.9 113.7 94.3 153.2 89.9 101.9 99.8 105.3 120.3 100.5 135.9 1.2 81.7 81.4 102.3 101. 125.1 75.7 145.5 3.2 96.8 95.8 99.2 98.9 118.6 70.9 138.6 2.4 1.2 109.3 99.3 89. 101.3 96.3 113.2 9.5 1.2 82. 73. 85.7 94.2 83.4 2. 59.2 37.2 84.8 87.3 99.8 98.7 93.8 79.6 55.6 81.8 77.1 87.7] | [ 7.6 9.5 148.3 135.8 117.7 104.1 128.4 145.9 46.7 10. 1.3 1.2 3.6 2.1 6. 1.8 117.2 68. 71.9 86.2 101.5 89.6 112.1 2. 125.9 80.5 1.2 2. 59. 65.1 93.9 7.9 7.3 76.7 53.8 126.4 1.3 92. 91.3 90.9 65.7 45.6 79.5 44.4 3. 4.4 129.2 94.8 106.5 47.7 88.1 70.8 1.3 79.3 59.5 7.4 2.8 41.5 80.4 82.2 81.8 70.9 39.8 39.2 38.8 59.2 2. 35. 2. 55. 75.7 4.7 5. 3.2 58.5 74.7 104.5 107.6 39.1 71.3 58.9 1.1 59.2 105.7 3.3 76. 5. 2. 2.7 4. 2.7 1. ] | [230. 246.7 162.9 546.8 57. 74. 92.3 1. 2.5 64.9 152.1 174.8 130.6 125.7 137.8 139.2 150.6 132.7 128.7 1.1 157.3 123.5 7. 1.8 7. 3. 73.9 1.2 5. 4.6 157.8 9. 1.8 1.1 8. 70.7 54.9 9.1 1.3 5.6 2.5 104.1 3.1 6.4 3.2 1.2 41. 57.4 42.1 74. 44.6 8. 86.3 1.2 88.8 1.2 3. 1.1 5.3 43.5 113.6 2.3 2. 173.1 1.8 1.3 5. 3.1 6.4 1.2 3.2 3.7 145.4 1.3 98. 6. 1.1 105.1 111.1 83. 106.9 1. 192. 1.2 6.4 2.8 102.3 6.2 6. 154.3 2.5 121. 90. 60.2]
8.001 | karakuzu_polymtl_NIST | https://osf.io/bfh4m/download/ | 20200227_karakuzu_polymtl_NIST/20191213_UNF/20191213_karakuzu_UNF_FS_NIST_Magnitude_T1map.nii.gz | Magnitude | Agah Karakuzu | Unite de neuroimagerie fonctionelle | Siemens | Prisma fit | 3 | NIST | 130 | 42 | 22.1 |  |  | t2_tse_dark-fluid_cor | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2400] | 287 |  |  |  |  | [1881.6 1877.1 1842.6 1889.2 1890.3 1845.7 1902. 1839.5 1875.2 1880.9 1880.3 1864. 1890.1 1852.6 1870.8 1859.5 1819.2 1879.7 1895.3 1901.8 1871.2 1860.2 1834.3 1890.7 1927.9 1880.9 1920.6 1890.9 1915. 1912.2 1882.8 1890.4 1846.8 1900. 1893.6 1938.9 1866.3 1888. 1874.5 1888.3 1870.3 1854.8 1883.9 1849.4 1854. 1918.6 1916. 1914.3 1913.2 1904.9 1871.2 1853.4 1878.2 1920. 1866.8 1849.8 1905.2 1893.1 1921.5 1961.5 1921.7 1843. 1862.2 1908.9 1893.6 1846.5 1900.4 1884.1 1952.9 1886.6 1934.8 1869.8 1865.2 1859.4 1853.6 1909.5 1914.8 1893.9 1885.8 1913.3 1900.3 1844.3 1904. 1933.5 1879.3 1844.1 1856.3 1860.5 1889.2] | [1390.1 1361.7 1384.5 1386.5 1380.3 1388.8 1340.7 1415.3 1349.7 1387.8 1376. 1359.4 1366.6 1388.1 1402.9 1381.7 1390.6 1363.2 1384.8 1362.4 1372.2 1370.1 1362.7 1363.4 1366.2 1402.5 1374.2 1371.3 1355.8 1361.3 1352.1 1359.2 1360.4 1395.9 1374.5 1389.1 1364.4 1396.1 1371.3 1368.8 1359.9 1414.7 1353.4 1383.5 1356.7 1385.5 1363.8 1400.4 1381.7 1390.3 1363.5 1385.1 1355.7 1373.9 1362.3 1377.3 1367.8 1402.8 1395.8 1398.2 1401.9 1380. 1366.4 1367. 1370.9 1370.7 1367.8 1394.5 1363.7 1376.7 1363.1 1357.1 1384. 1347. 1360.3 1396.8 1376.7 1364.5 1379.4] | [ 974.4 968.2 956.8 952.4 953.8 960.1 983.9 955.2 953.2 965.3 956. 967.1 966.4 952.1 985.5 964.9 962.6 995.3 958.7 951.6 957.2 962.2 979.2 947.8 1000.6 953.2 993.8 970.6 949.1 976.5 972.7 959.7 993.7 970. 988.8 965.3 977.1 951.8 964.5 953.5 949.6 965.8 973.4 971.5 985.9 972.4 977. 977.9 967.1 978.8 961.5 958.6 959.6 955.8 989.4 978. 978.3 985.2 968.8 966.5 960.1 957.2 962.5 973.8 973.6 992.5 984.1 966.8 976.6 967.4 957.6 966.2 977.6 988.7 992.3 963.7 968.8 958.4 976.2 985.5] | [695.6 672.3 669.1 662.5 673.2 680.8 674.2 680.7 704.6 711.9 709.5 709.5 700. 677. 673.6 706.1 707.4 694.4 676.8 675.4 677.9 697.4 712.1 695.9 705.6 699.3 683.8 699.9 699.6 710.2 702.6 699.8 685.9 707.3 712.9 689.7 711.8 704.6 698.7 699.4 694.4 709.9 710.4 688.7 701. 707.3 709.9 699. 711.3 715.9 711.2 695.1 716.6 699. 694.6 712.8 701.2 714.1 698.7 702.3 703.6 715.4 702.5 709.7 696.4 723.5 708. 715.9 702.1 726.7 699.1 717.2 700.6 717. 712.6 703. 710.1 695.3 700.4 703.2 714. 707.6 715.3 706.9 717.2 717.8 715.3 700.5] | [511.7 504. 499.1 500.2 499.6 512.6 510.9 505.6 507.8 512. 517.4 512.6 503.9 507.1 513.7 500.3 511.4 507.3 501.2 502.7 507.9 512.2 503.8 503.4 506.3 504.5 512.7 500.5 510.2 510.1 511.7 503.3 511.5 508.4 503.9 513.4 507.1 508.9 502.2 499.8 504.7 505.2 506.2 517. 504. 518.2 507.3 510.2 505.6 520.9 504.2 515. 507.1 519. 505.9 518.4 506.3 516.3 510.1 509.6 507. 512.7 503.1 517. 509.4 511.3 513.4 510.4 513.2 513. 515.3 504.5 504.5 512.7 507.8 520. 514.2 507.9 505.9 503.9 508.3 518.1 508.2 514.9 518.1 515.1 515.1 514.8] | [343.8 345.1 344.6 346.3 350. 345.3 347.8 354.9 341.8 353.3 349.4 347.8 349.2 349.7 345.8 341.7 348.3 348.1 353.3 352.9 349.4 354.1 350.8 355.5 348.1 351.7 352. 356.3 357.5 352.1 353.3 356.5 354.4 351.2 350.4 348. 345.8 346.3 348.6 358.5 349.6 355.7 356.7 354.9 354.9 349.2 351.9 351.6 356.4 350.3 356. 354. 360.5 355.2 356.9 355. 353.8 352.4 354.5 345.4 348.6 356.5 355.5 362.3 353.8 353.7 357.8 355.2 351.2 353.4 354.1 353.5 359.1 354.6 359.8 350.4 350.4 349.2 349.2 349. 358.4 352.4 356.9 354.1 351.1 353. 351.4 352.5 358.3 345.3 348.2 351.9] | [259.7 258.6 262.8 260.8 263.8 260.1 262.9 258.9 261.1 258.9 256.5 255.7 258.8 263.8 264.9 258.7 261. 256.2 263.1 260.9 262.2 258.4 257.4 262.7 261.9 257.5 259.9 259.5 257.3 253.3 254.4 258.5 252.9 254.4 260.4 258.6 255.5 258.1 253.8 264.4 259. 257. 257.6 259.3 260.5 258. 255.4 254.9 254.9 252.5 255. 258.2 251.4 260.1 260.2 256.1 260.2 255.2 256.4 260.1 259.1 255.1 257.9 257.9 257. 252.6 258.9 252.5 251.7 255.6 256.3 258. 254.8 254.7 253.3 256.1 257.3 258.9 253.4 253.3] | [193. 188.6 187.6 188.4 187.7 190.1 189.2 191. 189.1 189.1 191.5 189.8 191.8 189.5 191.9 186.8 192.3 185.1 191.4 184.9 193.4 185.4 192.9 189. 194.6 186.5 193.2 184.1 192.7 184.3 188.9 186.5 194.2 187.7 195. 189. 189.7 189.3 189.3 191.4 189.8 191.7 189.8 189.1 189.2 195.6 186.7 613.9 188.4 185.9 189.6 186.1 191.8 615. 188.2 650. 193.5 185.3 188. 191.3 192.3 189.1 186.6 184.7 195.4 612. 193.3 187.3 185.2 185.5 184. 187.7 195.5 539.7 611.4 188.3 187.5 189.8 603.7] | [218.8 205.4 198.5 193.7 199.6 217.3 217. 228.1 205.8 218.3 220.2 208.5 206.7 218.4 214. 219.1 200.6 223.7 197.4 192.3 205.4 209. 204. 203.2 214. 237.2 208.7 224.5 203.6 233. 192.2 218.1 186.9 206.2 210.8 235.1 193.8 225.1 195.1 205.2 209.9 214.4 197. 205.2 211.5 224.2 202.3 210.3 223.1 228.1 222.9 196.1 212.5 187.3 238. 194.5 210.4 203.9 193. 192.9 198.1 211.1 194.8 199. 252.4 212.7 188.6 200.9 206.2 207.8 204.6 183.8 191.9 233.7 228.3 224.8 193.6 189.5 190.8 134.6 211. 230.2 201.8 220.4 218.5 214.7 215.3 200.1] | [136.3 146.8 141.6 139.1 148.5 145.2 128.6 146.7 135.2 138.6 145.1 131.7 145.5 144.6 139.4 144.3 135.7 147.3 136.3 132.4 147.1 125.6 154.6 128.8 139.2 139.7 138.2 134.2 135.7 133.6 138.2 130.4 145.1 136.1 143.8 132.5 134.8 132.4 138.3 139.9 139.3 128.5 139.4 128.7 133. 144.8 124.5 139.3 131.2 127.3 141.3 122. 140.6 129.3 132.4 137.5 138.7 132.3 131. 128.8 121.9 131.4 134.7 130.8 147.9 124.2 121.2 133.9 127.4 124.9 114.2 122.6 116.1 152.6 150.8 129.6 121.7 122.4 118.6 115. 129.8 146.2 139.8 140.1 133.6 130.2 135.5 136.9] | [117.7 102.3 102.3 114.1 106. 144.4 97.8 126.9 108.4 104.4 106.3 104.9 102.2 134.7 131. 107.4 112.7 105.2 116.2 115. 120.9 107.7 117.3 139.4 123.8 115.2 108.6 121.1 106.5 114.9 105.3 118.4 111. 142.2 121.9 114.6 112.7 123. 110.9 128.8 110.8 120.5 113.9 124. 139.8 126.5 118.7 121.3 121.9 119.5 108.5 125.8 125.4 126.7 123.3 133.2 135.7 102.6 126.8 117. 128.5 123.3 124.8 133.9 130.2 153.4 117.4 109.5 111.6 124.8 121.2 128.8 122.9 109.6 133.5 134.2 136.3 114.7 114.4 124.7 120.9 120.1 132.4 139.3 125.2 122.2 129.7] | [1860.9 167.8 177.6 181.2 2513. 191.7 15. 170. 157.2 153.4 202.4 9.5 145.8 142.1 136.7 162.7 71.4 257.2 179. 134.9 130.5 161.7 131.6 133.3 137.7 159.4 153. 156.2 120.9 117.7 137.3 135.8 163.1 148. 126.3 138.7 126.1 134.9 159. 153.1 125.6 115.8 123.5 136.2 156.2 129. 131.2 128.3 140.5 171. 106. 131.5 131.8 123.4 133.1 124.5 155.2 110.4 139.6 156.9 134.1 124.2 144.9 105.8 130.5 147. 142.2 125.9 125.7 151.7 125.7 142.4 134.6] | [201.2 185.9 190.3 208.7 208.6 186.1 153. 146.8 157.9 151.6 173.6 207.3 204.4 170.1 145.4 160.7 175.9 171. 175.2 162.3 146.4 208.4 156. 177.4 169.5 152.5 157.6 157.6 185. 174.2 168.1 208.2 168.1 181. 154. 166.9 171.2 180.1 167.2 169.7 196.1 165.2 200.8 174.4 152.4 154.9 177.1 139.3 166.8 141.5 192.4 173. 190.1 197.3 227.3 162.6 168.9 156.1 168.4 138.7 181.5 148.8 181.2 157.2 157.4 156.9 198. 147.2 174.7 162.2 154.3 173.2 163.7 188.9 169.8 174. 197.8 167.3 154.2 172.6 167.8 183.9 180.9 174.6 170. 176.1 178.1 168.4 168.9 172.2] | [223.1 210.9 217.8 288.5 237.2 224.6 263.2 263.6 220.3 292.6 408.2 219.3 215.5 226.7 220. 188.4 238.6 215.6 250.1 299.7 186.7 255.7 162.4 354.6 211.6 258.4 214.2 255.7 253.9 257.1 422.7 266.4 275.1 210. 250.8 192.9 256.5 213.2 258.3 256.6 302.6 325. 347.6 259.9 292.1 276.8 286.2 253.5 301.4 264.9 263.8 454.3 284.5 338.6 371.7 311.4 253.1 224.8 263.1 299.9 207. 305.8 350.1 290.4 245.1 290.9 347.8 338.9 286.6 219.5 260.9 369.4 470.2 343.8 270.9 242.7 218.6 251.2 347.1 381.5 314.1 436.2 422.7 430.2 413.2 307.5]
8.002 | karakuzu_polymtl_NIST | https://osf.io/bfh4m/download/ | 20200227_karakuzu_polymtl_NIST/20200227_MNI/20200227_karakuzu_mni_FS_NIST_Magnitude_T1map.nii.gz | Magnitude | Agah Karakuzu | Montreal Neurological Institute | Siemens | Prisma fit | 3 | NIST | 130 | 42 | 22 |  |  | t2_tse_dark-fluid_cor | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2400] | 285 |  |  |  |  | [1854.5 1894.4 1910.8 1923.8 1926.8 1848.9 1913. 1873.2 1906.2 1920.8 1827. 1963.4 1827.8 1937.1 1891.1 1864.1 1905.2 1895.9 1866. 1911.1 1883.5 1896.9 1878.8 1951.9 1928.2 1906.7 1870.5 1915.5 1923.4 1883.5 1889.5 1946.6 1902.9 1910.5 1928.3 1924.5 1904.8 1876.6 1897.3 1923.7 1893.3 1904. 1896.8 1916.2 1914.2 1910.9 1895.7 1918. 1911. 1896.5 1947.4 1948.9 1930.3 1917.7 1940.5 1929.4 1936.7 1927. 1948.4 1908.6 1890.9 1892.5 1909.6 1943.8 1947.5 1949.2 1917.3 1997.3 1925. 1932.4 1895.4 1926.1 1913.3 1905. 1913.5 1940.6 1900.2 1962.7 1950.7 1895.8 1896.1 1889.4 1855.8 1883. 1925.8 1900.6 1974.1 1928.6 1881.9 1914.2 1970. 1968.1 1918.1 1946.7 1943.3 1938.7 1940.6] | [1293.4 1417.6 1400.4 1328.7 1342.2 1368.3 1311.8 1405.1 1375.6 1316.9 1391.7 1408.5 1406.2 1411.9 1432.3 1349.7 1411.5 1336.9 1418.9 1374.1 1392.8 1392.6 1376.5 1387.2 1317.4 1387.3 1354.4 1386. 1382.8 1360.5 1401.5 1386.1 1390.4 1390.7 1302.3 1421.2 1360.9 1403.7 1394.5 1415. 1381.8 1365.6 1384.1 1407. 1284.4 1417. 1374.9 1410.6 1376.8 1385.6 1381.1 1416.9 1378.5 1398.5 1266.8 1399.9 1364. 1396.2 1376.2 1381.8 1380.7 1372.8 1370.7 1404.2 1375.2 1382.5 1388.1 1411. 1384.8 1405.6 1410.7 1389.8 1383.4 1415.3 1334. 1387.3 1386.1 1399.5 1399.1 1366.3 1402.1 1336.4 1377.9 1389.5 1395.4] | [823.3 859.9 841.8 826.1 823.8 830.4 835.4 875.1 818.7 818.9 835.2 831.6 836.3 825.3 824.9 904.1 830.7 839.4 847.3 844.4 848.5 858.1 856.1 855.1 881.9 866.6 843.7 856.1 840.7 858.3 875.2 869. 846.4 870.2 934.2 848.5 882.1 873.2 876.2 866.3 859.3 874.6 889.9 862.8 946.3 847.1 913.5 877.8 893.3 863.5 907.1 855.9 898.4 874. 954.8 854.9 941.3 887.1 918.2 866. 884.2 878.6 891.5 879. 961.3 859.8 919.1 887.2 901. 911.5 901.7 910.8 868.6 900.8 903.2 870.6 931.5 877.2 887.1 889.9 877.5 895.1 865.5 874.5 919.2 917.8 913.5 905.1] | [684.8 468. 453. 449.4 447.4 447.8 592.2 447.2 440.2 593. 439.2 439.7 596.6 595.7 596.5 443.6 607.3 446.8 441.9 440.9 441.4 441.3 440.9 611.7 446.3 621.4 446.7 440.9 439.1 599.1 600.2 437.7 438.8 439.7 620.1 595.9 443.2 603.6 592.3 589.8 591.2 587.4 599.7 601.9 641.9 446.6 600.6 610.1 611.3 608.5 611.3 610.2 598.6 607.7 643.3 590.1 627.1 619.9 621. 608.7 609.2 617.4 620.3 626.8 651.3 605. 637. 617.9 626.7 624.3 632.1 624.3 623.3 636.1 610.6 643.1 621. 647. 622.5 626.6 640.1 622.9 636.2 638.4 629.7 645.5 643.1 622.7] | [444.4 427.3 418.8 420.2 421. 431.3 448. 422.1 426.7 437.8 447.8 442.3 429.8 423.5 464.5 427.9 436.3 437.2 433.6 436.9 432. 441.8 442.8 424.9 449.4 431.7 441.4 434. 450.3 451.4 448.1 435.6 444.1 439.6 435.7 449.9 442.9 455.1 453.3 437.6 452. 457.5 438.6 455.1 440.1 474.6 448.6 463.7 447.6 473.3 445.8 458.6 445.8 462. 445. 475.6 451.6 468.2 447.5 460. 443.8 464.6 451.6 467.2 449.5 466.2 460.7 465.5 469.5 454.5 463.2 465. 449.8 464.8 449.6 470.9 451.7 458.2 458.2 455. 452.3 462. 455.6 469.9 464.7 460.4 458.4 467.8] | [324.3 313.8 307.1 303.5 302.2 335.9 315. 317.1 319.4 325.4 320.3 310. 345.2 316.2 322.7 323.2 316.2 317.3 315.3 316.4 314.4 342.5 332. 322.5 326.1 322. 326.9 327. 323.1 314. 317.6 310.8 358. 321.1 335.1 324.3 334.6 324.3 324.7 324.8 325.1 317.3 322.4 363.1 320.2 343.2 325.3 335.5 325. 341.2 321.3 329.4 320.8 332.8 359.6 323.2 337.7 331.8 338.5 325.9 325. 326.2 332.4 322.4 326.4 342. 332.6 328.2 335.4 330.3 334.1 328.9 332.5 328.4 328.2 319.5 344.7 326.8 330. 328.6 326.2 323.5 323.5 325.1 323.4 338.4 324.9 327.6 329.6 331.4 329.8 320.6 331.3 321.5 316.3 315.5 315.3] | [233.8 218.7 216.5 216.1 212.9 211.3 239.6 222.6 233.3 227.6 232.6 229.7 229.7 221. 256.6 224.5 238.5 232.7 226.7 228.7 225. 223.7 230.7 223.3 249.9 227.5 241.9 233.2 243.8 235.5 236.3 234.8 226.2 231.5 244.6 239.3 243.6 240.1 238.2 233.4 232.1 236.7 231.2 239.5 243.7 245.5 246.2 241.1 239.1 244.8 237.8 235.4 233.5 241.4 249.2 231.6 249.4 236.8 243.9 230.4 234.1 240.7 230. 238.2 256.9 227.4 243.6 235.8 239.8 240.9 238.5 231.6 233.3 230.8 243.2 229.9 238.3 233.8 230.4 228.2 229.7 231.7 237.8 227.3 234.2 233.8 237.1 229. ] | [165.3 137.6 146. 156.6 153.8 151. 179.8 137.3 155.1 154.7 149.9 151. 148.1 157. 177.5 162.2 147.3 163.6 156. 164. 163.1 158. 154.8 157. 188.3 148.8 161.2 154.4 165. 158.2 156.4 165.3 158.7 168.2 189.1 148.8 170.4 157.2 168.5 162.9 166.9 164.6 162. 170.7 186.4 157.1 165.6 162.2 164.7 161. 159.4 165.1 160.2 168. 175.5 171.6 149.8 169.6 160.5 164.9 164.5 161.2 164.8 163.5 185.5 155.4 155.8 164.5 160.4 158.4 155.3 165. 151. 178.5 157.2 152.8 161.5 157.2 159.1 153.7 170.3 161.6 151.9 150. 157.4] | [129.1 106.3 103. 107.9 105.6 97.8 144.9 93.3 118.4 108.9 112.3 112.4 116.9 112.1 246.4 118.5 113.1 128.1 117.2 123.5 120.6 111.9 122.3 113.4 247.2 103.5 131.7 114.6 132. 121.6 122. 123.2 114. 132. 244.5 103.4 224.9 119.2 134.2 120.9 125.5 124.3 116.3 132.5 259. 104.6 207. 121.4 132.1 119.5 122.6 122.4 115.7 127.2 263. 119.3 121.9 127.2 120. 123.6 122.7 115.6 122.3 117.2 223.3 229.8 105.7 128.8 114.7 112.4 113.2 112.6 119.6 103.4 240.3 128.7 109.3 110.3 103.7 105.3 105.6 107.3 218.3 132.8 112.2 103.6 104.7 113. ] | [157.8 136.4 142.3 153.1 142.4 146.2 165.4 128.7 144.4 142.6 134.1 136.9 136.7 151.6 165. 165.8 150.9 147.3 147. 158.6 158.8 154.8 139.9 150.9 160.2 152.3 146.1 135.1 144.4 134.2 126. 150.4 134.5 147.9 168.5 137.9 139.4 143.2 138.8 147.8 148.7 141.5 137.7 142.6 168.1 141.9 141.4 127.5 138.9 127.6 125.7 137.6 127.1 138.2 161. 141. 144.5 134.8 136.9 140.9 143.6 132.2 128.8 139.2 153.4 150.9 125.2 133.2 138.5 124.5 128.7 123. 130.7 123.7 150.5 150.6 120.1 129.7 142.4 140.4 119.4 120.2 145.8 151. 131.1 121.3 121.3 128.5] | [132.8 112.1 103.8 124.5 129.1 98.1 106.6 114.9 113.5 108. 107.6 141.7 110.6 117. 119.3 98.6 107.7 104.6 110.2 144.6 127.2 136.4 108.5 105. 107. 118.1 113.1 114.3 114.1 136.5 140.8 126.2 122.6 112.3 116. 111.4 107.5 108.8 118.3 109. 138.3 140.1 127. 133.8 116.6 110.7 107.1 112.6 116.6 122.7 116.7 153.8 131.7 135.8 109.9 128. 102.7 119.5 114.1 124.7 118.8 125.7 119.6 144.7 108.4 125.2 121.1 119.2 119. 121.5 119.6 114.7 145.2 132. 114.4 115.7 121.6 117.6 126.9 121.6 121. 133.3 140.1 124.7 118.3 118.7 120. 110.1 126.1 137.2 135.7 130.4] | [172.5 151.5 168.1 136.3 175.4 10. 159.6 150.4 179.1 176.2 13. 138.2 127.9 131.7 122.3 252.7 181.9 9.5 130.3 148.6 122.5 177.9 119. 167.3 150.7 170.8 149.4 13. 126.4 130.4 143.4 131.8 173. 138.5 168.8 150.4 146.4 143.3 143. 139.4 127.5 115.6 135.5 183.4 144.8 139.6 127.3 145.8 159.4 105.3 143.4 123.4 110.1 151.7 124.8 142.9 127. 143.1 149. 149.4 128.1 129.4 107.5 140.2 154. 137.2 129.4 136.8 156.2 172.3 140.8 145.1] | [172.3 176.2 162.6 156.4 173.2 195.7 177.9 149.9 159.8 171.4 153.4 154.5 192.4 207.2 163. 152.9 174.3 173. 167.6 168.2 162.5 143.4 188.5 166.6 188.6 164.5 168.7 183.4 162.4 172.5 179.9 165.4 234.2 174.8 188.9 156.3 170.1 153.9 155.4 175.3 164.7 182.6 162.8 235.7 175.7 197.4 150.4 169.4 143.1 183.1 148.2 180.7 176.3 178.1 161.6 233.4 152.6 168.9 141.2 160.6 147.9 180.5 169.4 184.5 164.8 195.9 182.2 164.9 172. 177.1 187.5 179.4 207.7 165.3 207.9 185.5 177.6 181.1 164.9 192.1 185.3 193. 190.3] | [207.7 230.4 248. 185.9 230.2 246. 187.8 177.1 172. 239. 215.8 225.3 218.2 212.6 226.3 274.6 219.8 224.9 225.4 255.5 217.9 236. 233.3 203.2 206.1 225.6 197.9 201.4 235.5 402.9 289.6 242.6 247.3 249. 248.6 236. 244.7 207.3 212.7 272.6 453.5 288.3 290.1 244.3 273.2 206.2 192.8 240.2 245.4 255.5 247.8 506. 305.8 257.3 255.4 245.7 250.9 268.7 244.6 284.1 263.1 265.9 421.1 247. 260.4 252.9 226.4 216.2 240.7 278.7 207.7 320.3 469.8 432.9 293.6 265.8 290. 319.9 263.5 249. 284.4 423.1 382.1 299.2 277.6 280.3 362. ]
9.001 | madelinecarr_lha_NIST | https://osf.io/8nugw/download/ | 20200227_madelinecarr_lha_NIST/20200227_madelinecarr_lha_NIST_Magnitude_T1map.nii.gz | Magnitude | Madeline Carr | Liverpool Hospital Australia | Siemens | Magnetom Skyra | 3 | NIST | 130 | 111 | 19.72 |  |  | SE_IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 130 |  |  |  |  | [1925.8 1910.3 1978.1 1892.9 1918.4 1933. 1867.5 1947.6 1918.6 1843.7 1862.8 1886.2 1893.8 1875.3 1869.5 1884.4 1879.1 1889.1 1895.3 1878.7 1894.8 1876. 1903.7 1985. 1828.9 1837.4 1875.6 1891.1 1865.9 1892.4 1906.5 1886.4 1889.3 1907.7 1879.2 1864.3 1886.4 1907.2 1926.4 1873.4 1914.6 1867.3 1893. 1911.4 1911.2 1899.9 1863.6 1894.5 1860.6 1872.1 1851.3 1882.3 1879.2 1879.8 1912.1 1928.8 1940.4 1879.8 1902.2 1890.3 1916.9 1928.4 1908. 1909.4 1864.4 1837.5 1951.2 1882.3 1919. 1880.3 1850.2 1841.3 1844.3 1914.9 1980. 1830.7 1934.7 1910.6 1815.7 1899.5 1938.4 1904.2 1889.7 1883.5 1836.6 1951.1 2004.3 1895.6 1823.1 1805.2 1839.2 1936.5 1794.5 1948.8 2025.7 2026.8 2002.7] | [1323.8 1346.6 1361.2 1340.7 1381.6 1371.5 1341.2 1351.6 1353.3 1349.7 1328.5 1348.2 1329. 1343.4 1359.6 1333.9 1376.2 1351.5 1378.5 1371.8 1378. 1357. 1348.8 1358.5 1350.6 1348.2 1360.8 1346.3 1338.6 1325.2 1344.1 1326.1 1352.1 1356.9 1345.9 1353.4 1358.2 1345.8 1365.5 1363.7 1372.2 1351.7 1363.8 1348. 1365. 1357.9 1355.8 1362.9 1325.6 1354.1 1338.1 1335.1 1333.5 1330.7 1330. 1364.6 1370.9 1330.7 1367. 1333.9 1376.9 1372.2 1368.7 1358.8 1356. 1340.4 1353.4 1346.3 1366.7 1335.5 1369.1 1338.2 1326.5 1341.1 1341.6 1336. 1365.4 1341.7 1276.2 1371.5 1348.8 1329.1 1400.1 1378. 1374.2 1361.2 1389.5 1318.4 1287.5 1377.8 1366.8 1300.8 1322.2 1320.3 1319.6 1317.8 1362.1] | [ 964.1 936.6 943.1 963.3 948.7 963. 958.8 957. 967.3 966.6 971.8 976. 972. 982.3 929.8 976.9 960.5 977. 993.8 971.6 990.7 987. 988. 978.5 945.8 986.4 985.4 984.8 976.8 989.2 998.1 982.3 997.2 979.9 954.7 1004.1 986.3 993.5 984.5 1001.1 1000.5 992.1 990. 993.6 944. 989.5 982.5 1004. 997.5 1004.1 988.8 1001.1 993.7 989.9 988. 982.3 998.8 992.1 983.6 996.2 997.7 1000.5 990.5 1020.2 1008.4 953.8 983.2 1004.7 992.5 995.9 991.8 1005.1 1017.6 984.3 1010.4 973.7 965.4 995.3 1000.9 1005.5 1012.2 960.6 1009. 1015.7 959.2 973.9 973.8 961.9] | [452. 623.2 621.1 616.9 626.7 614.5 452.7 630.1 626.5 614.4 619.4 610.6 612.4 624.1 628.3 602.3 631. 632.3 633.4 631.5 633.3 636.5 629.7 628.3 616.3 629.8 627.7 634.3 630.3 635. 623.6 630.3 628. 642.8 612.3 633.6 638.8 642. 636.9 646.2 642.5 639.1 641.5 642.8 629. 637.3 637.2 644. 640.9 636.6 630.4 647.1 635.7 645. 666.5 612.4 656.4 644.8 640.2 659.1 652.1 649.4 647.6 655.4 691.9 636. 631.5 668.5 659.2 650.1 642.8 655.5 661.9 617.7 678.5 641. 618.8 647.5 662.9 665.4 640.8 617.6 682.1 683.6 653.1 639.6 640.6 652.5] | [480.2 488. 490.9 496.2 487. 489.6 477.7 493.1 481.4 485.2 482.8 484.7 485.6 482. 478.4 490.3 483.2 492.3 483.1 484.9 489.3 489.5 490.2 485.7 480. 489.9 485.2 484.4 481.8 483.8 482.6 486.1 487.3 484.6 481.9 491.3 481.2 486.4 484.7 484.6 487.1 486.5 482.8 483.8 481.7 488.5 485.9 487. 483.8 478.1 482.1 483.5 479.7 484.2 481.7 484.4 479.8 491. 481.4 486.1 487.4 484.8 484.9 483.9 487.2 484.6 486.7 486.9 485.5 484.5 487.4 479.3 484.6 485.2 486.5 488.2 485.9 490.8 484.2 480.9 483.5 486.1 487.5 489.7 482.9 484.3 489.1 489.4 481.9 485.2] | [339.1 339.2 341.2 340.2 340.5 340.7 349.2 353.3 357.3 352.9 354. 349.8 339.7 351.8 342.7 342.4 336.2 339.6 334.9 341.6 340.1 347.2 340.8 351.2 340.8 354.8 341.2 354.1 345.4 353.3 342.9 348.2 346. 342.7 352.2 336.8 343.4 334.5 341.7 334.2 343.7 339.1 350.6 347.8 342.8 351.1 342.7 351.9 342.8 353.8 344.2 350.8 342.2 347. 345.5 341.7 350.4 337.4 343. 336.2 341. 335.4 346.3 339.9 346.5 343.2 345.2 341.1 348.1 345.7 358. 342.4 352.4 343.3 351. 342.8 346. 347. 342.8 340.4 351.5 340. 339.9 337.4 341. 335.3 341.6 339.8 351.4 341.2 346.8 353. 349.9 351.3 349.3 350.3 349.2 350.4 337.8 338.9 342.7 342.8 343.1 342.6 344.3 346.4] | [247.4 246. 243.7 249.1 246.8 245. 247.3 240.8 243.7 238.8 241.1 239.1 244.6 236.7 238.2 248.8 240. 248.5 241.5 248.8 241. 247.2 241.3 243.7 239.7 249.9 236.2 242.9 235.8 241. 238.6 241. 237.8 244.6 240. 246.7 242.1 248.6 240.8 247.4 242.2 245.9 242.3 246.2 240.3 245.8 238.5 243.9 238.4 242. 236.6 242.8 237.9 243.4 240.5 245.6 241.9 246.5 242. 250.2 244.3 246.8 242.5 243.6 240.4 244.6 241. 237.2 239.2 242. 238.1 239.5 238.4 244.8 243.8 243.9 243.9 243.9 245.6 243.9 240.6 244.2 241.2 242.2 240.8 239.6 241. 243. ] | [175. 170.8 167.4 163.7 163.4 165.4 175.1 172.3 164.6 178.9 179.2 178.6 175.8 180.3 492.1 171.7 165.5 173. 172.4 166.7 173.6 166.3 170.3 175. 176.2 155.1 177.1 170.1 182.9 176.2 180.1 178.8 180. 173.7 175.7 158.8 174.9 167.1 173. 165. 168.4 164. 173.7 164.7 171.2 168.2 160.5 172.8 172.9 173.4 172.1 180. 170.1 178.2 168.7 175.1 167.5 156.7 168.9 161.7 164.5 161. 161.7 162.2 166.1 160.2 166.6 173.4 148. 161.4 166. 169.9 167.8 168.9 168.9 165.6 165.3 151.6 174.5 147.7 159.9 156.6 155.2 150.5 154.4 152.8 158.1 160.5 167.5 137.3 154.2 157.1 163.6 163.9 155.9 153.3 142.8] | [114.7 117.2 120.6 117. 112.1 116.5 117.4 125.5 167.1 163. 169.6 167.3 170.2 171.5 165.7 124.8 163.1 116.6 119.1 118.6 118.3 116.6 114.7 164.7 119.2 123.8 111.5 125.2 170.3 171.2 122.6 177.1 123.2 154.2 118.6 120.8 110.5 120. 115.2 121.4 114.6 117. 117.1 120.1 117.1 169.8 116. 175.5 119.3 173.5 121.8 173.3 120.6 172.1 117.4 171.3 113.5 122.6 120.7 119.3 116.2 120.7 118.3 162.4 115.3 124.1 150.5 124.1 120.7 176.4 168.7 125.7 115.8 164.7 126. 118.7 161.8 117.3 117.7 116.9 114.1 156.1 113. 125.5 113.8 164.4 163.2 117.6 160.9 109.4] | [ 97.2 89.2 93.4 104.1 92.8 98.2 95.8 77.8 77.1 74.4 83.5 62.2 86.5 64.3 84.3 84.4 78.7 92.7 83.9 103.3 79. 100.5 78.5 84.8 86.5 79.6 78.7 90. 65.6 92.9 74.8 89. 74.3 92. 86.2 95.2 76.9 103.7 85.3 99.6 79.9 101.8 69.9 105.8 88.8 86.3 76.4 86.7 79.6 76.1 71.5 99.9 66.8 94.5 100.5 87.5 98.3 73.4 83.7 90.2 77.7 91.1 81.5 79.5 103.1 86.3 84.9 84. 68. 80.5 74.1 71.7 79.8 89.6 98.7 93.6 96.3 84.9 81. 80.4 89.9 99.5 98.5 98.5 90.6 92.3 92.3 93.3] | [62.8 4. 9.5 1.3 6.3 78.6 3. 69.8 2.6 89.4 78.7 80.9 56.9 70.4 2.7 3.3 2. 7.9 1.2 1.2 3.4 48.2 63.2 52.1 56.2 1.2 88.1 1.6 86.8 52.3 63.9 1.2 85.8 2. 75.3 1.2 4. 3.3 1.4 3.2 60.7 1.6 64.1 92.3 1. 79.6 5. 91.2 1.8 97.5 4.2 84.7 5.1 77.8 93.8 2. 1.1 2. 72.1 8. 44. 2. 70.5 3.3 67.5 90.4 1.2 49. 1.1 88. 57. 89.5 8. 82.6 3.2 92.1 3.4 62.7 2.7 2.4 8. 9.5 1.5 56.7 2.4 67.7 9.5 6.7 1. 61.9 73. 57.3 2. 63.4 70.3 2. 3. 2. 82.5] | [ 1.2 1. 2. 2.5 66.3 48.3 74.8 98. 46.4 60.9 1.2 2.8 1. 1.3 3.1 1.3 7. 1.2 3.1 84.7 3.3 3.6 3.1 108.4 1. 83.8 2.5 98.2 1.2 100.9 3.2 86.2 3.3 6.1 2.5 63.3 1.2 1.2 2.7 75.7 103.9 1.1 83.3 3.4 100.7 4.1 94.2 3.6 81.5 4.2 86.9 86.4 1. 2.1 1.3 1.8 2.8 6.5 2.1 1.4 1. 75. 2.7 63.6 2.1 51.6 88.8 2. 93.7 46.8 86.6 6.4 78. 54. 59. 3.3 1. 2.5 2.5 2.5 2.5 1.2 1.2 2.5 1.2 75.8 1. 6.4 1.8 89.2 59.4 91.5 1.3 62.4 64.8 6.4 2.5 3. 4.8 3.2 2.7 74.1] | [ 1.6 2.4 1.2 64.2 7.2 87.2 2.8 113.9 2.4 107.9 3.1 2.5 98.8 1.9 73.6 1.1 3.1 3.4 60.3 3.1 47.1 3.6 1.2 2.6 90.7 6.7 105.8 1.1 91.5 2. 74.4 116.2 1.1 110.4 1.8 58.6 1.6 2. 2.7 2.7 1.8 61.9 119.6 4. 82.9 4. 94. 1.9 95.5 62.8 103.4 45.3 78.5 1.2 1. 3.3 6.4 2.8 3.1 56.9 3.6 62. 4.9 2. 3.3 3.8 45.4 63.7 2.5 102.2 70.5 103.3 1.8 9.5 92.4 1.2 3. 4. 1.3 9.5 1. 2. 3.1 2.4 51.6 2.5 2.7 7.3 2.4 1.2 10. 54.5 1.2 8.1 3.6 3.3 4. 1.2 3.2] | [ 2.6 2.6 1.2 7.3 5. 52.3 1.1 78.1 9.5 144.8 1.2 13. 3.4 4. 95.5 2. 1.2 4.5 1.8 2. 90.2 8. 2.7 3. 51.1 101.4 9.1 68.6 1.8 8.6 2.8 1.5 66.3 4.2 4.9 1.1 1.9 1.1 2. 55. 58.1 2.5 112.3 133.9 1.8 132.2 2.7 79.3 2. 117.8 4.1 67.1 2. 66.7 141.5 8.5 49.6 3.6 89. 5.6 105.8 1.3 63.4 2.1 4. 8.9 4.2 124.8 2.7 81.6 3.3 135. 1.9 112.4 2.6 7. 2.1 9. 2. 83.9 1.7 1.3 2. 99. 2.5 53. 6.4 2.8 1.6 119.9 1.1 57.6 57.7 1.2]
10.001 | wang_MDAnderson_NIST | https://osf.io/kszy2/download/ | 20200302_wang_MDAnderson_NIST/day1/20200302_wang_MDAnderson_day1_NIST_Magnitude_T1map.nii.gz | Magnitude | Chenyang Wang | MD Anderson | GE | DV26.0_R03_1831.b | 3 | NIST | 130 |  | 19.57 |  |  | IR Spin Echo TI=50,400,1100,2500 | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 122.109 |  |  |  |  | [1813.8 1698.3 1719. 1734.1 1684.5 1728.3 1689.5 1862.9 1944.3 1923.4 1928.4 1815.4 1664.8 1513.9 1579. 1669.2 1802.6 1903.8 1772.9 1746.9 1717.8 1744.2 1923.6 1985.6 1811.3 1625.8 1512.9 1539.2 1751.5 1712.9 1710.9 1719.9 1692.6 1777.7 1923.4 1990.1 1936.3 1703.2 1883.5 1987.8 1869.3 1883.4 1596.7 1552.4 1694. 1587. 1807.9 1746.6 1760.4 1905.3 1789.8 1638.1 1598.1 1846.3 1751.3 1552.7 1802.4 1702.9 1706.7 1776.1 1732.8 1679.5 1661.6 1812.6 1766.7 1751.3 1756.4 1622.6 1822. 1743.3 1950.4 1758.5 1745.8 1655.1 1686.6 1748.3 1578.8 1708.2 1578.2 1843.6 1688.2 1561.4 1955.9 1762.4 1637.9 1581.2] | [1294.8 1305. 1225.8 1260.4 1359.9 1358.8 1249.7 1274.5 1317.6 1235.9 1242.5 1262.4 1307.2 1294.1 1253.7 1251. 1246.3 1239.4 1289.4 1326. 1263.3 1309.7 1314.3 1300.2 1270.1 1309.2 1308. 1297. 1324.6 1318.8 1308.3 1294. 1265.6 1310.5 1289. 1297.6 1253.2 1299. 1378.2 1273.8 1275.9 1253.9 1233.5 1342.8 1288.9 1346.1 1342.9 1341. 1428.9 1284. 1272.9 1285.3 1260. 1339.6 1231.8 1260. 1296.9 1262.4 1305.8 1266.7 1267.5 1296.1 1336.1 1327.8 1303.7 1259.8 1249.2 1223.6 1278.7 1314.5 1290.3 1294.3 1315.6 1263.7 1292.4 1270.2 1270.3 1323.7 1347.6 1328.6 1275.8 1241. 1243.5 1259.6 1301.3 1269.9 1307.4 1277.1] | [942.9 918.5 922.3 931.8 889.3 911.3 933.8 905.1 896.2 922.4 931.6 905.4 938.4 935.9 872.6 919. 886.8 909.6 940. 936.7 897.9 920. 944.8 945.4 880.8 899.6 881.4 904.6 891.2 890. 884.1 910. 943.2 914.9 918.8 865.2 857.7 900.6 902.9 932.2 938.5 936.2 932.8 911.6 890.6 885.5 934.1 922.7 910.7 918.7 921.1 925.3 930.9 946.2 914. 923.4 945.7 905.6 908.7 906.2 918.7 905.4 901.6 917.7 932.2 890.4 915.3 921.6 923.6 911.1 922.8 926.8 935.5 932.8 881.1 932.8 928.7 925.8 916.1 915.6 906.2 935.9 940.3 867.5 900.1 941. 946.6 895.1] | [659.9 670.1 663.3 665.5 658.4 672.9 677.2 715.4 727.8 659.4 676.6 657.3 661.8 669.1 669.5 685.6 690.4 667. 651. 678.5 664. 654.3 689.2 667.5 657.9 669.5 672.4 671.2 674.7 654.1 665.1 674. 668.9 657.5 659.9 659.7 654.3 665.5 659.4 652. 669.7 662.9 629.4 647.8 674.5 666.1 659.3 672.1 667.3 662.5 668. 654.8 652.1 671.4 726.8 662. 665. 671.8 680. 688.3 663.4 645. 667.4 680.3 701.6 649.2 672. 663.7 665.8 678.1 660.2 653.8 662.5 651.1 661.2 651.6 641.7 655.2 669.4 666.9 659.3 665.6 676.3 662.4 649. 656.8 661.4 649.8] | [483.5 477.7 479.2 474.9 468.2 469.3 485.6 479.6 484.4 486.9 477.5 478.5 477.8 479.4 474.7 485.9 468.6 481.3 480.7 466.3 474.9 476.6 475.8 469.3 479.6 479.9 470.7 471.9 483. 484. 486.1 485.2 485.2 467.8 473.5 475. 483.9 470.9 473.8 477.5 473.1 474. 486. 483.3 469.6 471.1 476.1 469.4 474.3 475.9 476. 475.6 483.1 480.4 470.8 476.7 475.3 470.3 474.5 471.5 482.5 488.1 476. 462.9 473.4 482.2 486.7 478.4 472.6 457.4 465.7 487.4 474.9 475.6 470.3 476.4 469.3 476.2 473.8 472.9 487.6 474.3 472. 475. 474.6 472.4 472.6 471.9] | [353.7 338.3 333.4 339.1 354.2 335.8 342.3 329.9 331.7 345.3 354.8 361. 351.7 338.8 324.5 350.1 356.2 346.9 337.6 324.4 324.6 340.7 339.6 350.6 357.5 342.6 341.7 336.6 349.9 330.1 332.5 338.5 339.9 340.6 330.1 337.6 357.1 349. 347.6 340.1 345.4 337.3 339.6 338.8 341.8 345.3 342.8 333.5 349.7 346.3 357.7 345.5 338.7 335.7 336.4 337.8 341.1 342.1 340.6 339.6 346.1 343.6 336.6 345.1 344.2 339.2 344.4 346.5 350.4 345.8 342.8 327.2 327.4 332.3 338.8 346.5 346.2 333.9 323.1 333.4 340.5 339.5 341.1 333.9 339.8 337.9] | [238.1 231.8 243.3 242.4 231.1 236. 240.2 246. 237.7 237.8 244.4 243.2 242.4 242.2 243.1 237.8 245.7 242.1 232.5 236. 236.8 239.2 232.6 232.9 239.9 240.6 241.2 242.9 243. 240.9 236. 240. 236.5 242.2 245.9 249.9 245.6 243.4 240.9 241.3 242.6 243.7 241.7 245. 242.7 240. 241. 249. 243.4 240.5 233.6 233.5 233.3 240.9 245.1 246.4 242.2 237.6 238.4 244.2 236.2 240. 238.9 237.9 245.5 245.6 243.8 234.5 236.5 238.6 237.4 238.7 232.7 233.3 234.6 239.5 239.2 239. 236.8 244.1 245. 238.8 235.7 235.1 244. 248.9 247.9 241.2] | [178.6 380.3 364.9 172.9 179. 341.7 169. 175.2 320.7 171.8 159.1 302.6 329.1 168.7 366.9 160.4 328.8 371.6 173.9 169.2 351.7 319.1 165.3 168.6 305.2 166.9 182.1 318.4 279.9 157.3 178.2 350.7 306.2 163. 170.6 304. 320.2 168.3 315. 286. 169.5 173.2 282.9 321.1 182.1 173.5 319. 174.3 173. 298.3 352.6 191.7 177.3 299.4 310.6 171.1 182.4 296.6 173.2 177.7 319.3 314.8 289. 279.7 287.4 298.2 291.5 271.9 289.7 179.1 291.9 278.8 309.2 308.5 179.7 294.1 169.9 181. 315.7 171.3 168.7 174.9 180.7 291.1 186.6 300.4 165.9 168.8] | [183.1 188.1 162.9 117.9 111.1 101.4 85.7 125.3 195.7 183.7 123.8 112.5 112.6 121.1 175. 126.5 106.8 127. 173.3 169.1 137.3 111.9 107.2 109.3 141.9 140.4 121.5 107.2 101. 137.8 155. 172.6 138.3 123.1 103.1 143.7 174.1 156.5 120.3 130.3 112.4 143.7 125.4 159.9 135.6 179.4 159.4 150.8 123.5 114.9 123.8 129. 88.9 126.8 112.1 146.9 133.4 114.5 128.7 127.5 145.9 167.5 132.6 162.4 133.9 152.5 132. 129.5 132.6 120.8 133.9 116.8 111.5 131.9 127.7 120.8 141.1 163.2 130.6 130.2 115.1 121.1 143.7 136. 143.6 143. 129. 134.7] | [ 56. 123.8 104.4 97.4 97.5 86.7 84.9 77.6 88.3 88.2 80.4 77.6 81.4 85. 86.8 94.4 100.5 93.8 83.2 89.6 76.1 63.4 91.2 101.4 113.7 86.5 89.2 109.1 72.5 96.1 99.9 68. 92.2 95.4 79.4 90.2 1.2 4.9 9.5 86.8 90.6 66.9 97.7 102.1 70.5 80.9 62.4 108.2 114.7 109.4 85. 9.5 79.8 106.9 92.5 108.7 77.4 105.4 87.6 99.5 108.5 91.4 55.5 93.1 74. 98.6 89.3 120.6 98. 88.9 62. 9.5 2.7 80. 74. 4.5 92.3 83.6 73.1 90.8 99.3 102.2 96.8 108.4 91.5 93.7 99.4 79.2] | [ 8.1 2. 1.2 5. 82.8 111.1 118.8 95.4 2.4 92. 76.3 1. 59.3 100.3 2.7 2.1 4.5 1.1 85.4 102.9 63.1 2.6 78.1 89.7 82.8 1.6 58.2 97.7 2. 2.1 9.5 99.8 2.7 2.8 86.4 60. 9.5 84.8 88.4 1.2 91.9 98.3 118.4 100.4 90. 112.4 88.9 1.3 91.1 92. 71. 1.6 60.6 1.8 92.4 47.4 84.5 114.5 1.2 125.2 96.5 75.6 3.4 4.1 2. 81.9 1.1 88.4 2.1 1.3 73.2 120.8 88. 1.1 1.2 115.2 86.2 2.8 72. 1.1 1.1 2.7 86.2 106.2 7. 2. 1.2 3.3 2. 64.9 1.9 6.1 5.1 3.1] | [ 74.2 97.7 83.2 1.2 2.5 103.9 164.9 6.3 3.3 2.8 3. 2.5 7. 2.6 89.8 79.7 93. 2.4 3. 72.8 3.6 107.2 117.3 86.9 1.8 8. 96.1 80.4 2.4 66.4 70. 81.9 1.2 117.4 1.1 2.7 2. 81.1 5. 1. 2.4 86.2 75. 3.3 84.4 116.6 96.9 1.9 139.6 70.4 99.4 2.6 9.5 60.6 1.3 135.5 98.8 70.8 1.8 118.1 6.1 1.3 1.2 60.9 79.9 132.8 89. 82.7 92.7 1.8 1.6 117.5 2.4 4.6 1.2 6.1 3.3 1.9 2.8 4.7 4.4 2.5 2.6 101.3 62.8 71.7 3.6 3.6 89. 2.8 2.6 2.4 8. 4.1] | [ 4.1 1.4 3. 2. 70.3 4.2 68.7 45.9 42.1 2. 1.5 1.8 72.2 1.1 3.2 7.9 122.4 7.7 1.3 3.6 78.3 84.7 77.5 3. 5. 71.4 73.6 3.6 77.4 4.1 1.8 3.1 77.7 99.7 6.3 2. 81.2 2.7 2. 92.8 75.2 52.8 7.6 1. 4.4 5. 3.7 96.8 3.2 2. 82.7 62.9 3.2 1.2 6.6 1.4 3.3 127.9 87.5 1.2 2. 1.2 2.8 5.3 109.2 1.3 78.4 76.9 93.4 59.2 83. 2.5 3.1 3.1 7.6 2.5 70.4 91.1 1.1 9.5 85.4 2.3 2.2 3.4 3.6 1.8 8.7 102.7 140.8 3.2 1.3 2. 1.8 1.2] | [112.7 71. 2.8 1.4 6. 228.5 77.6 2.7 1.1 2.5 2.7 2.8 4.9 59.5 195.8 5.6 154.7 873.4 2.7 1.1 723.6 283.1 155.1 1.9 7. 80.5 490.1 123. 2.3 340.1 237.1 1.6 152. 238. 1.8 4.7 5. 6.6 61.6 2.3 4. 182.7 2.1 5. 80.2 1.6 468.4 244.8 6.5 87.8 78.2 9. 119.6 94.2 130.6 236.6 62.1 531.8 1.3 2.7 6.3 134. 96.6 1.5 9. 1.4 3.3 94.6 70.9 124.4 391.5 193. 78.6 64.3 254.3 87.2 3.2 3. 74.7 96.3 111. 145.9 53.9 767.1 4.4 3.4 1.2 1.4 4. 1.8 5.8 3.1 1. 1.8]
10.002 | wang_MDAnderson_NIST | https://osf.io/kszy2/download/ | 20200302_wang_MDAnderson_NIST/day1/20200302_wang_MDAnderson_day1_NIST_Complex_T1map.nii.gz | Complex | Chenyang Wang | MD Anderson | GE | DV26.0_R03_1831.b | 3 | NIST | 130 |  | 19.57 |  |  | IR Spin Echo TI=50,400,1100,2500 | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 122.109 |  |  |  |  | [1822.1 1712.3 1732.5 1753.7 1695.7 1742.4 1707.4 1889.2 1959.7 1930.6 1944.5 1836.7 1675.7 1522.6 1591.6 1675.3 1811.2 1937.6 1788.1 1769. 1739.2 1756.3 1938.8 2013.3 1822. 1637.4 1534.6 1546.1 1760.7 1729. 1724.4 1735.5 1708.9 1817.4 1941.6 2006.3 1969.5 1729.9 1899.8 2012.1 1891.9 1903.6 1614.3 1566. 1706.4 1596.5 1830.3 1760. 1780.7 1932.1 1797.5 1645.5 1613.4 1858.5 1766.4 1573.2 1813.1 1713.8 1728.4 1800.8 1749.1 1690.1 1676.4 1826.4 1775.2 1772.5 1768.7 1628.9 1849.7 1772.5 1980.1 1776.1 1767. 1672. 1697.5 1766.7 1588. 1719.5 1586.9 1875. 1705.3 1570.4 1976. 1774.8 1648.8 1585.6] | [1300.9 1311.4 1230.5 1269. 1365.4 1360.4 1255. 1277.9 1322.7 1237.5 1247.8 1268.8 1313.3 1299. 1257.9 1255.1 1249.9 1243.3 1293.3 1330.2 1266. 1315.3 1319.5 1306.3 1272.6 1312.5 1310.8 1299.5 1329.7 1324.7 1313.2 1299.2 1271.5 1314.7 1298.9 1303.7 1258.9 1303.2 1383.5 1279.3 1283. 1258.6 1237.5 1347.1 1298.1 1352. 1347.1 1347.3 1436.6 1288.2 1276.3 1290.2 1265.3 1344.4 1236.5 1264.7 1301.1 1266.2 1312.4 1270.8 1271.5 1297.3 1340.8 1334.3 1311.6 1268.1 1254.4 1228.1 1285.7 1320.3 1296. 1295.7 1318.4 1268.2 1304.1 1275.2 1275.1 1331.7 1354.6 1337.5 1279.7 1243.4 1249.6 1264.6 1307. 1272.6 1315.4 1286.3] | [945.1 921.7 924.1 932.8 889.9 911.5 935.4 908.2 900.1 925. 935.4 908.7 939.9 937.2 875.7 921.1 888.6 911.7 942.1 938.8 898.9 922.5 948.2 947.7 882.8 901.4 881.4 906. 893.5 892.4 884.8 911.3 945.4 918. 919.6 866.7 860.3 902.6 905.1 936. 940.3 937.1 933.6 912.7 891.9 886.4 936.5 923.7 912.2 921.1 923.1 926.3 933.9 948.2 916.5 926.5 947.8 907.5 910.6 908.9 921.1 907.3 903.8 919.7 934.3 892.7 917.7 923.2 925. 913.1 925.9 928. 936.8 935. 880.9 933.8 929.8 928.3 919. 917.4 908.5 937.7 945.1 869. 902.4 943.2 949.1 898.2] | [661. 671.9 666.6 670.2 661.8 674.5 680. 684.3 679.3 662.4 678.4 657.4 665.4 675.1 672.1 687. 691. 672.5 663.6 681.5 665.4 655.8 690.6 668.9 660.8 671.7 673.9 672.7 676.8 665.8 669.7 675. 670.2 659.9 661.2 661.3 656.5 669.9 665.1 654.8 676.8 665.7 629.7 649.2 675.5 673.1 663.8 674.3 672.1 663.2 668.2 655.9 653.1 675.1 704.1 668.4 668.9 673.8 682. 690.8 664.3 646.5 668.9 682.7 704.4 651.1 676.7 665.5 668.2 681.3 661.2 656.4 664.3 652.3 662.3 658.3 643.4 658.2 673.6 668.1 659.5 666.8 681.7 663.5 649.8 658.5 663. 650.1] | [484.5 478.7 479.8 475.5 469. 469.8 486.6 480.8 485.2 487.3 477.7 479.4 478.8 480.2 475.3 486.7 469.4 482.2 481.6 466.6 475.8 477.1 476.3 469.8 480.3 481.3 471.6 472.1 483.9 485. 486.6 486. 485.6 468. 473.7 475.9 484.9 471.5 474.4 478.4 473.6 474.4 486.6 483.8 469.7 471.5 476.5 469.9 474.9 476.6 477. 476.1 483.6 481.2 471.8 477.3 475.7 471.1 474.8 471.9 483.2 488.6 476.6 463.1 473.9 482.7 487.6 479.3 472.7 457.5 466.2 488.1 475.8 476.2 470.9 477.2 469.8 476.6 473.9 473.1 488.2 475.1 472.6 475.8 475.2 473. 473.5 472.6] | [354.2 339.2 334.3 339.9 354.7 336.8 343.1 330. 332.3 345.6 354.7 361.5 352.9 339.1 325. 351.4 357.2 347.6 338. 324.1 324.6 341.1 339.6 350.8 357.3 342.5 342.4 337. 349.9 330.3 332.8 338.8 340.3 340.7 330.1 338. 357.4 349. 348. 340.9 345.9 337.5 340.6 339.1 342.1 345.6 342.9 333.5 349.7 347.2 358.2 346. 339.2 336. 336.4 337.7 341.1 342.3 340.6 340.1 346.2 343.9 336.7 345.3 344.4 339.2 344.6 346.7 350.8 345.8 343. 328. 327.9 332.3 338.8 346.4 346.4 334.1 323.6 333.6 340.7 340. 341.5 334.3 340.1 338.2] | [238.1 232. 243.5 242.5 231.2 236.1 240.2 246.3 238.1 238.2 244.3 243.2 242.2 242.3 243.3 238. 246. 242.1 232.8 235.8 237. 239.1 232.7 233.1 240.3 240.6 241.3 243. 243. 240.8 235.9 240.2 236.6 242. 246.2 249.9 245.4 243.7 240.9 241.1 242.8 243.8 241.6 245.2 242.6 240.4 241.1 249.1 243.4 240.5 233.9 233.4 233.1 241. 244.8 246.1 242.2 237.5 238.3 244.1 236.1 239.9 238.9 238.1 245.4 245.5 244. 234.4 236.5 238.7 237.5 238.6 232.7 233.3 234.7 239.8 239.3 239.1 237.3 244. 244.7 238.9 236. 235.1 244.1 249.5 248.1 241.2] | [178.9 194.9 190.4 172.4 179.2 186.6 168.2 175.6 183.7 171.9 159.1 176.8 181.4 168.3 194.6 160.6 180.9 192.2 173.8 169.1 188.2 183.2 164.8 168.6 173.5 166.6 182.4 183.8 171.1 157.1 178.1 195.8 179.2 162.6 170.9 177.5 178.3 167.9 180.5 171.7 169.4 173.4 171.7 178.3 181.6 173.5 182.3 173.9 173.3 175.9 191.8 191.1 176.8 177.5 181.9 171.2 182.6 179.4 173.3 177.2 182.6 181.1 175. 171.6 173.6 171.6 176.3 170.3 172.1 179.2 172.6 168.8 180. 180.3 179.8 182.7 169.7 180.9 185.3 171. 168.3 174.6 180.8 173.7 186.6 174.8 165.2 168.9] | [153.3 152.8 135.8 117.5 111.4 101.5 85. 126.2 158. 148. 107.2 112.1 113.8 105. 143.5 109.6 107.2 127.9 140.6 138.2 116.3 111. 107.4 110.9 123.3 121.4 121.1 106.6 100.2 137.8 131.3 139.8 137.2 121. 103. 143.9 143.5 130.7 119.6 130.3 112.4 120.5 125.5 135.1 115.9 147.4 135.3 126.2 123.1 115.3 123. 128.5 89.1 127.2 98.3 123.3 113.6 99.9 109.9 106.4 120.9 137. 113.7 136.7 117.6 128.7 131.8 129.3 115.8 103.2 111.9 116.1 110.8 112. 109.7 121. 140.7 135. 130.8 130.2 114.6 120.7 120.3 113.6 120.9 123.9 129. 134.8] | [ 54.8 123.4 104.7 97.1 96.7 86.6 85.2 78. 88.9 87.9 79.1 76.9 82.3 85.3 87.1 95.1 100.7 93.9 82.4 90.4 75.5 62.9 91.1 102. 112.5 86.3 87.9 108.8 73.3 95.7 100.3 67.9 92.1 95.3 79.9 89.5 10. 1. 1. 85.8 89.6 67.4 97.1 100.9 70.4 81.5 61.6 108.3 114.8 109.5 84.7 9.1 79.9 105.2 91.6 107.8 77.1 105.5 87.8 99. 108.4 90.2 55.6 92.4 75.3 98.4 89.3 120.1 98.6 87.3 60.8 2. 2.9 80.1 73.9 10. 91.8 84. 73.6 90.7 99.5 102.1 96.5 108.3 91.6 94.3 98.9 79.1] | [ 2.6 1.2 1.6 4.4 86.6 110.2 135.2 106.6 3. 103.7 81.4 1.2 67.1 99. 4. 2.6 2.8 2.4 85.7 116. 73.4 6.6 79.9 89.2 82.8 1.1 50.6 97.7 6. 7.6 6.6 100.2 1.8 10. 97. 62.4 4.2 93. 102. 6.1 95.6 112.5 118.3 100. 98.1 107.7 86.8 2. 100. 103.3 78.2 1. 67.4 9.5 93.4 42.8 92.2 134.4 10.1 123.1 108.8 83.5 2.4 7. 1.2 83.7 1.3 101.6 3.4 3.5 86.8 120. 89.4 3.4 5. 116.1 84.7 10. 81.4 2.8 8.1 5.9 86. 109.4 10. 2.6 2.8 4.9 9.5 71.8 4. 2.1 2.4 2.8] | [ 90.1 131.7 114.4 9.5 2.7 109.8 164. 10.7 3. 9.5 4. 4.4 1.3 9.5 127. 105.2 125.9 3.3 4.1 88.2 7. 114.1 112.5 107.6 1.8 50.8 127. 99. 1.8 78.4 83.6 85.2 9.5 119. 2.1 2. 3.6 82.2 2.4 1.2 1.2 114.3 89.9 7. 94.8 113.2 94.4 3.3 140.2 85.8 109.1 3. 2.4 66.3 10. 134. 140.8 93.3 9.5 120.2 1. 1.2 10. 91.2 90.1 128.6 91.9 108.8 92.5 4. 1.1 122.8 9.5 1.6 2. 2.8 1. 2.4 4. 3.2 10.2 9.5 1.2 105.9 73.2 92.3 6.9 3.3 85.9 1. 4.3 10. 10. 8.8] | [ 10.9 3. 1.8 1.8 119.9 4. 110.8 4.4 6. 4. 1.3 5.2 128.7 58.9 1. 1. 139.4 80.9 65.9 5.2 133.6 147.6 70. 121.7 8.5 112.9 132.6 10. 112.6 3. 3.2 6.4 117.8 187.4 81.4 70.7 165.5 104.1 3.7 104. 88.4 100.7 4.7 1.1 62.7 10. 4.9 105.2 2.6 9.2 135.9 75.1 65.3 1.1 7. 7.1 3.5 122.5 139.8 62.6 2.3 1.9 10. 56.3 138.6 5. 10. 110.3 86. 91.8 122.4 10.9 4.4 9.3 9.5 1.5 108. 154.1 9.1 7.8 146.1 10. 1.4 4.9 5.7 2.3 2. 124. 165. 1.2 10.6 3.1 10. 84.5] | [386.3 189.2 2. 7. 69.3 299.9 128. 1. 190.5 10. 3.7 1.3 10. 119.9 248.2 6.6 96.8 649.3 123.8 4.2 621.5 197.5 99.4 7.7 162.3 183.9 337.4 1.1 4. 246.3 234.1 4.8 221.6 262.4 40.8 295.3 174.7 238.9 561. 149.7 7. 109.8 7.3 184.5 133.4 5.2 715.8 211.4 8.6 265.2 148.7 6. 212.4 253.2 10.7 114. 112. 536.3 1.2 3.7 5. 210.9 229.4 252. 198.5 9.8 2. 242.6 1.2 152.8 213.1 210.7 203.6 156.8 2.3 5. 5. 2.1 10. 273.8 328.7 160.5 309.2 238.2 3.7 10.7 2.8 4.7 110.4 80.5 1.3 178.9 2. 9.5]
10.003 | wang_MDAnderson_NIST | https://osf.io/kszy2/download/ | 20200302_wang_MDAnderson_NIST/day2/20200302_wang_MDAnderson_day2_NIST_Magnitude_T1map.nii.gz | Magnitude | Chenyang Wang | MD Anderson | GE | DV26.0_R03_1831.b | 3 | NIST | 130 |  | 19.515 |  |  | IR Spin Echo TI=50,400,1100,2500 | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 122.109 |  |  |  |  | [1856.2 1864. 1781.1 1757.3 1818.6 1894.5 1897.5 1936.2 1886.4 1823. 1865. 1836. 1848.8 1873.3 1862.5 1813.8 1792.3 1845.1 1928.5 1865.3 1817. 1774.7 1758.2 1778.8 1845.9 1819.3 1779.6 1857.8 1841. 1870.7 1905. 1886.5 1795.5 1815.4 1815.1 1855.2 1844.9 1827.6 1859.3 1751.7 1818.4 1837.4 1823.5 1868.6 1866.1 1852.9 1922.6 1914.4 1874.6 1865.1 1785.9 1773.9 1751.3 1760.8 1807. 1828.8 1846.1 1897.2 1894.4 1867.8 1825.8 1856.3 1822.3 1763.3 1852.8 1850.9 1892.3 1922.5 1840.8 1836. 1846.7 1793.3 1884.3 1865.9 1846.6 1927. 1808.1 1848.9 1790.3 1796.6 1822.4 1885.3 1950.7 1795.1 1804.8 1838.6 1848.4 1805.6 1782.1 1852.5 1878.4 1828.6 1810.8 1771.6 1800.9 1842.1 1782.6] | [1327.4 1329.2 1278.9 1260.4 1276. 1317.1 1283.9 1315. 1402.3 1327.5 1305.7 1329.2 1358.9 1338.6 1355.9 1294. 1303.5 1334.6 1317.1 1303.8 1324.2 1337.9 1293.8 1284.8 1338.3 1306.6 1320.4 1318.6 1334.1 1281.5 1307.7 1336.1 1339.1 1330.6 1315.7 1301.4 1323.6 1310.1 1357.8 1325.3 1339.9 1312.3 1321.8 1339. 1309.2 1321.8 1366.3 1328.1 1352. 1333.2 1311.5 1295.6 1318.7 1333.4 1351.8 1335.3 1318.8 1293.3 1331.6 1361.5 1322.5 1309.8 1326.8 1317.2 1324.5 1315.1 1315.9 1343.4 1347.7 1342.2 1327.8 1311.6 1331.2 1318.9 1302.2 1333.4 1315.6 1311.2 1324.1 1347.6 1332.4 1324.8 1317.9 1336.3 1333.1 1339.7 1344.6 1360. ] | [932.1 935.5 935.1 947.7 940.9 940.8 928.2 936.2 951.3 942.2 929.2 916.5 927.5 947.8 933.9 934.9 935.1 928.6 937.1 927.2 946.1 952.7 950.9 939.1 938.3 930.5 940.8 938.6 944.4 916.7 940.9 945.1 941.8 953.6 950. 921.1 931.4 932.9 943.8 935.1 934. 929. 932.3 937. 949.6 929.3 950.8 932.3 927.7 933.2 920.7 933.1 951. 934.3 939. 926.6 942. 936.8 941.5 961.6 958.6 939.6 956.4 925. 934.6 944. 935.1 950.7 928.2 945.6 956.8 920.5 946.4 938.3 931.9 918.5 962.6 937.8 943.4 954.3 945.7 944.9 932.1 948.7 930.3 936.8 934.6 944. ] | [674.1 682.7 680.3 679.1 674.6 678.7 683.6 666.8 672.8 675.4 664.8 664. 674.5 665.9 673.1 676.3 686.5 679.4 675.5 663. 664.4 680.8 673.4 678.1 685.7 663.9 688.3 681.5 671.8 672.4 672.6 664.8 662.2 669.8 687.8 668.7 673.1 673.6 683.4 675.2 687.2 683.2 684.9 676.3 674.5 672.1 679.3 677.8 685. 662.5 685. 684.1 682. 662.2 676.5 673.1 694.6 686.6 670.5 661.2 683.4 675.3 681.4 676.3 666.6 683.9 681.4 675.6 673.3 672. 677.4 666.3 670.3 695.4 679. 667.2 668.7 669.9 669.5 685.8 689.9 671.1 673.2 673.1 670.4 667.2 670.8 692.1] | [490.6 496.5 490.6 490.6 489.4 487.1 479.7 478.2 474.9 482.2 487.5 479.1 481.4 479.1 486.9 486.8 487.6 484.4 484.2 482.4 481.3 485.7 484.3 483.9 484.9 485.4 480.6 485. 480. 481.7 480.1 481. 485.6 479.9 489.3 488.7 481.7 480.2 480.7 477.4 485.9 475.2 478. 484.3 492.4 474.8 478.9 483.2 484.7 473.5 489. 473.6 483.3 496.1 491.2 467.6 473.4 477.1 480.9 470.7 480.9 479.1 480.5 480.4 487.8 475.8 478.5 489.4 478.7 476.3 482.9 479.8 489.6 502.9 476.4 469.8 493.6 488.8 482.9 488.9 486.7 477.2 471.2] | [336. 344.6 342.4 346.8 342.2 336.4 326.2 344.1 348.2 344.8 348.7 351.5 341.3 331.2 320.7 336.5 347.9 342.7 336.5 348.7 347.8 340.8 341.8 352.1 349.6 347.3 341.3 339.8 332.5 345.9 343.9 348.3 347.7 340.3 339. 339. 341.7 338.2 337.1 338.8 348. 343.5 341.2 342.3 344. 348.9 343.4 339.2 341.1 338.1 342. 345.4 348.1 341.5 348.4 343.6 342.1 341.9 339.8 342.9 344.1 343.9 339.7 348.3 344.5 346.7 339.8 342.7 349.3 347.6 342.8 344.1 342.3 343. 349.2 342. 337.5 343.7 347. 340.4 340.9 344.3 340.5 339.5 338.8 344.1 354.5 344.4 336.2 339.2 348.8 348.6 344.3 345.7 337.7 337.7 344.2] | [251.6 242.1 239.2 243.5 243. 246.1 234.6 233.9 236.7 243.9 241.5 238.3 238.2 245.8 246.2 240. 236.8 237.7 239.9 237.7 243.1 243.7 239.9 242.6 236.9 237.7 239. 242.5 240.4 238.5 239.5 240.7 240.5 238.9 244.2 244. 240.8 244.4 239.5 238.2 236.6 243.6 246.5 241.9 238.5 239.8 238.9 245.3 240.9 241.5 238.8 243.5 243.6 241.4 237.3 237.5 237. 240.1 234.9 238.1 241.4 243.2 244.2 242.1 240.1 238. 239.1 241.2 237.8 241.8 242.9 240.7 240. 238. 243.9 244.2 240.1 239.7 242.8 241.5 244.7 244.5 240.3 243.8 248.7 245.3 243.2 243.6] | [261.4 175.1 178.1 177.8 178.9 175.1 172.5 175.6 177. 174. 174.2 176.1 174. 174.1 173.1 169.7 282. 289.9 170.7 180.4 298.4 173.6 278. 174. 174.9 173.8 272.9 174.8 164.8 274.9 174.1 173.7 179.6 171.4 171.6 171. 175. 181.3 174.8 179.4 275.8 176.2 275.1 172.7 179.1 178.1 171. 271.9 171.3 168.4 270.7 176.9 176.2 288. 171.8 283.8 176.8 173.6 177.7 178.3 179.9 178.1 175.6 172.6 273.7 270.5 178. 171.7 172.2 174.2 175.4 171. 269.9 175.1 175. 173. 175.4 177.6 177. 285. 175.3 177.8 175.7 173.9 172.9 173.6 172.9 171.3] | [116.1 113.5 140.5 118.4 143.2 122.7 115.2 124.4 150.3 125.1 118.3 120.1 141.2 119.7 123.9 122.7 141. 117.9 112.4 132.4 120.2 121.5 141.5 143.4 121.5 123.6 142. 145.1 113.1 109. 119.5 125.9 120.4 116.6 122.1 123.3 115.9 152.7 133.5 118.1 118.9 121.2 122.5 119.5 117.8 121.3 121.9 149.3 131.5 121.3 147.5 139.6 144.5 121. 115. 123.7 144.8 149.7 123.5 113.9 122. 135.4 137.9 143.2 125.1 123.9 138.1 141.3 112.5 125.5 141.9 118.1 142.2 124. 120.3 149.8 149.1 119.6 122. 140.2 122.1 122.1 123.2] | [ 84.8 92.3 80.8 83.8 82.3 83.1 88.6 88.5 96.4 87.6 86.4 99.2 87.1 74.2 87.7 86.6 87.3 91.4 80.5 61.3 80.2 79.7 85.2 101.2 97.5 87.8 90.8 85.5 93.7 102.6 87.1 61.1 72.7 85.2 92.8 83.5 99.1 79.9 87.7 93.1 83.4 73.5 86.4 91.2 103.1 82.6 84.1 84. 95.1 86.2 99.1 95.2 87.4 80.9 103.3 95.1 88.3 99.2 77.3 2.7 75.8 86.1 83.7 79.8 51.9 83.4 80.3 84.7 67. 101.3 95.1 88. 77.2 89.8 2.7 97.6 96.4 65.1 87.6 75.7 82.9 93.2 4.8 2.5 76.3 84. 81.1 64.9] | [58.2 43.1 4. 62. 2.8 10. 63.3 56.9 2.6 9.5 1.9 62.7 1.3 1.2 61. 67.4 60.4 80. 66.5 1.3 79.6 4. 60. 1.8 55.7 6. 79.3 46.8 59.1 96.5 3.3 70.9 66. 78.4 3.4 1.8 1.2 70. 49.8 61.2 92.5 1.8 81.3 3.3 1. 64.1 78.6 58. 70.1 76.7 57.4 64.7 3.1 73.2 2.4 2.7 59.1 9.5 2. 65.1 65. 1.3 85.4 51.7 64.6 89.8 90.9 84.9 72.3 77.2 68. 59. 3.4 2.4 1.2 6.6 6.1 2.9 10. 1.3 63.9 1.1 2.8 73.7 83.2 66.5 77.8 78.3 1.6 1.2 3.2 3.4 2. 2.4] | [ 2.5 65.6 60.9 3. 79.4 51.6 4.9 1.8 8. 1.8 1. 1.6 3.3 9.1 62.1 8.1 3.1 51.1 3.1 1. 3.2 2.5 71.4 2.4 2.3 45.9 2. 79.6 59. 3. 1.2 9.1 9. 5. 2. 78.9 50.2 2.5 10. 3.3 1.1 72.3 91.6 53.5 3. 3. 54.8 4.7 85.3 74.3 115.9 1. 2.5 3. 3.1 4.2 69. 91.9 1.3 2. 1.8 92.4 4. 58. 54.4 76. 3.3 2.1 4.6 1.9 94.8 118.8 9.5 1.8 1.3 3.4 2.8 1.4 62.7 1.2 3.3 3.3 3.4 105.8 39.2 9.4 3. 63.9 73.5 2.2 3. 2.5 53.1 1. ] | [ 79.9 116.5 72.5 109.1 122.2 2. 1.8 3.6 1.5 1.2 9.5 3.4 3.1 112.5 74.4 103.4 101.8 116.5 80.6 87.8 74.6 72.9 1.9 77.2 4. 4.9 2.5 2.8 1. 1.3 2.6 4. 10.2 95.7 2. 1.2 71.5 3. 99.8 4.6 110.4 67.6 128.1 1.3 7.9 2.6 3. 6.6 3.4 63.8 1.1 68.8 5. 108.3 2.6 114.3 83.4 9.5 1.8 105.2 113.1 1.8 74.5 89.6 131.6 5. 57.5 1.8 4. 2.4 83.2 90.2 62. 67.9 2.1 4.3 86.9 1.8 2. 2.5 1.2 73.4 58. 3.6 2.8 2.5 85.7 4.3 66.9 59.1 5.3 3.3 2.1 2.5] | [ 69.7 145.3 209.7 85.5 50.1 76.6 2.8 7. 2. 63.2 62.2 126.7 120.1 4. 5. 66.2 210.4 169.9 63.5 82. 2.4 9.6 174.8 4. 8.3 1.9 2.8 2.7 2.5 127.4 149.4 176.7 3. 1.7 6.3 71. 66.1 75.6 1.2 67.3 2.9 3.6 63. 2.5 120.9 4.1 44.2 7. 59.1 2. 129.7 158. 183.8 108.4 3.3 1.1 4.6 106.3 54.6 153.7 121.6 76.2 2.5 3.6 1.2 121.1 40.6 70.9 2.5 3.3 1.1 5.6 155.2 140.1 95.7 1. 1.1 1.8 68.6 133.8 1.8 1.1 7. 2.5 179.2 3.1 1.2 4.8 51.3 62. 138.4 3.3 1.3 2.8]
10.004 | wang_MDAnderson_NIST | https://osf.io/kszy2/download/ | 20200302_wang_MDAnderson_NIST/day2/20200302_wang_MDAnderson_day2_NIST_Complex_T1map.nii.gz | Complex | Chenyang Wang | MD Anderson | GE | DV26.0_R03_1831.b | 3 | NIST | 130 |  | 19.515 |  |  | IR Spin Echo TI=50,400,1100,2500 | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 122.109 |  |  |  |  | [1865.1 1868.9 1784.4 1761.3 1822. 1900.7 1906.3 1944.2 1890. 1827.6 1870.2 1843.7 1854.6 1879.2 1868.4 1818.9 1799.4 1853.2 1934.5 1870.8 1822.4 1777.5 1762.8 1782.8 1851.8 1825.8 1784.1 1863.1 1849.1 1875.5 1911.7 1891.8 1799.7 1820.3 1819. 1858.3 1850.3 1830.9 1863.5 1759.7 1824.4 1842.4 1827.2 1872.9 1869.5 1857.5 1926.2 1919.4 1878.6 1869.8 1791.9 1775.3 1755.2 1763.8 1811.5 1833.5 1850.4 1900.7 1898.8 1873.2 1831.4 1862.7 1825.8 1765.9 1856.3 1853.9 1897.5 1927.4 1845.7 1841.1 1852. 1798. 1889.3 1871.9 1852.3 1933.4 1813.5 1854.8 1795.1 1802.3 1827.6 1891.2 1956.1 1800.5 1808.3 1843.4 1854.5 1810.3 1787.6 1856.6 1881.8 1834.1 1815.1 1776.3 1807. 1846. 1785.6] | [1330.3 1332.4 1280.5 1261.8 1279. 1319.3 1285.4 1317.1 1405.3 1329.5 1307.9 1331.9 1361. 1341.2 1358.9 1296.1 1305.4 1337.8 1320.3 1306.2 1325.7 1340.7 1295.7 1286.5 1341.1 1308.4 1322.4 1320.2 1337.2 1283.7 1309.2 1338.5 1341.4 1332.4 1317.7 1304.6 1324.9 1311.5 1359.8 1328.4 1343. 1314.4 1323.7 1341. 1311. 1324.3 1367.9 1330.2 1354.5 1335. 1314.2 1297.4 1322.1 1336.2 1354.2 1338.1 1321.1 1295.3 1332.9 1364. 1324.4 1311.4 1328.8 1318.9 1326.2 1316.9 1319.5 1347.2 1348.8 1344.2 1329.8 1312.5 1333.1 1321.8 1305.4 1336.5 1317.8 1311.9 1326.5 1350.9 1335.6 1326.5 1320.8 1338.8 1335.4 1342.6 1347.1 1361.9] | [932.5 936. 935.7 948. 941.6 941.9 929.1 937.3 952.2 942.9 929.8 917.5 928.5 948.3 934.3 935.9 935.6 928.9 938. 927.8 947. 953.2 951.7 940. 938.7 931.3 942.1 939.7 945.4 917.1 942.2 946.3 943. 953.9 950.5 922.4 932.4 933.2 944.4 935.8 934.6 930.2 933.3 937.2 949.9 930.3 952.1 933.4 928.4 933.3 920.9 933.8 951.2 934.9 939.6 926.7 942.2 937.8 942.6 962.7 958.8 939.9 956.4 925.6 935.9 944.8 935.7 951.5 928.8 946.9 958.2 920.9 946.7 939.1 932.9 919.8 964. 938.5 943.5 954.3 946.7 945.7 932.2 949.1 929.5 936.8 934.8 944.3] | [674.6 683.4 680.5 679.5 675.1 679.3 683.5 667.3 673.5 675.5 665.6 664.5 675.2 666.2 673.6 676.8 687.2 679.7 675.6 663.9 665.3 681.4 674.4 678.3 686.4 664.4 688.9 682.2 672. 673. 673.2 665.3 663.5 670.7 688.1 669.1 673.4 673.9 683.5 675.7 687.8 683.5 685.9 676.5 674.8 672.5 679.7 678.2 685. 662.7 685.3 684.3 682.6 662.7 676.9 673.2 694.8 687.4 670.8 661.7 683.9 675.9 681.9 676.5 666.8 684.6 681.5 675.8 673.3 672.2 677.7 667.1 670.7 696. 679.3 667.5 669. 670.5 670.3 686.1 690.3 671.7 673.5 673.7 670.8 667.3 671.1 693.5] | [490.7 496.7 490.9 490.7 489.6 487.3 479.7 478.3 474.8 482.3 487.5 479.3 481.6 479.2 487.2 486.9 487.7 484.4 484.2 482.3 481.5 485.8 484.6 483.9 485.1 485.9 480.8 485.1 480.2 481.9 480.4 481.2 485.8 479.9 489.5 489. 481.8 480.2 480.7 477.5 486.1 475.4 478.1 484.4 492.5 474.7 479. 483.2 484.7 473.9 489.1 473.7 483.3 496.3 491.7 467.8 473.7 477.3 480.9 471. 481.4 479.6 480.7 480.5 487.8 475.9 478.8 489.9 478.8 476.4 482.9 479.9 489.7 503.1 476.5 470.1 493.8 489.1 483.2 489.2 486.8 477.2 471.1] | [336.3 344.8 342.4 347.2 342.6 336.4 325.9 344. 348.3 345. 348.8 351.7 341.3 331.2 321. 336.5 347.8 342.9 336.7 348.8 347.8 340.9 341.8 352. 349.6 347.5 341.2 339.9 332.5 345.9 344.1 348.5 347.9 340.4 339. 339. 341.7 338.2 337.1 338.7 348.3 343.7 341.2 342.3 343.9 349. 343.5 339.3 341.1 338.4 342.2 345.8 348.3 341.5 348.5 343.5 342.1 342.1 339.9 342.9 344.2 343.9 339.9 348.6 344.8 346.7 339.8 342.6 349.3 347.7 342.7 344.1 342.4 343. 349.4 342.3 337.5 343.7 347.1 340.5 341.2 344.5 340.5 339.6 339. 344.4 354.7 344.3 336.4 339.3 348.9 348.7 344.2 345.7 337.8 337.6 344.2] | [251.8 242.2 239.3 243.4 243.2 246.3 234.7 233.9 236.6 243.8 241.4 238.6 238.6 245.9 246.2 240. 236.8 237.7 240. 237.5 243. 243.9 240. 242.6 236.9 237.8 239. 242.6 240.5 238.4 239.6 240.6 240.6 239. 244.3 243.9 240.6 244.4 239.6 238.3 236.7 243.6 246.4 241.8 238.7 239.9 239.1 245.4 240.8 241.6 238.8 243.4 243.6 241.3 237.3 237.6 237.2 240.1 234.8 238. 241.3 243.2 244.2 242.2 240. 238.1 239.2 241.1 237.8 241.8 242.9 240.8 240.1 237.8 243.9 244.3 240.1 239.8 242.9 241.5 244.6 244.5 240.4 243.7 248.6 245.4 243.2 243.6] | [170.2 175. 178.1 177.7 178.8 175. 172.5 175.8 176.9 173.9 174.1 176.1 173.8 174. 173. 169.7 178.7 181.1 170.6 180.4 183.7 173.6 175.7 173.9 174.8 173.9 175.7 174.8 164.8 175.8 173.9 173.7 179.6 171.2 171.6 171.1 175.1 181.2 174.6 179.5 174.7 176.2 175.4 172.5 179.1 178. 171.2 175.5 171.2 168.4 173.6 176.9 176.3 179.8 171.8 176.9 176.6 173.6 177.7 178.3 179.9 178.1 175.6 172.5 172.3 172.3 177.9 171.6 172.2 174.1 175.4 171. 173.7 175.1 175.1 173. 175.4 177.7 177. 178.8 175.3 177.8 175.6 173.9 172.8 173.6 172.8 171.2] | [116. 113.5 122.5 118.2 124.7 122.5 115. 124.3 129.8 124.9 118.2 120. 122.4 119.4 124. 122.6 123. 117.9 112.3 116.4 120.1 121.4 123. 124.4 121.4 123.5 123.6 125.5 112.9 108.9 119.5 125.8 120.2 116.5 122. 123.2 115.9 131.1 117. 118.1 118.9 121.1 122.3 119.5 117.8 121.2 121.8 129.7 115.2 121.2 126.5 121.8 126. 120.9 114.9 123.6 125.5 129.8 123.5 113.8 121.9 118.6 120. 124.1 125.2 123.8 121. 123.1 112.4 125.4 122.9 118. 122.9 123.9 120.2 129.2 128.7 119.6 121.8 122. 122. 122. 123.1] | [ 84.8 93. 82.2 85.1 82.5 83.4 89.6 88.2 96.8 88.2 86.7 99.2 87.5 74.6 87.7 86.7 87.5 92.5 81.1 62. 80.7 80.2 85.3 101.9 97.5 87.7 90.8 86.2 94.3 102.5 87.5 61.4 73. 86.1 93.4 83.7 99.3 80.3 88.5 94.1 84.2 74.1 87.2 91.3 102.8 83.8 84.6 84.6 95.2 86.3 99.7 96.3 88.1 81.3 103.3 95.3 88.5 99.2 77.2 2.6 75.9 86.7 84.1 79.7 52.2 83.4 80.8 85.5 68. 102.2 95.7 88. 77.5 90.2 10.6 98.1 96.4 65.8 88.6 76.3 83.4 93.5 2.7 2. 76.7 84. 81. 65. ] | [58.3 49.5 10.9 68.1 40.8 5. 64.3 58.5 1. 10. 1.3 64.3 10. 4. 63.1 75.9 66.8 89. 67.7 9.5 80.1 11. 67.1 10.9 62.4 10. 80.4 46.2 65.4 97.5 3.4 71.6 68.2 80.4 2. 1. 10. 71.9 52.4 62.4 93.1 7.6 90.8 2.7 1.8 71.7 88.5 64.5 71.1 77.9 64.9 71.2 1. 80.8 3. 2.8 64.8 4.9 10. 71.2 66.9 4.8 85.6 56.3 72.2 90.1 91.1 85.8 73. 77.6 76.7 61.2 1.3 1.2 5. 3.5 4. 9.5 1.1 1.8 64.4 3.7 9.5 74.8 83.3 67.4 78.7 78.4 8. 2.8 2.5 4.8 5.7 7.6] | [ 3.1 71.5 78.2 5.8 85.6 69.1 1.9 2.6 5. 3.3 3.3 4. 3.2 2. 69.7 61.2 3.3 66.8 2.7 2.7 53.8 57.9 75. 3.7 1.2 66.2 1. 82. 67.8 4.3 8. 2. 1.8 9.5 2.4 101. 60.7 7.3 5. 61. 56. 78.4 94.8 67.8 54.7 3.6 61.7 1.9 89.2 77.9 117.4 2.3 9.5 8.3 8.3 4. 73.4 93.4 6.3 2.5 8. 94.5 4. 64.5 66.2 80. 42.6 51.7 1.9 3.3 97.6 120.5 2.9 10.9 4.9 4.1 4. 3.3 68.8 2.5 8. 6. 10. 106.8 48.2 10.2 1.3 80.3 93.9 1.3 9.5 5.7 64.2 9.5] | [123. 118. 88.7 110.4 124.5 10.7 7.1 1.4 5. 4.9 9.5 62.4 6.7 111. 114.7 112.7 102.2 110.3 82.7 98.6 108.1 105.7 2. 81.6 3.1 4.1 9.5 6.2 6.3 1.9 2.7 3.7 1.4 99.4 8. 3. 105. 8.5 100.6 8. 109.8 104.4 133.7 4.3 4.3 1.1 7.1 67. 1.2 99.2 1.6 104.8 9.5 109.7 7.2 115.9 89.3 1.1 6.7 112.8 118.5 11. 107.4 93.5 134.6 1.4 73. 2. 5. 4. 92.7 95.5 90.2 102.6 9.5 8. 85.8 5.3 67.5 1.6 5. 94.5 88.6 9.5 6.7 10. 92.8 10.2 83.1 89.7 6. 10. 1.3 5.3] | [109.4 135.6 204.6 189. 77.8 138. 10.9 2.4 83.9 10.7 89.9 127.8 114.1 1.9 10.6 110.9 208.9 169.5 83.4 88.5 66.5 2.8 163.7 105.2 79.3 9.5 3.6 1. 10.7 122.7 139.8 146.4 10.2 2. 4.1 135.6 85. 135.5 102.7 121.9 51.5 4. 125.8 83.9 129.1 10.7 109. 10. 121.3 2. 122.5 142. 166.7 144.7 1.8 3.6 2.2 129.5 83. 127.1 96.8 114.2 110.1 80.1 9. 120.6 81.8 131.3 10.6 5. 5.9 89.9 164.6 165. 126.6 4.2 4.5 85.9 109. 142.8 88.1 9.2 4.8 10. 153.8 10.3 10.7 1. 107.7 111.7 140.4 10.6 10.2 1.3]
11.001 | ngmaforo_ucla_NIST | https://osf.io/ab2zg/download/ | 20200305_ngmaforo_ucla_NIST/20200305_ngmaforo_ucla_Skyra_noFS_NIST_Magnitude_T1map.nii.gz | Magnitude | Nyasha G. Maforo | University of California, Los Angeles | Siemens | Skyra VE11C | 3 | NIST | 130 | 46 | 21.5 |  |  | t2_tse_dark-fluid_cor | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 17 | [50, 400, 1100, 2400] | 283 |  |  |  |  | [1977.3 1884.1 1991.3 1978.9 1846.1 2064.6 2154.7 1972. 1931.1 1912.4 1870. 2034.4 2275.8 2060.9 2076.7 2030.5 2011.8 1989.5 1853.6 1899.5 1862.9 1961.9 2118.6 1914.3 1913.8 2016.2 2048.6 2036.1 1978.2 1922.7 1896.6 1986.9 2077.6 2164.3 2073.7 1983.8 1944.2 1967.1 1881.3 1940.2 1853. 1915.1 1944.5 2156.5 2087. 2145.6 2106.2 1965.2 2019.8 1959.5 1993.5 1889.1 2020.1 2057.6 2000.7 1951.1 2023.2 2011.3 1940.9 1930.3 1934.9 1976.9 2035. 2027.7 1909.3 2128.5 2050.5 2078.5 2137.8 2095. 1914.8 1865.1 1933.8 1991.8 1900. 1846.9 2035.3 2007.3 2095.7 1953.8 1900.1 1937.7 1960.2 1965. 1936.6 1998.1 2043.1 1891.9 1906.8 1839.7 1808.5 2002. 2065.8 2088.6 2151. 2027.8 1988.2] | [1347.8 1389.3 1361.3 1381.9 1383.3 1382.3 1380. 1366.2 1442.2 1368.9 1320. 1321.8 1359.4 1400.9 1470.9 1436.2 1368.9 1329.7 1369.3 1361. 1382.1 1366.3 1303.9 1359.1 1438.1 1359.8 1380.1 1347.9 1361.8 1308.7 1383.6 1428.4 1332. 1392.8 1321.8 1325.8 1375.2 1345.4 1369.4 1421.5 1439. 1370.6 1304.6 1367.6 1325. 1350.2 1431.2 1361.7 1351.5 1391. 1376.2 1394.3 1404.4 1405.7 1454.4 1313.7 1359.8 1365.4 1400.6 1380.1 1326.5 1372.2 1404.2 1414.3 1392.7 1386.4 1365.3 1406.3 1384.8 1352.2 1349.7 1317.9 1333.5 1387.4 1375. 1360. 1432.3 1354.1 1378.1 1444.3 1334.7 1390.5 1358.6 1403.3 1390.2 1368.4 1354.2 1373.6] | [ 979. 970.4 946.4 951.8 980.3 994.1 972.3 982.1 982.3 995.4 993.7 981.8 976.6 998.4 985.1 990.4 1037.1 1024.1 1010.4 1005.7 1004. 984.1 1007.8 982.7 984.4 1002. 1014.8 1040.1 1004. 1004.5 1035.8 995.7 1004.2 1007.9 966.7 1011.2 991.4 987. 1021.2 1025.3 997.4 1011.5 1009.8 1012. 965.4 998.5 984.8 981.9 1021.1 1006.7 1003.6 1021. 981.2 1033.6 961.6 980.1 1003.5 1032.3 997.9 947.3 973.3 981.9 974.3 1004.4 954.5 979.7 992.9 963. 995.6 1022.5 973.7 975.7 1008.9 957.7 975.4 982.1 992.8 970.8 972.8 991.4 1017.5 1014.7 955. 989.1 971. 973. 987.9 975.3] | [742.8 668.2 690. 678.3 680.9 706.3 701. 695.5 699.1 729.2 704.6 698.8 678.3 664. 742.5 688. 727.6 694.7 676.9 676.9 690.6 708.7 692.4 694.8 707.5 719.2 691.3 699. 713.3 716.6 696.8 681.5 720. 680.3 674.8 704.9 692.5 734.1 717. 701.1 687.9 683.4 706.6 683.1 682.2 721.1 684.1 708.4 689.8 688.3 700.8 687.1 698.4 679.4 695.6 726.5 684.2 693.9 709.5 685.4 707.2 669.8 701.6 681.3 719.1 697.1 704. 683.2 702.3 687.8 684. 684.9 698. 687.8 698.3 707.7 725.4 703.4 683.7 706.4 711.4 690.3 697.9 690.6 703.3 701.6 683.8 666.2] | [490.1 482.1 454.4 464.4 463.1 456.1 488.1 476.2 473.4 468.9 469. 471.3 469.4 470.4 476.5 488.1 452.7 473.4 474.3 453.7 441. 451.3 473.7 458.2 487.8 489.7 470.6 455. 458. 471.1 463.2 459.1 461. 482. 481.1 472.5 479. 459.4 457.2 458.9 449.2 467.8 453.3 469.7 493.5 487. 478.5 481.6 468.5 471.5 466. 486.6 461.5 467.8 481.7 498.3 469.3 491.5 459.8 467.7 468. 460.4 451.1 462.8 460.3 495.7 472.5 480.8 482.5 477.7 463.2 475. 481.2 460.6 472.2 494.9 468.4 469.1 479.5 469.6 475.1 457.8 456.3 488.3 480.4 474.5 467. 474. ] | [345.5 318.7 319. 336.8 347.7 333. 346.2 356. 351.7 353.6 347.7 338.7 337.3 349.6 344.7 341. 332.2 329.7 342.9 357.7 354.7 338.9 346. 349.3 349.7 349.3 347.5 343.7 355.7 339. 338.2 338.1 358.4 342.9 337.9 340.3 328.6 327.1 325.4 342.6 333.3 331.9 357.6 364.2 339.2 338.9 331.8 331.8 331.8 337.7 340.4 348.4 335.7 355.4 347. 337.8 351.4 344.4 349.6 325.4 337.5 328. 346.4 338.5 346.3 357.8 340.5 342.8 329.1 342.7 339.5 345.1 318.8 343.8 342.7 351.2 342. 338.7 329.6 333.6 340.8 341.9 336. 351.5 337.1 338.7 342.3 344.7 336.1 342.9 347.2 345.4 335. 339.6 336.9 339.6 327.9] | [256.4 254.9 251. 257.1 260.5 254.7 251.7 249.5 246.5 238.4 239.8 242.7 235.2 250.7 246.9 246.6 239.5 249.6 249.7 241. 245.6 241. 236.7 250.2 260.9 243.9 248.4 236.2 236.5 230.3 228.4 239.2 242.4 229.9 250.7 244.9 243.6 240.4 245. 242.5 245.3 232.3 247.3 240. 243.1 251.8 239.3 235.3 237.1 231.2 244.1 234.7 241.4 232.1 240.2 251.8 240.9 236.4 242.3 236.7 242.6 232.4 240.4 220.2 244.8 235.8 232.7 233.5 242.9 238.2 233.1 231.1 241.5 221.2 230.3 241.5 247.7 228.8 225.1 235.5 243.3 235.4 237.3 244.9 246.8 246.1 237.9 225.5] | [182. 181.6 182. 184.3 183.7 177.9 178.2 181.2 181.1 183.4 186. 184.2 190.8 177.1 181.8 173.5 190. 165.1 173.8 181. 162.6 172.8 187.3 173.5 170.5 636.7 171.6 483.8 178.7 183.5 189.6 179.6 178.4 182. 174.9 197.8 176.2 188.5 168. 184.8 167.6 184.8 172.4 590.1 175.5 190.8 181. 188.1 179.1 835.4 164.7 184.2 173.5 191.9 175.4 186.9 173.4 612.7 172.5 168.8 181. 498.8 163. 181.3 173.2 183.4 182.6 177.1 190.5 186.5 187.6 164.8 180.8 175.8 179. 189.1 181. 174.4 164.7 173. 177.9 176.8 169.2 176.7 187.3 178.9 185. 180.2] | [199.4 179.3 185.2 226.7 142.2 121. 202.2 121.5 231. 194.2 190.6 190.3 203.4 202.7 224.9 123. 214.5 206.4 183. 203.2 208. 189.3 195.3 188.1 196.8 202.8 202.9 188.5 220.5 187.6 191.7 212.9 164.3 216.2 205.7 247.2 197.7 210.7 180.5 203.4 193.8 205.6 189.6 210.6 196.4 198. 171.7 199.4 164. 125.7 177.1 180.6 180.6 195.8 186.6 185. 199.9 192.7 223.1 181.8 182.8 202.9 166.8 218.5 240.7 182.7 220.9 175.3 170.6 199.1 183.2 125.5 183.1 192.9 211.7 195.9 195.9 179.2 125.4 152.4 205.3 209.4 183. 156.6 171.9 157.2 173.4 187.9] | [114.9 136.4 128.9 135.2 147.2 136.1 133.1 144.5 136.3 98.8 105.6 129.4 142.9 154.3 126.4 113.3 4. 116.4 132.3 132.7 124.8 104.7 95. 126.3 140.7 10. 122.6 127.9 97. 99. 106.2 134.7 92.4 3.2 97.6 96.6 105.1 9.5 120. 134.7 119.3 100.8 115.5 114.6 89. 136.9 108.8 118.3 118.9 99.5 115.5 131.9 121.3 130.2 112. 118.2 115.2 127.9 88.5 115.3 92.7 123.3 104.1 124.5 101. 115.8 119.7 91.9 119.3 110.5 121.7 122.2 103.6 125.8 110.9 137.6 77.6 119.8 99.8 107.4 74.2 115.1 117.2 131.6 92.8 79.2 85.1 81.1] | [107.1 113.7 126.7 134.4 76.6 2. 91.7 90.2 113.6 75.6 121.4 84.9 106.8 111.7 117.4 111. 138.1 95.9 86.7 140.3 115.4 128.3 142.4 74.2 113. 111.3 3.2 1. 122.1 115.8 110.3 126.2 110.3 125.1 126. 96.5 116.6 97.9 113.2 86.5 96.1 132.1 126.5 121.2 121.4 91.7 77. 106.5 121.5 108.9 125. 123.9 129.1 104.7 120.2 120.3 135.4 139. 101. 133.2 119. 102.1 148.4 136.2 106.2 81.9 118. 115. 101.2 122.5 108.9 104.9 108.6 109.6 127.5 116.4 120.5 116.1 123.8 102.8 117.4 129.5 104.8 150.2 129.9 121.3 104.8 128.6 141.2 122. 116.1 130.6 115.8 108.9] | [165.7 148. 153.6 95.4 139.5 126. 137.4 119. 120.2 123.9 3.2 103.4 1.8 119.9 1.3 88.3 116.9 116. 78.3 120.4 139. 109.9 136. 104.5 107.4 148.4 57.1 94.5 105.7 128.5 122.1 141.5 107.2 95.9 129.9 78.3 92.5 131.1 133.7 143.8 127. 144.5 147.9 108.5 151.4 157.2 123.5 144.6 103. 153.6 1.2 115.3 116.4 105. 121. 131.6 130. 113.9 144.7 4.7 147.8 97. 139.7 130.2 136.2 169.1 117.4 154.6 135.5 148.2 128.7 65.2 119.5 87.7 129.8 131.9 148.5 123.8 100.4 131.1 53.6 159.6 130.4 121.1 108.7 131.7 162.6 102.6 115.5 100.4 136.5 104.3 129.8 147.6 86.1 165.3 127.8 78.1 109.3 124.6 120.2 1.1 154.8 118. 86.9] | [133.1 146.7 193. 171.7 168.5 153.2 229.7 234.3 159.1 129.2 105.1 145.9 173.9 166.2 224.4 158.4 99.7 126.3 139.3 164.2 84.7 55.7 193.4 147.4 166.5 160.5 87.9 82.9 169.5 172.2 2.4 142.3 222.3 169.5 151.7 159.7 93.9 130. 166.6 123.2 107.3 166.1 114.6 176.1 106.8 196.4 194.5 199.3 191.9 217.8 141.1 164.2 191.2 146.9 202.3 219.4 219.9 126. 159.5 110.2 184.2 115.1 207.8 138.6 125. 170.8 100.1 169.8 115. 113. 159.9 102.2 129.6 158.7 130.2 54.3 149.6 185.3 184.6 92.9 127.9 131. 58. 80.2 174.7 4. 173.5 122.9 125.1 162.1 153. 147.7 202.9 162.9 146. 152. 1.2 2.5 140.5 95. 179.5 96.2 155.1 166.2 86.5] | [236.1 219. 245. 306.2 119. 65. 262.5 212.9 224.1 114.9 165.2 188.3 120. 335.1 196. 149.1 320.9 103.9 1.3 122.2 369.3 300.6 242.1 204.9 268.8 103.5 235.7 246.2 283.1 309.3 215.7 341.9 530.5 543. 216.9 289.5 280.9 286. 383.7 267.2 222.5 321. 191.4 224.1 159.9 205.1 328.3 276.5 212.3 344.9 435. 241.2 305.7 213.6 134.9 270.8 533.1 543.4 321.1 473.1 623.4 613.6 219.6 382.6 69.5 417.6 431.2 596.6 261.3 394.2 301.4 726.5 277.1 510.7 401.5 313.7 202. 197.4 84.3 198.3 9.1 386.3 392.2 505.7 296. 229. 371.3 480.7 313.6 112.6 253.6 264. 191.5 258.8]
11.002 | ngmaforo_ucla_NIST | https://osf.io/ab2zg/download/ | 20200305_ngmaforo_ucla_NIST/20200305_ngmaforo_ucla_Skyra_FS_NIST_Magnitude_T1map.nii.gz | Magnitude | Nyasha G. Maforo | University of California, Los Angeles | Siemens | Skyra VE11C | 3 | NIST | 130 | 46 | 21.5 |  |  | t2_tse_dark-fluid_cor | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 17 | [50, 400, 1100, 2400] | 283 |  |  |  |  | [2142.4 1891.8 1829.4 1964.1 2091.6 1959.9 2089.3 1982.7 1905.1 1951.5 1901.5 1963. 2050.3 1972.2 2085.9 1901.4 1930.1 2018.1 1959.3 2041.6 2077.9 1972. 2092.8 2009.1 1976.2 1881.3 2048.6 1978.2 1926. 1966.2 2040. 2027. 2136.8 2159.7 2021.3 2013.4 2004.7 2040.9 1916.6 1976.9 1909.3 2013.7 1975. 1996.4 1924.9 1998.1 1985.2 2055.4 1991.8 1851.3 1961.9 1929.9 2027.8 1970.3 1994.3 1915.5 2057.5 1855.7 1976.9 2053.2 1954.1 1850.6 1883.7 1911. 1848.5 2054.7 1955.1 2083.2 1979.9 2024.4 1875. 1972.6 1901.2 1943.7 2004. 1951.2 2071. 2126.4 1993.6 1998.4 1866.6 2086. 2070.7 1925.7 1978.6 1999.5 1969.1 1984.1 1990.8 2012.8 2037.9 1940. 2018.1 1889.9 1933.3 1907.2 2004.8] | [1304.9 1346.3 1400.1 1396.6 1367.2 1357.5 1279.1 1395.9 1436. 1366.4 1332.5 1334.6 1377.5 1421.6 1440.4 1351.3 1388.1 1328.5 1329.9 1334.9 1365.4 1413.5 1377.3 1335.2 1384.5 1428.6 1374.1 1318.8 1457.1 1459.9 1388.7 1364. 1370.8 1422.6 1421.8 1401.5 1390.3 1348.7 1331.4 1382.7 1408. 1343.3 1315. 1388.2 1379. 1315.8 1384.1 1429.5 1371.4 1344.7 1368.3 1387.6 1325.9 1348.4 1408.7 1347.1 1335.5 1371.5 1410.1 1412.8 1374. 1377.4 1350.6 1366.2 1466. 1408.9 1326.4 1355.3 1321. 1326.4 1380.2 1397.1 1388.5 1379.2 1473.8 1427.6 1434.4 1405.8 1388.7 1388.9 1339. 1331.2 1402.8 1418.4 1437.9 1359.3 1285.6 1336.8] | [ 967.6 1000.8 954.7 963.1 1002.2 964.4 977.5 945.9 993.7 987.5 1003.7 966.1 941.1 939.1 993.3 1001.5 997.4 999.5 1005.2 1018.5 940.3 962.2 1011.8 993.3 967. 985.7 972.1 1004.8 983.2 1010.7 995.1 1003.9 1029.2 1024.3 1009.3 1005. 999.7 986.8 964.5 1004.2 993. 1000.4 975.1 1014.6 997.7 990.9 1009.4 994.3 969.6 995. 985. 1014.8 985. 998.4 951.1 1015.2 974.2 1007.6 1005.6 978.9 992.8 1008.8 992.8 975.2 950.7 975.5 949.4 964.2 1008. 989. 969. 986.9 988.3 966.9 938. 1010.7 981.1 982.3 1007.7 995.1 1014.9 996.7 982.8 968.5 983.7 1005.4 1007.9 980.9] | [712.6 681.5 674.8 667.2 679.1 700.4 708.4 690.1 710.4 701.7 695.3 709. 688.7 682.7 709.8 695.9 720.2 678.2 678.1 689.4 667.6 686.7 704.4 699. 688.4 721.3 691.9 689.4 702.9 709. 688.5 684.3 719.1 672.4 666.8 717.2 679.3 710.2 683.9 689.6 709. 691.6 676.8 664.8 670. 705. 671.5 709.7 702.5 688. 700.1 672.2 674.8 686.5 688.7 704.8 697.5 701.3 713.9 689.1 711.4 681.5 705.2 668.3 723.2 701.4 712.8 670.2 691.1 696.2 661.9 677. 715.3 662. 701.9 717.3 722.1 694.3 701.5 709. 706.9 684.3 693. 680.4 689.4 703.1 690.2 673.5] | [491.5 481.6 461.4 459.6 453.4 448.5 492.4 481.2 467.2 469.5 470.6 470.5 460.4 463.5 474. 509.2 464.7 479.4 470.8 457.9 456.4 449. 465.9 449.8 483.1 480.6 466.3 472.6 467.4 478.7 469.9 453.8 451.6 468.8 495.9 468.6 476.2 473. 472. 474.6 441.3 466.4 461.1 476.1 495. 478.5 471. 468.7 466.2 478. 449.5 474.2 456.4 466.8 477.9 505.9 471.4 489.2 467. 460.1 468.1 469.6 452.4 472.2 453. 497.3 469.2 486.9 493.1 463.3 466.5 479.2 479.8 477.7 468.6 487.1 466.4 471.8 475.1 480. 472. 458. 463. 479.5 472.3 464.9 474.1 468.3] | [345.3 342.5 343.3 337.3 342.4 339.7 341.6 341.5 336.4 337.5 349.4 348.5 344. 324.9 333.9 353.4 341.2 338.1 340.4 359.5 353.9 345.2 340.4 340.3 339.4 338.8 325.2 336.7 348.1 335.3 355.9 340.8 341.1 339.1 332.7 346.9 338.1 343.4 339.2 358.2 339.7 347.2 343.1 336.4 341. 331.7 348.9 314.2 336. 332.2 344.2 340.3 335.2 350.2 339.9 354. 337.3 351.6 339.7 337.9 324.6 337.7 333.2 340.1 350.4 342. 349.7 337.8 325.3 333.7 326.6 339.6 347.1 325.2 357.1 333.8 333.1 343.1 337.3 336.8 335.8 348.2 335.1 332.8 359.7 327.6 334.5 335.9 334.6 327.1 337.9 352.9 347.5 324.9 333.2 345.8 340.7] | [241.5 243.8 236.3 235.7 238.9 254. 242.6 250.5 247. 246.1 246.1 240.5 236.1 256.9 242.3 251.1 234.6 226.2 232.3 231.2 242.2 252.9 239.1 258.7 250.5 240.1 235.8 250.4 245.5 249.8 239.2 239.1 227.2 247.5 247.2 235.7 240.8 246.9 228.8 243.7 236.1 236.3 234.9 244.3 251.6 247.1 239.7 236.7 235.9 245.1 230.2 241.8 230.7 243.2 244. 248.9 241.7 239.5 232.7 226.5 222.8 239.4 224.4 256.1 230.6 244.1 247.5 240.4 231.2 229.2 231.4 232.7 239.3 242.3 238.5 243.3 238.4 236.2 243.7 230.4 231. 244.8 234.5 227.6 221.9 237.6 247.8 240.4] | [196.5 191. 187.4 188.5 181.5 171. 193.1 169. 170.1 179.1 173.5 177.7 188.6 166.8 195.1 174.4 179.3 183.7 184.9 192.4 176.3 178.2 181.1 179.1 200.4 172. 193. 170. 182.8 167.5 613.2 171.8 194. 169.8 183.7 175.3 184.1 174.5 203.2 171.6 194.9 172.5 194.1 171.2 186.2 182.8 196. 177.9 173.4 180.2 195.3 171.6 192.8 175.6 590.3 181.9 178. 189.2 181.6 183. 177.4 182.7 182.7 178.3 174. 184.8 184.9 173.5 164.2 171.9 181.2 181.2 173.4 674.4 175. 187.9 184.5 183.8 180.9 172. 172.7 188. 179.7 176.3 171.9 169.3 181. 191.9] | [220.9 207.4 197.3 182.4 185.3 191.7 208.2 177.5 197.3 195. 195.5 125.6 201.2 204.9 198.1 196.6 179.2 226.9 183.5 198.5 214.3 128.4 217.8 112.4 233.2 185.5 215.8 181.2 196. 181.3 195.3 195.1 200. 189.1 223.8 182.5 234.9 196.1 212.3 196.9 129. 184.3 195.3 200.3 211.1 134.9 190.1 209.7 154.9 184.8 216.1 184.7 227.8 179.5 176.1 218.1 170.4 192.8 214.7 211.9 187. 196.8 214.9 115.6 176.1 171.4 204.4 198.4 124. 119.1 185.7 192.4 175.7 188.4 198.7 176.2 121.3 178.5 173. 117.8 123.8 209. 220.5 196.7 204.5 205.6 189.5 213.1] | [128.1 103.5 134.1 110.2 121.6 152.1 81.7 105.4 94.3 122.5 134.6 97.3 117.1 137.7 83.9 122.6 129. 105.2 95.2 98. 129.3 108.2 100.6 134.5 113.9 112.7 83.3 132.5 123.5 91.3 117.3 125.1 93.4 119.5 119.7 102.8 129.3 113.8 111.4 119.2 128.1 139.5 135.5 97.1 156.7 114.2 130.8 109.6 148.6 124.3 107.2 114. 142.6 126.5 132.3 99.7 121.6 125.1 106. 128.8 133.1 98.7 137.5 134.6 72.9 117. 65.1 113.5 106.4 119.3 81.4 81.8 79. 124.9 145.4 101.6 102. 89.5 101.9 110.3 86.2 64.1 86. 115.2 100. 108.7 114.9 3.6] | [ 8.3 71.5 85.6 103.1 120.5 105.4 113.4 95. 68.3 9.5 128.8 117. 9.1 95.4 120.4 1.8 107.5 120.5 72.7 106.2 130.9 92.5 90.5 106.6 113.2 131.1 111.5 123.1 94.4 102.6 119.9 119.1 113.8 105.5 103.8 126.9 87.3 102.7 125.2 117.1 122.8 113.1 130.6 137.1 131.1 146.5 131.8 105.6 126. 126.5 83.8 123.9 122.1 104.3 127.9 103. 122.4 143.6 122.3 126.1 103.2 109.4 104.2 104.8 104.1 88.9 119.4 127.9 111.4 107.4 130.5 90.7 93. 131.8 125.9 115.5 51.7 116.8 130.1 107.6 100.9 141.6 121.7 107.1 131.3 110.6 126.4 89.4 116.3 119.6 143. 134. 113.1 118.5] | [ 9.8 141.8 128.1 110.9 104.4 56.2 92.3 111.9 127.2 87.5 140.7 131. 94. 123.5 139.5 90.1 102.5 1.2 126.9 91.7 147.3 1.1 10.3 176.6 82.9 90.8 126.6 74.1 160.2 117.3 139.1 164.8 93.9 132.6 112.8 140.1 54. 80.5 81.5 132.3 101.7 143.6 132.1 80.3 112.9 110.7 130.5 82.9 145.6 72.2 135.2 118.9 139.8 96.8 142.1 115.8 138.3 119.4 149.1 153.1 122. 159.2 127.1 113.2 126.4 139.9 140.8 89.8 156.8 85.1 105.4 114. 102.2 79.1 148.3 200.8 111.7 148.7 115.7 97.1 111. 107. 122. 112.8 147.8 133. 89.9 129.2 159.6 141.8 96.5 185.9 140.7 124.1] | [223.2 178.7 125. 141.8 156.5 164.6 157.6 153. 121.7 69.9 110.1 112.8 132.9 116.3 96. 172.8 84.3 121.2 4. 174.6 108.2 1.6 147.4 94.2 116.3 2. 130.5 130.2 182.1 169.8 112.9 217.7 131.8 192.3 144.9 170.9 2.1 122.9 138.1 128.6 143.2 150.3 174.4 133.9 158.5 103.1 92.3 112. 195.5 124.6 170.9 142.4 165.7 191.7 186.8 183.1 202.8 166.5 136.4 135.2 164.4 180.6 151.3 171. 169. 88.8 149.6 198.4 121.5 180.9 97.6 53.2 2.5 107.1 61.5 192.7 151.1 205.7 156. 183.9 110.7 157.6 140.7 90.9 216.4 177.7 97. 163.3 121.7 159.4 39.4 130.6 186.5 89.7] | [184. 220.6 99.5 170.7 599. 113.5 103.1 105.6 121.3 58.2 234.8 136.4 108.5 219.4 316.9 3. 269.5 40.2 161.8 398.6 118.1 332. 248.4 254.9 181. 305.9 129.5 226.8 200.8 286.2 274.2 188. 473.1 273.1 234.1 244. 244.5 311.4 354.2 280.5 257.4 574.8 308.8 306.4 111.4 276.4 221.4 245. 417. 255.8 437. 6.7 451.3 448.6 188.7 205.2 478.4 307.2 244.4 444.7 464.7 805.8 191. 240.8 4.2 64. 260.6 268. 230.1 173.7 270.6 290.8 164.1 316.3 266.4 266.4 255.4 394.6 293.6 4. 276. 593.8 161.4 176.4 221. 138.4 338.6 437.3 638.8 800.9 616.8 708.5 673.5 775. ]
12.001 | near_douglas_NIST | https://osf.io/qr6w2/download/ | 20200309_near_douglas_NIST/20200309_near_douglas_NIST_Magnitude_T1map.nii.gz | Magnitude | Jamie Near | Douglas Brain Imaging Centre | Siemens | Prisma fit | 3 | NIST | 130 | 42 | 22 |  |  | t2_tse_dark-fluid_cor | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 16 | [50, 400, 1100, 2400] | 287 |  |  |  |  | [1913. 1894. 1909. 1917.9 1930.1 1862.5 1906.3 1947.2 1954.6 1932.1 1943.4 1981.6 1871. 1933.6 1976.8 1949.2 1978.8 1981.9 1932.1 1921.2 1981.1 1894. 1941. 1938.1 1919.7 1924.3 1938.2 1989. 1982.2 1940.8 1945.8 1988. 1898.1 1955.9 1917.3 1967.2 1948. 1952.8 1961.3 1936. 1934.5 1938.3 1965.3 1943.7 1959.5 1953.3 1944. 1936.4 1944. 1960.5 1941.7 1921.4 1926.2 1940. 1971.9 1922.8 1948.2 1964.3 1973.2 1955.4 1910. 1907.9 1978.7 1945.5 1947.2 1992.5 1939.6 1962.4 1977.2 1986.5 1949.3 1947.8 1932.3 1968.6 1956.8 1977.7 1985.8 1998.7 1967.3 2002.4 1939.2 1928.8 1948.1 1950.1 1949.2 2005.6 1984.5 1965. 1954.1 1934.4 1895.5 1950.5 2010. 1965.1 1938.4 1962.8 1910.7] | [1447.3 1379.5 1323.3 1331.5 1375.5 1435.8 1428.1 1271.1 1307.7 1371.8 1362.2 1354.5 1313.4 1428.1 1416.4 1278. 1372.1 1377.3 1368.7 1371.7 1395.4 1406.5 1362.1 1424.7 1263.4 1380.2 1366.4 1344.1 1366.8 1400.8 1382.2 1348.6 1398.1 1338.9 1283.7 1367.6 1340.3 1398. 1374.2 1380.4 1387.9 1385.4 1395.7 1371.3 1347.5 1367. 1363.2 1399. 1360. 1408.7 1355.1 1398. 1377.5 1377.4 1339.5 1381.6 1372.6 1400.9 1387.6 1407. 1355.6 1390.9 1382.6 1404.6 1277.4 1379.5 1333.5 1390.7 1367.8 1347.6 1393.3 1385.4 1384.5 1373.5 1362.8 1356.4 1337.6 1382.8 1379.8 1390.3 1351.3 1385.6 1360.4 1360.6 1325.3 1314.2 1328.5 1376.6] | [842.8 819.5 817.1 819.7 828.9 850.6 825.2 821.9 836.9 841.9 837.4 827.2 815.4 824.6 836.6 844.8 836.6 824.3 821.9 827. 836.1 852.4 816.7 843.9 844.1 852.3 829.4 841.3 840.1 845.1 826.6 834.7 839.6 815.1 862.6 842.7 850.9 840.2 823.4 830.1 848.2 829.3 848.9 814.7 876.8 848.4 863. 838.6 886.2 826.2 839.6 828.7 857.3 828.3 877.1 848.2 865.7 841.1 839.4 825.1 844.2 824.2 852.1 829.3 863. 858.2 852.3 864.3 852.5 850.5 837.5 836.6 845.5 828.4 874.4 859.9 846.2 843.3 841.1 842.3 858. 824.7 860.8 860. 860.8 856.4 853.8 828.1] | [432.2 432. 433.6 432.8 433.8 437.1 440.4 437.4 434.1 432.5 433.3 435.2 598.8 432. 442.6 435. 432.9 438.1 437.3 436.5 436. 432.9 606.2 431.1 436.7 437.4 436.6 439.1 438.4 440.9 437.3 431.2 434.5 613.8 434.9 442.6 439.6 593.8 596. 595.9 438.2 434.4 432.2 609.3 598.8 438.7 595.8 444. 442.6 440.9 599.1 594. 429.2 605.1 606.5 438.2 599.7 598.6 603.7 440.3 604.5 596.7 430.1 602.9 610.9 592.7 609.2 593.3 597.2 607.3 434.8 606.1 430.1 612.1 603.1 431.9 608.3 609.8 432. 602. 614.2 429.9 609. 605.9 604.7 604.4 618.2 427.5] | [418.2 409.6 411.2 413.1 411.9 417.9 419.4 419. 425.8 424.7 427.4 431.8 426. 415.1 429.3 422.9 429. 417.7 419.9 422. 417.7 424.4 428. 419.3 416.7 435. 422.7 434.5 430.5 432. 435. 420.8 429.2 418.2 420.7 443.7 428.7 438. 421.1 428.6 426.2 431.4 422.8 433.5 426.5 449.6 432.2 437.1 430.1 443.8 423.4 437.6 425.5 440.8 425.6 447.6 429.8 446.6 427.7 426.1 436.1 432.4 431.9 431.6 425.8 438.9 431.1 432. 435.4 434.5 431. 424.6 435.5 420.9 442.4 426.9 438.2 428.7 424.5 423.7 426.8 431.9 423. 427.6 437.6 422.6 426.3 430.9 426.1 422.5 417.5 436.3 431.5] | [309.3 315.5 310.8 307.3 309.3 318.7 327.2 309.3 307.3 303.5 304.9 301.2 321. 329.6 304.6 316.3 311.7 311.3 314. 312.6 306. 303.5 337.2 309. 316.9 314.7 308. 313.2 313.2 309.8 311.1 306.4 318.3 326.9 309.8 320.8 312.5 319.8 312.3 313.6 317.6 306.4 314.6 326.5 319.3 319.6 321. 319.2 317.1 318. 310.4 316.1 310. 319.6 322.6 319.8 319.7 320.4 318.9 318.3 318. 311.4 318. 314.1 319.8 310.1 325. 311.1 320.7 314.5 320.6 312.7 314.8 316.9 312.6 313.4 326.6 312. 313.9 315.7 310.8 311.7 309.4 304. 310.4 322.6 306.6 309.1 308.1 313.5 310.7 305.6 316.9 303.5 298.9 299.7 297.5] | [209.4 225.1 226. 216. 212.5 211.6 208.6 231.2 218.8 208.1 208.6 211.7 210.1 204.9 220. 233.3 224.5 214.5 225.2 217.8 219.2 215.3 215.1 216.3 209. 244.4 209.8 228. 215.3 215.8 217.3 216.1 211.4 212.3 222.7 232.3 213.6 226.7 215.2 227.9 211.7 217.2 222.4 209.2 228.9 224.5 223.9 225.2 219.9 218.6 215.5 219.1 212.1 220.3 226.5 223.8 222.3 223.4 220.3 218.4 219.9 222.2 210.6 221.4 214.9 228.5 209.5 223.9 216.4 226.1 208.6 214. 220.4 210.1 231.8 204.4 217.9 214.5 215.7 214.9 216.8 209.7 214.6 204.4 211.7 208.8 210.3 209.6] | [148.3 151.5 152.7 142.9 137.6 136.5 144.3 156.1 150.7 131.6 131.4 135.1 134.5 129.4 167.1 154.5 157.3 127.4 141.9 144.1 138.8 136.7 142.7 140.1 142. 165.5 135.1 142.7 142.8 137. 145.5 143. 134. 147.1 153.4 156.9 133.7 145.7 140.4 147.1 139.2 141.1 148.6 139.5 159.9 147.4 142.8 146.7 147.8 138.9 147.5 141.3 147.5 141.6 158.6 146.5 140.5 142.9 144.5 139.3 143.8 138.4 146.8 139.1 148.5 154. 130.8 143.3 138.5 146.8 141.4 140.9 145.5 138.6 159.3 132.1 133.1 136.9 132.3 136.6 133.8 129.2 151.7 123.2 129.2 134.4 137.6 134.3] | [250.9 204.6 129.6 128. 125.5 126.5 230.9 120.1 118.8 112.3 100.3 94.7 98.1 111.6 251.7 115.9 125.1 108.4 85.3 85.8 91.6 88.3 87. 104.1 122. 132.1 108. 77.4 97.9 90.8 94.6 91.9 98.9 89.3 113.8 133.5 99.1 109.2 99.1 100.2 100.1 100.1 94.1 103.9 133.2 112.9 94.7 105. 103.8 107.1 94.1 98.8 107.8 100.6 142.7 97.7 113. 98.3 110.4 88.9 110.5 86.8 110.5 96. 142.2 96. 101.6 96.4 105.3 89.5 100.3 102.1 108.7 93.3 108.7 111. 79.9 99.6 108. 101.2 96.9 96.7 95.6 87.7 93.4 64.8 89.6 99.6 70.4 97.4 79.3 74.5 2.7] | [139.4 152.2 154.2 159. 162.9 157.5 148.6 151.6 138.2 130.5 123.1 130.4 144.2 158.2 133.2 98.3 123.4 126.5 141.6 132.9 139.8 133.9 130.3 162. 158.1 92.6 135.7 131.7 115.4 133.8 127.1 137.6 137.6 131.5 134.6 120.7 133.8 141.4 137.9 145.1 142.6 130.8 142.7 129.2 130.5 147.8 122.4 138. 123.6 123.6 127.6 138.3 124.1 145. 116. 132.8 126.6 135.1 125.9 148.6 123.6 132.8 122.5 135.2 106.1 132.4 117.1 132.6 124.3 121.7 126.4 118.7 123.8 124.6 125.6 122.6 123.4 132.1 134.9 128.5 117.2 131.4 105.8 119.4 117.7 114.3 116.3 116.5] | [113.2 126.1 125.9 116.3 131.9 122.4 109.9 102.2 100.1 105.3 129.7 130.4 100.2 95.3 106.7 106.9 111.3 101.2 104.5 112.3 143.8 103.4 107.8 114.8 100.9 99.6 97.2 111. 113.7 128.9 133.2 101.3 118.6 94.4 112.3 115.8 112.1 111. 111. 116. 142.1 125.1 109.8 106.2 116.7 106.2 106.5 97.1 119.2 100.9 118.7 144.3 124. 110.1 106.6 121.3 105. 127.5 107.6 115.5 100.7 123.5 136. 129.8 110.9 115.3 108.6 116.5 98.7 106.7 114.7 107.2 139.7 105.4 131.4 113.8 123.1 123.8 123.2 114.8 124.7 143.6 114.6 113.6 121.4 121.4 112.1 119.4 123.7 124. 114.4 117.2 122.2 120.3] | [1.223e+02 1.937e+02 1.782e+02 1.397e+02 1.577e+02 1.446e+03 2.899e+02 1.344e+02 1.728e+02 5.600e+01 1.551e+02 1.615e+02 1.529e+02 8.105e+02 3.251e+02 4.000e+00 1.790e+02 1.600e+01 1.601e+02 1.391e+02 1.471e+02 1.558e+02 1.323e+02 7.450e+02 2.465e+02 1.011e+02 1.959e+02 1.600e+01 1.326e+02 1.251e+02 1.401e+02 1.638e+02 1.528e+02 1.850e+02 9.500e+00 1.559e+02 1.703e+02 1.300e+01 1.232e+02 1.365e+02 1.041e+02 1.330e+02 1.752e+02 1.153e+02 1.415e+02 1.525e+02 1.409e+02 1.485e+02 1.100e+00 1.141e+02 1.117e+02 1.218e+02 1.215e+02 1.587e+02 1.113e+02 1.240e+02 1.388e+02 1.483e+02 1.521e+02 1.347e+02 1.218e+02 1.249e+02 1.273e+02 1.179e+02 1.509e+02 1.283e+02 1.071e+02 1.157e+02 1.296e+02 1.531e+02 9.950e+01 1.255e+02 1.112e+02 1.150e+02 1.104e+02 1.118e+02 1.375e+02 1.025e+02 1.426e+02 1.440e+02 1.267e+02 1.175e+02 1.377e+02 1.046e+02 1.206e+02 1.425e+02 1.411e+02 1.156e+02 1.185e+02 1.431e+02 1.329e+02 1.431e+02 1.458e+02 1.405e+02] | [185. 170.3 163.8 173.2 205.9 181.5 151.3 148.9 156.2 138.9 157.2 186.9 184.5 140.6 127.2 149.2 152.8 148.8 157.2 152.9 143.2 194. 144.8 167.2 158. 154.4 160.2 150.8 175.6 178.8 169.5 192.3 166.3 172.5 151.3 159.8 145.7 165.8 169.4 135.8 177.8 160.7 189.6 165.7 153.9 147.7 147.2 142.6 178.5 142.4 164.8 177.4 164.6 189.3 156.2 159.8 137.2 148. 127.4 141.2 139.6 173.6 172.6 157.1 189.1 181.8 147.9 157.4 158.6 163.9 175.2 165.2 181. 164.2 190. 157.8 196.8 165.6 149.1 167.8 157. 183.3 160.1 159.2 158.1 176.3 179.4 177.5 167.2 150.5 101.8 167.1 168.3 109.1] | [250.1 196.8 158.8 246.5 248.8 162.2 194.1 233.4 255.8 209.7 431.2 230. 197.6 249.3 247. 152.5 181.4 205.3 229.1 318.6 213.5 254.6 219.5 190.4 241.3 278.6 197.4 244.2 223.9 452. 271.7 272.1 225.6 211.8 212. 208.5 188.8 230.1 234.2 277.3 412.4 293.3 271.8 166.3 261. 192.3 271.4 270.1 261.7 252.7 248.6 437.7 321.9 288.6 140.2 229.3 270.6 204.4 185.7 239.8 228.4 293.4 353.8 285. 298.5 230.4 257.5 291. 328.6 253.8 262.6 274.2 424.9 305.2 257.6 318. 267.4 237.9 231.9 219. 296.4 377.3 295.6 255.9 259.1 280. 317.1 298. 250.3 379.2 333.6 268.3 226.5 215.4]
13.001 | jalnefjord_sahlgrenska_NIST | https://osf.io/tw6eq/download/ | 20201802_jalnefjord_sahlgrenska_NIST/20201802_jalnefjord_sahlgrenska_NIST_Complex_T1map.nii.gz | Complex | Oscar Jalnefjord | Sahlgrenska University Hospital | Philips | Achieva dStream 5.6.1 | 3 | NIST | 130 | 103 | 21.7 |  |  | IR | Spin-echo inversion recovery | ['256 256 1'] | ['1 1 2'] | 2D | 2550 | 14 | [50, 400, 1100, 2500] | 435